## TAPP
Metadata Item | Description | Procedure-Level Tier | Analysis-Level Tier | Data Type | Example / Allowed Content | Comments | Last Update | Keyed By | Literature Assessment
1. Procedure Identification | [8 empty] | N
Procedure Name | A short descriptive name for this analytical procedure, including a version number. Should identify the instrument, the technique, and the scope of what is measured — analyte, material, or feature. | Basic | Read-Only | Text (free) |  |  | 2026-07-24 | (none)
Technique | Top-level analytical technique identifier. | Basic | Read-Only | Controlled list | Solution MC-ICP-MS |  | 2026-06-18 | (none)
Procedure Author | Person(s) or laboratory responsible for developing and registering this procedure. ORCID recommended for individuals. | Basic | Read-Only | Text (free) |  |  | 2026-07-24 | (none)
Laboratory | Name of the laboratory or institution hosting the instrument. | Basic | Editable | Text (free) |  |  | 2026-06-18 | (none)
Laboratory ID | Persistent identifier for the laboratory (e.g., ROR ID). | Advanced | Editable | URI / DOI |  |  | 2026-06-18 | (none)
Procedure Start Date | First date this procedure configuration was used in production. | Basic | Read-Only | Date |  |  | 2026-07-24 | (none)
Funding Source for Procedure Development | Grants and other funding sources that supported instrument acquisition, major upgrades, procedure development, and associated personnel time. Include grant numbers and funding agencies where applicable. | Advanced | Read-Only | Text (free) |  |  | 2026-07-24 | (none)
Procedure Reference(s) | DOI or URL for peer-reviewed publications or technical reports describing, validating, or benchmarking this procedure. | Advanced | Read-Only | URI / DOI |  |  | 2026-07-24 | (none)
Procedure DOI | DOI or unique persistent identifier for the registered procedure used in this analysis, generated upon procedure submission to a data repository. This field is mandatory when submitting analytical data, to encourage formal procedure registration and ensure traceability. If a DOI has been applied for but not yet minted, enter "pending". | N/A | Basic | URI / DOI |  |  | 2026-07-24 | (none)
Session Identifier | Identifier for the analytical session this record describes — the laboratory's own run, sequence or batch identifier as generated by the instrument or acquisition software. The analysis record corresponds to one session, which may cover several samples, and this is the link back to the raw instrument files. Distinct from any persistent identifier a repository mints on submission. | N/A | Basic | Text (free) |  |  | 2026-08-13 | (none)
Analyst | Name(s) of the analyst(s) who performed the analysis session. ORCID is recommended for persistent identification. | N/A | Basic | Text (free) |  |  | 2026-06-18 | (none)
Analysis Start Date | Date on which the analytical session began. For sessions spanning multiple days, use the date of the first session. Format: YYYY-MM-DD. | N/A | Basic | Date |  |  | 2026-06-18 | (none)
Analysis End Date | Date on which the analytical session ended. May equal Analysis Start Date for single-day sessions. Format: YYYY-MM-DD. | N/A | Basic | Date |  |  | 2026-06-18 | (none)
Funding Source for Analysis | Grants and other funding sources that supported the specific sample analyses and data collection reported here. Include grant numbers and funding agencies. Distinct from procedure development funding, which may have been awarded years earlier. | N/A | Basic | Text (free) |  |  | 2026-07-24 | (none)
Coupled Technique(s) | Other analytical techniques applied to the same sample(s) whose results are intended to be interpreted together with data from this procedure. Document coupling with any technique whose results are functionally linked to this dataset — providing calibration inputs, complementary spatial context, or required companion measurements. Use the same controlled vocabulary as the Technique field. Enter "None" if no coupling is intended. | Advanced | Editable | Controlled list / Text | Solution Q-ICP-MS ¦ Solution SF-ICP-MS ¦ TIMS ¦ LA-MC-ICP-MS ¦ None ¦ N/A |  | 2026-07-24 | (none)
Coupling Description | Description of how this procedure is coupled with the technique(s) listed above. Include: (1) the functional relationship — what data or context flows between techniques, or how results are combined (e.g. which output from the coupled technique serves as input to data reduction for this technique); and (2) the analytical sequence — which technique is performed first and why (e.g. non-destructive before destructive). Required when Coupled Technique(s) is not "None". | Advanced | Editable | Text (free) |  |  | 2026-06-18 | (none)
Coupled Procedure DOI | Registered procedure DOI for the coupled technique named above. Provides a stable, citable link to the companion method independent of whether a dataset has been deposited. If the coupled procedure has not yet been registered, enter the DOI of a publication describing the coupled method, or "pending". Enter "None" if no coupling is planned. | N/A | Advanced | URI / DOI |  |  | 2026-07-24 | (none)
Coupled Dataset or Publication Reference | DOI or other persistent identifier for the co-registered dataset or publication where both datasets are reported together. Accepts: a dedicated dataset DOI (if separately deposited), a shared dataset DOI (if co-submitted in the same package), or a publication DOI. Use "same submission" if the coupled dataset is included in this data package, or "pending" if not yet assigned. If coupling is documented through a shared sample identifier only, that information is already captured in Sample Persistent Identifier (Group 2). | N/A | Advanced | URI / DOI / Text |  |  | 2026-06-18 | (none)
2. Samples | [8 empty] | N
Target Material | Type of geological or experimental material being analyzed (e.g., basalt, chondrite, seawater, mineral separate). Constrains the dissolution strategy, chromatographic separation, and matrix-matching requirements. | Basic | Read-Only | Text (free) | Basalt ¦ Chondrite ¦ Seawater ¦ Mineral separate ¦ Synthetic solution ¦ N/A ¦ None ¦ Other: specify |  | 2026-06-18 | (none)
Sample Name | Name or identifier of each sample analysed in this session, as used in the laboratory — a sample mount, section or aliquot counts as one entry where that is what the laboratory tracks. The analysis record corresponds to one session and may cover several samples; fields whose Keyed By column declares 'sample' take one value per entry. Should match the identifier used in associated publications or data tables. | N/A | Basic | Text (free) | e.g., 'Allende MS-1 digest A' ¦ 'NWA 8657 aliquot 2' ¦ 'BHVO-2 replicate 3' |  | 2026-08-12 | defines: sample
Sampling Unit | The physical subdivision of the sample to which one row of reported values corresponds — the unit that is analysed and reported, as distinct from the sample as a whole. State the unit type at procedure level and the units actually analysed at analysis level. Where units nest (e.g. confined tracks within grains), state both levels. | Basic | Basic | Controlled list / Text | Whole sample ¦ Aliquot ¦ Grain ¦ Spot ¦ Analysis point ¦ Phase ¦ Sub-volume ¦ Region of interest ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-11 | defines: sampling unit
Sample Persistent Identifier | Globally unique, persistent identifier for each sample listed in Sample Name. IGSN (International Geo Sample Number) is the recommended standard for geological and cosmochemical samples, as used by Astromat, EarthChem and SESAR. Where a sample and its sub-samples are separately registered, record the identifier at the level actually analysed. | Advanced | Basic | URI / IGSN |  |  | 2026-08-12 | sample
Sample Description | Brief description of sample provenance, form, or preparation state relevant to this analysis. | N/A | Basic | Text (free) |  |  | 2026-06-18 | (none)
Sample Preparation Method | General approach to mechanical sample preparation prior to dissolution (e.g., whole-rock crushing and powdering, mineral separation). Documents how the solid material was conditioned before acid digestion. | Basic | Read-Only | Text (free) | Whole-rock powder ¦ Mineral separate ¦ Dissolved residue ¦ Pre-digested liquid ¦ N/A ¦ None ¦ Other: specify |  | 2026-06-18 | (none)
Digestion Acid(s) | Acid mixture used for sample dissolution. Record acid type(s) and concentration(s) where known (e.g., concentrated HF–HNO3 in Parr bombs is standard for silicate dissolution). For isotope-ratio work, complete matrix dissolution is required to avoid isotopic fractionation during partial dissolution. | Basic | Read-Only | Text (free) | HF–HNO3 ¦ HF–HNO3–HClO4 ¦ HNO3–HCl ¦ HNO3 only ¦ Aqua regia ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-10 | preparation step
Digestion Vessel Type | Type of vessel used for acid digestion. Vessel type constrains maximum temperature and pressure and determines suitability for refractory mineral dissolution (e.g., Parr bombs required for complete silicate digestion at high temperatures). | Basic | Read-Only | Text (free) | Savillex beaker ¦ Parr bomb ¦ TFE/TFM bomb ¦ Carius tube ¦ HPA-S ¦ Open beaker ¦ N/A ¦ None ¦ Other: specify |  | 2026-06-18 | (none)
Number of Digestion Steps | Total number of distinct acid digestion steps required to dissolve the sample. Some procedures require multiple sequential steps — for example, an initial open-beaker HF–HNO3 dissolution followed by a second closed-vessel step for residues retaining refractory minerals, or by an aqua regia reflux to destroy fluoride complexes. | Basic | Read-Only | Integer | 1 ¦ 2 ¦ 3 |  | 2026-07-24 | defines: preparation step
Digestion Temperature | Temperature at which acid digestion was carried out (°C). Procedure specifies the target temperature. | Basic | Editable | Numeric (°C) | e.g., '90–120 °C (hotplate, Savillex beaker)' ¦ '150–190 °C (Parr bomb, refractory phases)' ¦ '220 °C (HPA-S)' |  | 2026-08-10 | preparation step
Digestion Duration | Duration of the primary acid digestion step (hours or days). | Basic | Editable | Text (free) | e.g., '12 h' ¦ '24 h' ¦ '48 h (refractory accessory phases)' ¦ '3 days (multi-step with intermediate dry-down)' |  | 2026-08-10 | preparation step
Final Solution Matrix | Acid type and concentration of the final dissolved sample solution introduced to the instrument. Directly affects instrument response and ion transmission. For isotope-ratio procedures, acid molarity must be precisely matched between samples and bracketing standards to avoid matrix-induced mass bias offsets. Procedures targeting HFSE may add trace HF to the final solution to keep those elements in solution. | Basic | Read-Only | Text (free) | 0.3 M HNO3 ¦ 2% HNO3 ¦ 0.5 M HCl ¦ 3 M HCl ¦ N/A ¦ None ¦ Other: specify |  | 2026-06-18 | (none)
Sample Aliquot Mass or Volume | Mass (mg) of solid material digested or volume (mL) of liquid taken for dissolution. Used for yield calculations and concentration back-calculation. | Advanced | Editable | Numeric (mg or mL) | e.g., '50 mg (whole-rock silicate)' ¦ '25 mg (mass-limited meteorite)' ¦ '100 mg (low-abundance trace elements)' ¦ '1–5 mL (aqueous sample)' |  | 2026-08-10 | sample
Chromatographic Separation Applied | Whether chromatographic separation of element groups (e.g., using ion exchange resin) was performed prior to analysis. Separation may be required to remove matrix elements that cause variable mass bias, to eliminate isobaric interferences, or to isolate an element group before measurement (e.g., Fe and Cu–Zn for isotope dilution, Hf separation from REE, removal of Cr and Ni from Fe fractions). It is usually mandatory for isotope-ratio procedures and optional for concentration measurement. Record resin type and elution matrix conditions if applied. Record 'None' if the direct digestion solution is analyzed. | Basic | Read-Only | Controlled list | Anion exchange resin AG1-X8 (HCl) ¦ Cation exchange resin AG50W-X8 ¦ TRU resin ¦ UTEVA resin ¦ AG-MP-1 (HBr) ¦ Thiol resin ¦ Multi-step: specify ¦ None ¦ N/A ¦ Other: specify |  | 2026-06-18 | (none)
Isotope Dilution Spike | Isotopic spike material added to the sample before or during digestion for isotope dilution (ID) analysis or double-spike mass bias correction. For double-spike procedures, the spike must be added and equilibrated before digestion to correct for any isotopic fractionation during dissolution and column chemistry. For concentration ID, the spike must be added before digestion for complete equilibration. Record spike name or certificate reference, spike isotopes, and whether added pre- or post-digestion. Record 'None' if ID or double-spike technique is not used. | Basic | Read-Only | Text (free) | IRMM-3636 (233U–236U double spike) ¦ 117Sn–122Sn double spike (in-house) ¦ 57Fe–58Fe double spike ¦ None ¦ N/A ¦ Other: specify |  | 2026-07-24 | (none)
3. Instrument & Software | [8 empty] | N
Instrument Make and Model | Manufacturer and model name of the MC-ICP-MS instrument. Determines the collector array geometry, available mass resolution modes, and detector complement. | Basic | Read-Only | Text (free) | Thermo Scientific Neptune ¦ Thermo Scientific Neptune Plus ¦ Nu Instruments Nu Plasma II ¦ Nu Instruments Nu Plasma 1700 ¦ Nu Instruments Nu Sapphire ¦ IsotopX Phoenix ¦ N/A ¦ None ¦ Other: specify |  | 2026-06-18 | (none)
ICP-MS Type | Type of ICP-MS system. Multi-collector sector-field instruments combine a magnetic sector and electrostatic analyser with a fixed array of Faraday cups (and optionally ion counters) that simultaneously collect multiple isotope beams, enabling high-precision static multi-collection. | Basic | Read-Only | Controlled list | Multi-collector sector-field (MC-ICP-MS) ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-11 | (none)
Instrument Serial Number or Lab Identifier | Serial number or laboratory-internal identifier for the specific instrument unit. Supports traceability to instrument service records. | Advanced | Editable | Text (free) |  |  | 2026-06-18 | (none)
Torch Type | Type of plasma torch used (e.g., standard quartz, high-matrix, low-flow). | Advanced | Read-Only | Text (free) | Standard quartz ¦ High-matrix ¦ Low-flow ¦ N/A ¦ None ¦ Other: specify |  | 2026-06-18 | (none)
Interface Cone Configuration | Geometry and designation of the sampler and skimmer cones installed during analysis. For Neptune-type instruments, H-cones (standard) and X-cones (high-sensitivity, ~3× higher transmission) are common options; the choice affects ion beam intensity and is a procedure design decision. Jet cones combined with an interface pump upgrade are used on some instruments for enhanced sensitivity. | Basic | Read-Only | Text (free) | Ni H-sampler + Ni H-skimmer (standard) ¦ Ni H-sampler + Ni X-skimmer ¦ Ni Jet-sampler + Ni X-skimmer (with interface pump) ¦ N/A ¦ None ¦ Other: specify |  | 2026-07-24 | (none)
Sampler and Skimmer Cone Material | Material composition of the sampler and skimmer cones. Nickel (Ni) is standard for HNO3 matrices. Platinum (Pt) cones are used for HCl-rich matrices (e.g., 6 M HCl in Fe chemistry procedures) due to greater corrosion resistance. | Advanced | Read-Only | Text (free) | Ni sampler and Ni skimmer ¦ Pt sampler and Pt skimmer ¦ Ni sampler and Pt skimmer ¦ N/A ¦ None ¦ Other: specify |  | 2026-07-24 | (none)
Torch Depth | Distance between the load coil and the sampling cone tip (mm), also called injector depth or torch position depending on the instrument manufacturer. Affects ion transmission efficiency, oxide formation, and doubly-charged species production. The procedure specifies a target value optimised during initial setup; the analyst confirms or fine-adjusts during session tuning. | Advanced | Editable | Text (free) | 5 mm ¦ Standard position (manufacturer default) ¦ N/A ¦ None |  | 2026-08-12 | (none)
Guard Electrode | Whether a guard electrode (grounded shield electrode) is installed and active on the torch assembly. Capacitively decouples the plasma from the load coil, reducing secondary discharge and improving ion extraction efficiency. | Basic | Read-Only | Controlled list | On ¦ Off ¦ Not installed ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-11 | (none)
Faraday Cup Array Configuration | Description of the multi-collector Faraday cup array. Specify the total number of Faraday cup detectors (fixed and moveable), their labeled positions (e.g., L4, L3, L2, L1, Ax, H1, H2, H3, H4 for a 9-cup Neptune), and the presence and position of any ion counter (SEM) or Daly detector in the collector block. The number and span of cups constrains which isotope combinations can be simultaneously collected. | Basic | Read-Only | Text (free) | Neptune: 9 Faraday cups (L4–H4) + 1 SEM (axial or off-axis); Nu Plasma II: 16 Faraday cups + 3 ion counters ¦ N/A ¦ None ¦ Other: specify |  | 2026-06-18 | (none)
Faraday Cup Amplifier Resistor Values | Resistance values (Ω) of the feedback resistors in the Faraday cup amplifiers. Standard amplifiers use 10¹¹ Ω resistors, yielding 1 V per ~6.24 × 10⁶ ion counts per second. High-gain amplifiers (10¹² or 10¹³ Ω) are fitted to selected cups to improve signal-to-noise for very low-intensity beams (e.g., 234U, low-abundance spike isotopes). Report the resistor value per cup position, or note 'all 10¹¹ Ω' if uniform. | Basic | Read-Only | Text (free) | All cups: 10¹¹ Ω ¦ L1: 10¹³ Ω (234U); all others: 10¹¹ Ω ¦ N/A ¦ None ¦ Other: specify |  | 2026-06-18 | channel
Faraday Cup Gain Calibration Method | Method used to calibrate the relative gain (amplification factor) of each Faraday cup amplifier at the start of or during each analytical session. Relative gain differences between amplifiers cause systematic errors in isotope ratios if uncorrected. Common approaches: instrument internal gain calibration routine (measures signal in each cup sequentially using a common beam), amplifier rotation procedure, or measurement of a known isotope ratio solution in each cup. For instruments with SEM detectors, also describes the SEM-to-Faraday cross-calibration sequence. Specify frequency (per session, per block, weekly). | Basic | Read-Only | Text (free) | Instrument internal gain calibration routine; baseline and gain calibrated at session start and daily ¦ SEM-to-Faraday cross-calibration measured per session using known ratio solution ¦ N/A ¦ None ¦ Other: specify |  | 2026-07-24 | channel
Mass Resolution Setting | Mass resolution mode used in this procedure. MC-ICP-MS instruments allow selection of low resolution (LR; m/Δm ≈ 300–400), medium resolution (MR; m/Δm ≈ 3000–8000), or high resolution (HR; m/Δm ≈ 10,000–12,000). Most isotope ratio measurements use low resolution after chemical separation has removed interfering elements. Medium or high resolution is used when polyatomic interferences (e.g., ArN+, ArO+, ArOH+ on Fe isotopes) cannot be fully resolved by chemistry alone. Procedure registers the mode in use; analyst confirms at session start. | Basic | Editable | Controlled list | LR (m/Δm ≈ 300) ¦ MR (m/Δm ≈ 3000–8000) ¦ HR (m/Δm ≈ 10000–12000) ¦ N/A ¦ None ¦ Other: specify |  | 2026-07-24 | (none)
Collision/Reaction Cell (CRC) Configuration | Whether a collision or reaction cell is installed and its operating mode for this procedure. Most conventional multi-collector instruments have no cell — mass bias and isobaric interferences are handled entirely by chemical separation and mass-bias correction (record 'Not installed'). A small number of modern MC-ICP-MS instruments (e.g., Nu Instruments Sapphire, Thermo Scientific Neoma MS/MS) add a hexapole collision/reaction cell upstream of the collector array, operated in STD/KED/DRC modes analogous to quadrupole ICP-MS/MS instruments, to suppress polyatomic and isobaric interferences before simultaneous multi-collection. Specific gas types, flow rates, and voltages are documented in Group 4. | Basic | Read-Only | Controlled list | Not installed ¦ STD (standard mode, no gas) ¦ KED (kinetic energy discrimination, He gas) ¦ DRC (dynamic reaction cell, reactive gas) ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-11 | (none)
Nebulizer Type | Type and material of the pneumatic nebulizer used for sample introduction. Affects droplet size distribution and sample introduction efficiency. | Basic | Read-Only | Text (free) | PFA self-aspirating MicroFlow (50–100 µL/min) ¦ PFA-ST ¦ Glass concentric ¦ ESI Apex-HF nebulizer ¦ N/A ¦ None ¦ Other: specify |  | 2026-06-18 | (none)
Spray Chamber Type and Cooling Temperature | Type of spray chamber and cooling temperature if thermostatted (°C). Chamber type and temperature control aerosol droplet size and the solvent vapor load reaching the plasma. In some procedures a spray chamber is placed downstream of a desolvating nebulizer to improve signal stability. | Basic | Read-Only | Text (free) | Stable Introduction System (SIS): cyclonic + Scott double-pass ¦ Scott double-pass (glass) at 2°C ¦ Cyclonic at 5°C ¦ Spray chamber downstream of Aridus II ¦ N/A ¦ None ¦ Other: specify |  | 2026-07-24 | (none)
Desolvation System | Desolvating nebulizer or membrane desolvator used upstream of the plasma to reduce solvent load. Lowers oxide production rates and increases sensitivity, and is commonly used for low-abundance analytes or small sample sizes. Note the trade-off: desolvation can introduce additional instrumental mass bias instability relative to wet plasma introduction. Record 'None' if not used. | Advanced | Read-Only | Controlled list | Aridus I ¦ Aridus II ¦ Apex Q ¦ Apex IR ¦ Apex HF ¦ None ¦ N/A ¦ Other: specify |  | 2026-08-10 | (none)
Sample Uptake Rate | Peristaltic pump speed or self-aspiration flow rate for sample introduction (µL/min or mL/min). | Basic | Editable | Numeric (µL/min or mL/min) | e.g., '50–100 µL/min (self-aspirating PFA nebulizer)' ¦ '200–400 µL/min (micro-flow, peristaltic)' ¦ '1 mL/min (conventional concentric)' |  | 2026-08-10 | (none)
Acquisition Software | Name and version of the software used to control the instrument and acquire raw multi-collector data. | Basic | Editable | Text (free) |  |  | 2026-06-18 | (none)
Data Reduction Software | Name and version of the software used for data reduction: baseline subtraction, gain correction, mass bias correction (SSB, double-spike inversion, or internal normalization), interference corrections, and calculation of isotope ratios, δ-values, and concentrations. Include version number. | Basic | Editable | Text (free) |  |  | 2026-06-18 | (none)
4. Measurement Information | [8 empty] | N
Analytical Mode | Top-level declaration of the kind of measurement this procedure covers. Solution ICP-MS has a single routine mode — continuous nebulisation of a digested solution — so this TAPP defines no mode flag columns; the field is still required (Rule 3) so that a procedure record is self-describing and comparable with multi-mode TAPPs across the library. | Basic | Read-Only | Controlled list | Solution nebulisation (continuous) ¦ Flow injection |  | 2026-08-08 | (none)
RF Power | Radiofrequency forward power applied to the plasma (W). Controls ionization efficiency and oxide production rates. | Basic | Editable | Numeric (W) | 1200 ¦ 1250 ¦ 1300 ¦ 1350 |  | 2026-06-18 | (none)
Coolant (Plasma) Gas Flow Rate | Flow rate of the outer (coolant) argon gas stream (L/min). Influences plasma temperature and oxide ion formation. | Basic | Editable | Numeric (L/min) | 13–16 L/min |  | 2026-06-18 | (none)
Auxiliary Gas Flow Rate | Flow rate of the intermediate (auxiliary) argon gas stream between torch body and injector tube (L/min). | Basic | Editable | Numeric (L/min) | 0.8–1.1 L/min |  | 2026-06-18 | (none)
Nebulizer Gas Flow Rate | Flow rate of the carrier argon gas delivered through the nebulizer (L/min). Controls aerosol transport and strongly influences signal sensitivity and stability. Adjusted daily to optimize signal. | Basic | Editable | Numeric (L/min) | 0.8–1.1 L/min |  | 2026-08-10 | (none)
Make-up Gas Flow Rate | Supplementary argon flow added downstream of a desolvation system to maintain total gas delivery to the plasma (L/min). Record 'N/A' if no desolvation system is used. Record 'N/A' where Desolvation System is 'None'. | Advanced | Editable | Numeric (L/min) |  |  | 2026-08-11 | (none)
Plasma Thermal Mode | Whether the ICP plasma is operated under normal (hot) or cool plasma conditions. Normal plasma (>1000 W RF) is standard for solution MC-ICP-MS. Cool plasma is rarely used in MC-ICP-MS due to stability requirements for high-precision isotope ratio work. | Basic | Read-Only | Controlled list | Normal plasma ¦ Cool plasma ¦ N/A ¦ None ¦ Other: specify |  | 2026-06-18 | (none)
Analyte | The chemical species this procedure is designed to determine, recorded at whatever resolution the chemistry is resolved — element(s) for this technique; valence species where a procedure resolves oxidation state; compounds where it resolves molecules. Isotopes are not analytes: isotopes of an element are the same chemical species. In solution MC-ICP-MS the analyte is typically a single element isotope system (e.g., Fe, U, Sn, Pb, Zn) or multiple elements when the measurement simultaneously yields isotope ratios and isotope dilution concentrations. The procedure registers the primary analyte element(s); at analysis level the analyst records the actual suite measured. The complete mass list and cup assignments are documented in Collector Configuration. | Basic | Editable | Text (free) |  |  | 2026-08-12 | defines: analyte
Reported Variables and Units | The final variable(s) this procedure reports and their units — distinct from the fields recording what was *acquired* rather than what is reported. A procedure may acquire many channels and report a small number of derived quantities; without this field a data consumer cannot tell which. Record every reported variable, including intermediate quantities reported alongside final ones (e.g. both the 206Pb/238U ratio and the 206Pb/238U date). Where a reported variable is a nominal property with no magnitude (e.g. a mineral identification), record the variable and give the unit as 'N/A — nominal property'. | Basic | Read-Only | Text (free) | e.g., 'Element concentration (ppm); oxide (wt%)' ¦ 'd56Fe (permil vs IRMM-014); d57Fe (permil vs IRMM-014)' ¦ '206Pb/238U date (Ma); 207Pb/206Pb date (Ma)' ¦ 'Mineral identity (N/A — nominal property); match score (0-1)' |  | 2026-08-12 | defines: reported property
Collector Configuration | Assignment of isotope masses to Faraday cup or ion counter detector positions for this procedure, and the analyte each position serves. Lists the mass number monitored in each collector position with its analyte element where it has one — e.g. 'L3=116Sn (Sn); L2=117Sn (Sn); Ax=120Sn (Sn); H1=121Sb (interference monitor, no analyte)'. Interference-monitor and internal-standard positions serve no analyte and carry no parent element; the analyte list itself is given by the Analyte field and is never inferred from the mass labels here. For static multi-collection procedures, one configuration applies throughout the measurement. For multi-dynamic procedures, list all configurations and the cycling sequence. Different element systems require different mass assignments within the collector array span. | Basic | Read-Only | Text (free) |  |  | 2026-08-12 | defines: channel per analyte
Number of Blocks per Measurement | Number of measurement blocks acquired per sample or standard solution introduction. In MC-ICP-MS, data are structured as blocks of cycles; each block typically begins with a baseline (on-peak zero) measurement before the analytical cycles. Multiple blocks per sample allow inspection of signal stability and within-measurement drift. Procedure specifies the standard number; analyst may adjust for samples with low signal or for high-precision requirements. | Basic | Editable | Integer | 1 ¦ 2 ¦ 3 ¦ 5 |  | 2026-07-24 | (none)
Number of Cycles per Block | Number of measurement cycles acquired per block. One cycle corresponds to a single set of simultaneous Faraday cup readings integrated for the duration specified in Integration Time per Cycle. Total integration time per sample = (Number of Cycles per Block) × (Integration Time per Cycle) × (Number of Blocks per Measurement). Procedure specifies the standard value; analyst may adjust within procedure bounds. | Basic | Editable | Integer | 25 ¦ 40 ¦ 50 ¦ 80 |  | 2026-07-24 | (none)
Integration Time per Cycle | Duration of signal integration per measurement cycle (seconds). Determines counting statistics per cycle. Longer integration times improve shot-noise precision but increase the impact of signal drift within the integration window. For high-gain (10¹² or 10¹³ Ω) amplifier channels, longer integration times are often required to accumulate sufficient charge. Procedure specifies the standard integration time; analyst may confirm or adjust within procedure bounds. Where different isotope channels use different integration schemes, record the time for each channel. | Basic | Editable | Numeric (s) | 4.194 s ¦ 8 s ¦ 10 s ¦ 16.8 s |  | 2026-08-12 | channel
Baseline Measurement Approach | Method and timing for measuring the ion beam baseline (on-peak zero) during acquisition, and the portion of the acquisition it is taken from. Three approaches are in use and the choice depends on how sample reaches the plasma: deflecting the ion beam (beam-off or electrostatic deflector); aspirating an acid blank, which applies only to solution introduction; or, for laser ablation, collecting a defined number of laser-off cycles at the start of the same block, since no blank can be aspirated mid-ablation. State which, and how many cycles or how long. | Basic | Read-Only | Text (free) | Beam deflector on-peak zero; 1 baseline block of 1 cycle (4.194 s) per measurement block ¦ On-peak zero measured before each sample by aspirating 0.3 M HNO3 for 10 cycles ¦ N/A ¦ None ¦ Other: specify |  | 2026-06-18 | (none)
Wash Time Between Samples | Duration of instrument rinse between successive sample solutions (seconds). Controls carryover and memory effects. For isotope-ratio work, complete washout is critical to prevent cross-contamination between samples with differing isotopic compositions. | Basic | Read-Only | Numeric (s) | 120 ¦ 180 ¦ 210 ¦ 300 |  | 2026-06-18 | (none)
Analysis Sequence | Description of the measurement order within a session: how sample solutions, bracketing standards, secondary reference materials, and blanks are interleaved. For SSB procedures, each sample is bracketed by two standards (standard–sample–standard); the bracketing standard is the same isotopic material as the zero-delta reference. For double-spike procedures, standards are spiked to the same level as samples before analysis. | Basic | Read-Only | Text (free) | SSB: sample bracketed by two standards (Std1–Sample–Std2) ¦ Double-spike: spiked standards bracketing spiked samples ¦ SSB + secondary RM every 5 samples ¦ N/A ¦ None ¦ Other: specify |  | 2026-07-24 | (none)
Collision Gas Type | Type of collision gas used in KED mode for polyatomic interference suppression, on collision/reaction-cell-equipped MC-ICP-MS instruments (e.g., Nu Sapphire, Thermo Neoma MS/MS). Record 'N/A' if the instrument has no cell or KED mode is not used. Record 'N/A' where Collision/Reaction Cell (CRC) Configuration does not include KED. | Basic | Read-Only | Controlled list | He ¦ H2 ¦ He+H2 ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-11 | (none)
Collision Gas Flow Rate | Flow rate of the collision gas in KED mode (mL/min), on collision/reaction-cell-equipped MC-ICP-MS instruments. Higher flow rates provide greater interference suppression at the cost of analyte sensitivity. Record 'N/A' if the instrument has no cell or KED mode is not used. Record 'N/A' where Collision/Reaction Cell (CRC) Configuration does not include KED. | Basic | Editable | Numeric (mL/min) | 0.5 ¦ 4.5 ¦ 5.2 ¦ N/A |  | 2026-08-11 | (none)
Cell Exit Discrimination Voltage | Kinetic energy discrimination offset voltage applied at the exit of the collision cell (V), on collision/reaction-cell-equipped MC-ICP-MS instruments. Controls the degree of polyatomic ion suppression. Record 'N/A' if the instrument has no cell or KED mode is not used. Record 'N/A' where Collision/Reaction Cell (CRC) Configuration does not include KED. | Basic | Editable | Numeric (V) | −2 ¦ −5 ¦ N/A |  | 2026-08-11 | (none)
Reaction Gas Type | Type of reactive gas introduced into the dynamic reaction cell (e.g., NH3, O2, CH4), on collision/reaction-cell-equipped MC-ICP-MS instruments. Record 'N/A' if the instrument has no cell or DRC mode is not used. Record 'N/A' where Collision/Reaction Cell (CRC) Configuration does not include DRC. | Advanced | Read-Only | Controlled list | NH3 ¦ O2 ¦ CH4 ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-11 | (none)
Reaction Gas Flow Rate | Flow rate of the reaction gas in DRC mode (mL/min), on collision/reaction-cell-equipped MC-ICP-MS instruments. Record 'N/A' if the instrument has no cell or DRC mode is not used. Record 'N/A' where Collision/Reaction Cell (CRC) Configuration does not include DRC. | Advanced | Editable | Numeric (mL/min) | N/A |  | 2026-08-11 | (none)
Internal Standard Element | Element added at a known concentration to all sample and standard solutions to serve as an internal normalization standard for mass bias correction (SSB + internal normalization procedures). The element must have a known isotopic composition and mass similar to the analyte. Examples: Ni or Cu doping for Fe isotopes (must verify complete matrix separation); Tl for Pb isotopes; Cd for Sb isotopes. Specify element and monitored isotope. Record 'None' if SSB without internal standard or double-spike method is used. | Basic | Read-Only | Text (free) | Cu (65Cu/63Cu for Fe procedures) ¦ Ni (62Ni/58Ni for Fe procedures) ¦ Tl (205Tl/203Tl for Pb procedures) ¦ Cd (113Cd/111Cd for Sb procedures) ¦ None ¦ N/A ¦ Other: specify |  | 2026-07-24 | (none)
Internal Standard Concentration | Target concentration of internal standard element(s) in all measured solutions (µg/L or ppb). Record 'N/A' where Internal Standard Element is 'None'. | Advanced | Read-Only | Numeric (µg/L) | e.g., '5 µg/L In' ¦ '10 µg/L Rh' ¦ '1–10 µg/L (In, Rh, Re multi-element)' ¦ '20 µg/L (high-dilution runs)' |  | 2026-08-11 | (none)
Oxide Production Method and Threshold | Method used to quantify plasma oxide production and the acceptance threshold applied before commencing analysis. Record both the monitored mass ratio(s) and the maximum allowed threshold(s). Measured values are recorded in Oxide Production. CeO+/Ce+ (m/z 156/140) is the standard monitor proxy. Stricter oxide thresholds may be required for analytes sensitive to oxide interferences (e.g., BaO+ on REE isotopes). | Basic | Read-Only | Text (free) | CeO+/Ce+ (m/z 156/140) < 0.5% ¦ CeO+/Ce+ < 0.3%; BaO+/Ba+ (m/z 154/138) < 0.1% ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-12 | (none)
Oxide Production | Measured oxide production ratio obtained during session-start tuning, for the proxy specified in Oxide Production Method and Threshold. Record the measured value and confirm whether the procedure threshold was met. | N/A | Basic | Text (free) | CeO+/Ce+ = 0.3% (threshold < 0.5% met) ¦ BaO+/Ba+ = 0.07% (threshold < 0.1% met) |  | 2026-07-24 | (none)
Peak Flatness Method and Threshold | Method used to verify that every collector sits on a flat region of its peak top simultaneously, and the acceptance threshold applied before analysis begins. Specific to simultaneous collection: a sequential instrument centres one mass at a time, whereas a multi-collector array must have all cups on flat peak tops at once or the measured ratio carries a peak-shape bias that no downstream correction removes. Typically assessed by scanning across the peak plateau, or by comparing ratios measured at slightly offset mass positions. | Basic | Read-Only | Text (free) |  |  | 2026-08-08 | (none)
Peak Flatness | Peak flatness actually achieved in the session, assessed against the criterion registered by the procedure. Companion measured value to Peak Flatness Method and Threshold, following the criterion-versus-measurement split used for Oxide Production. | N/A | Basic | Text (free) |  |  | 2026-08-08 | (none)
5. Data Processing | [8 empty] | N
Mass Bias Correction Strategy | Primary strategy used to correct for instrumental isotopic mass fractionation. In solution MC-ICP-MS, instrumental mass bias is typically 3–5%/amu and must be corrected before isotope ratios can be interpreted geochemically. Four main strategies: (1) Sample-standard bracketing (SSB): alternating sample and isotopic standard measurements; bias interpolated linearly between bracketing measurements. (2) Double-spike: a mixture of two enriched isotopes of the analyte element added before digestion provides an internal monitor of both instrumental and chemical mass fractionation. (3) Internal normalization: an element of known isotopic composition added to samples and standards; measured ratio of normalizing element used to calculate the mass bias factor applied to analyte ratios. (4) SSB + internal normalization: both strategies applied simultaneously for redundant bias correction. | Basic | Read-Only | Controlled list | SSB ¦ Double-spike ¦ Internal normalization ¦ SSB + internal normalization ¦ N/A ¦ None ¦ Other: specify |  | 2026-06-18 | (none)
Mass Fractionation Law | Mathematical law used to parameterize instrumental isotope mass fractionation as a function of mass difference. The exponential law (Maréchal et al. 1999) is most widely used and theoretically best justified for MC-ICP-MS. The linear law assumes fractionation proportional to mass difference. The power law is an alternative formulation. For double-spike procedures, the law is embedded in the inversion algorithm and must be consistent between spike calibration and sample data reduction. | Basic | Read-Only | Controlled list | Exponential ¦ Linear ¦ Power ¦ N/A ¦ None ¦ Other: specify |  | 2026-07-24 | (none)
Double-Spike Isotope Pair | The pair of enriched isotopes used to construct the double spike. Both isotopes must be naturally rare, free from isobaric interferences, and chosen to minimize error magnification in the inversion. The double spike composition must be accurately calibrated before use. Specify both isotope masses and the spike certificate or calibration reference. Record 'N/A' if double-spike method is not used. Record 'N/A' where the procedure does not use a double spike. | Basic | Read-Only | Text (free) | 233U–236U (IRMM-3636) ¦ 117Sn–122Sn (in-house calibrated) ¦ 57Fe–58Fe ¦ 61Ni–62Ni ¦ 204Pb–207Pb ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-11 | (none)
Double-Spike Mixing Ratio | Target proportion of double-spike signal relative to total analyte signal in the spiked mixture, expressed as spike fraction (0–1) or spike:sample ratio. An optimal mixing ratio minimizes error propagation through the double-spike inversion; the optimum is analyte-system specific and is typically determined using the Double Spike Toolbox or equivalent. The achieved mixing ratio may deviate from the target within acceptable bounds (typically ±20% of optimal); the double-spike inversion corrects for actual mixing ratios. Record 'N/A' if double-spike method is not used. Record 'N/A' where the procedure does not use a double spike. | Advanced | Editable | Text (free) | ~0.40 (spike fraction; optimal for 117Sn–122Sn) ¦ ~0.5 spike:sample ¦ N/A ¦ None |  | 2026-08-11 | (none)
Double-Spike Inversion Algorithm | Mathematical algorithm used to simultaneously solve for the true sample isotope ratio and the mass fractionation factor from the measured spike-sample mixture. Common implementations: iterative Newton-Raphson solution, matrix algebraic approach, or published software (e.g., Rudge et al. 2009 Double Spike Toolbox). Specify the algorithm or software reference and version. Record 'N/A' if double-spike method is not used. Record 'N/A' where the procedure does not use a double spike. | Basic | Read-Only | Text (free) | Rudge et al. (2009) Double Spike Toolbox ¦ Iterative Newton-Raphson (Compston & Oversby 1969) ¦ In-house spreadsheet implementing exponential-law inversion ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-11 | (none)
Internal Normalization Element and Isotope Ratio | Element and isotope ratio used as the internal normalization standard for mass bias correction. The measured ratio of the normalizing element is used to calculate the mass bias factor β using the chosen mass fractionation law, which is then applied to correct the analyte isotope ratios assuming β_norm ≈ β_analyte. Specify the element, the monitored isotope ratio (e.g., 113Cd/111Cd for Sb measurements), and the certified or consensus value used for correction. Record 'N/A' if internal normalization is not used. Record 'N/A' where the procedure does not use internal normalization. | Basic | Read-Only | Text (free) | 60Ni/58Ni = 0.38579 (from IUPAC 2021) for Fe-Ni doping ¦ 65Cu/63Cu = 0.44561 for Fe-Cu doping ¦ 113Cd/111Cd = 0.95542 (Abouchami et al. 2012) for Sb ¦ 205Tl/203Tl = 2.38714 for Pb ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-11 | (none)
Isotope Ratio Reported | Specific isotope ratio pair(s) reported as primary data output. The denominator isotope is typically the reference mass used in the delta notation (e.g., 54Fe for Fe isotopes, 235U for U isotopes). Report all ratio pairs routinely calculated and reported. | Basic | Read-Only | Text (free) | 56Fe/54Fe and 57Fe/54Fe ¦ 238U/235U ¦ 122Sn/118Sn ¦ 123Sb/121Sb ¦ 206Pb/204Pb; 207Pb/204Pb; 208Pb/204Pb ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-12 | reported property
delta or epsilon Value Reference Standard | International or community-accepted isotopic reference standard used as the zero-delta anchor for expressing isotopic compositions in delta (δ, per mil) or epsilon (ε, per ten thousand) notation: δⁱX = [(ⁱX/ʲX)sample / (ⁱX/ʲX)reference − 1] × 1000‰. The reporting reference is a per-study or per-publication decision and is specified at analysis level. Specify the standard name, lot or batch number where applicable, and the certified or consensus isotope ratio used for normalization. | N/A | Basic | Text (free) | IRMM-014 (Fe; Taylor et al. 1992) ¦ CRM-112a / SRM-960 (U) ¦ NIST SRM 981 (Pb) ¦ NIST SRM 3102a (Sb) ¦ Sn_IPGP (in-house Sn standard) ¦ JMC Lyon (Zn) ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-12 | analyte
Per-Analyte Calibration Strategy | Approach used to convert measured ion signals to elemental concentrations, where concentration data are also a product of the procedure (e.g., isotope dilution combined with isotope ratio measurement). For procedures producing isotope ratios only (no concentration output), record 'Not applicable (isotope ratios only)'. For double-spike procedures that simultaneously yield concentrations, record the ID calculation approach. Where different elements use different strategies, record each. | Advanced | Read-Only | Controlled list | Not applicable (isotope ratios only) ¦ Isotope dilution via double-spike inversion (concentrations co-determined with isotope ratios) ¦ IUPAC isotope dilution equations (separate ID spike) ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-12 | analyte
Signal Integration Interval Method | Method used to define which portion of the acquisition signal is integrated for each measurement. In MC-ICP-MS, all cycles within a block are typically used unless signal instability is detected. Specify the rule for including or excluding individual cycles. | Basic | Read-Only | Controlled list | All cycles included unless signal instability detected ¦ Visual inspection of cycle-by-cycle data; outlier cycles excluded manually ¦ Automated 2σ outlier rejection per block ¦ N/A ¦ None ¦ Other: specify |  | 2026-06-18 | (none)
Signal Integration Time | Total integration time per sample solution (s). Calculated from blocks × cycles per block × integration time per cycle. Recorded at analysis level as the actual total time used. | N/A | Basic | Numeric (s) |  |  | 2026-06-18 | (none)
Blank / Background Correction Method | Method used to subtract instrument background from analytical signals. In MC-ICP-MS, on-peak zero subtraction (aspirating the same acid matrix as samples while measuring background at the analyte masses with beam deflector or from baseline cycles at the start of each block) is standard. For isotope ratio procedures where procedural blank is minor relative to the sample signal, on-peak zero is often sufficient. For low-level analyses, procedural blank must be measured and subtracted. | Basic | Read-Only | Controlled list | On-peak zero (acid blank) ¦ Procedural blank ¦ On-peak zero + procedural blank ¦ N/A ¦ None ¦ Other: specify |  | 2026-07-24 | (none)
Spike / Outlier Filtering Approach | Criteria used to identify and exclude anomalous cycles or replicate measurements from the calculated isotope ratio mean. Filtering is applied after baseline subtraction and before mass bias correction. | Advanced | Editable | Text (free) | 2σ exclusion on per-cycle ratios ¦ Grubbs test ¦ Visual inspection ¦ None ¦ N/A ¦ Other: specify |  | 2026-06-18 | (none)
Isobaric Interference Corrections Applied | Whether mathematical corrections for residual isobaric or polyatomic interferences are applied in data reduction, supplementary to any interference suppression achieved by chemical separation or mass resolution. In solution MC-ICP-MS, chemical separation is the primary interference mitigation; mathematical corrections address residual contributions. Common examples: 204Hg on 204Pb; abundance sensitivity tailing of 238U onto 236U and 235U. | Basic | Read-Only | Controlled list | Y ¦ N ¦ Analyte-specific ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-12 | (none)
Interfering Species | List of isobaric or polyatomic species mathematically corrected in data reduction. Includes direct isobars (e.g., 54Cr+ on 54Fe+, 58Ni+ on 58Fe+), abundance sensitivity tailing (e.g., 238U tail on 235U and 236U in U isotope measurements), and hydride interferences (e.g., 238UH+ at mass 239). | Basic | Read-Only | Text (free) | 54Cr+ on 54Fe+; 58Ni+ on 58Fe+ ¦ 204Hg+ on 204Pb (corrected via 202Hg monitor) ¦ 238U abundance sensitivity tail on 235U and 236U ¦ None ¦ N/A ¦ Other: specify |  | 2026-08-12 | channel
Interference Correction Method | Mathematical approach used to calculate and remove residual interference contributions. For direct isobars, interference monitor masses (e.g., 202Hg to correct 204Hg on 204Pb) are measured simultaneously and used with natural abundance ratios to calculate the interference contribution. For abundance sensitivity, the tailing factor is measured using a pure standard. | Basic | Read-Only | Text (free) | Monitor mass correction: 202Hg signal × (204Hg/202Hg)natural to correct 204Pb ¦ Abundance sensitivity tailing factor measured weekly from pure U standard ¦ None ¦ N/A ¦ Other: specify |  | 2026-08-12 | channel
Isotope Dilution Data Reduction Method | Mass balance approach used to calculate sample isotope ratios and/or elemental concentrations from spike-sample mixture measurements. For double-spike procedures, isotope ratios and concentrations are co-determined through the spike inversion. For conventional ID concentration procedures (separate concentration spike), IUPAC ID equations are used. Record 'None' if neither isotope dilution nor double-spike methods are used. | Basic | Read-Only | Text (free) | Double-spike inversion (isotope ratios and concentrations co-determined) ¦ IUPAC isotope dilution equations (concentration from separate ID spike) ¦ None ¦ N/A ¦ Other: specify |  | 2026-07-24 | (none)
Memory Effect Mitigation | Procedure applied to minimize carryover of analyte isotopes between successive sample introductions. In MC-ICP-MS, extended rinse periods with the same acid matrix as samples (see Wash Time Between Samples) are the primary mitigation. At data processing level, documents any flagging or exclusion of measurements preceded by samples with significantly different isotopic compositions where carryover may be suspected. | Advanced | Editable | Text (free) | Extended acid rinse between samples; first measurement excluded if preceded by high-signal sample ¦ Blank measurement inspected before each sample; session halted if blank > 0.5% of sample signal ¦ None required for this analyte suite ¦ N/A ¦ None |  | 2026-07-24 | (none)
Uncertainty Level | The convention at which reported uncertainties are quoted — 1-sigma, 2-sigma, or a 95% confidence interval — and whether a measured spread is reported as a standard error or a standard deviation. A reported uncertainty is not interpretable without it: the same numeric value means different things at 1-sigma and at 95% confidence. State the convention applying to all values reported under this procedure, and state each separately where different reported quantities use different conventions. Distinct from Uncertainty Propagation Method, which addresses how component uncertainties are combined rather than how the result is quoted. | Basic | Read-Only | Controlled list / Text | 2σ (95% confidence) ¦ 1σ (68% confidence) ¦ 95% confidence interval ¦ 2 standard errors (2SE) ¦ 1 standard deviation (1SD) ¦ N/A ¦ None |  | 2026-08-10 | (none)
Uncertainty Propagation Method | The approach used to propagate analytical uncertainty through the data reduction chain to the final reported value. State which sources are included in the propagation: counting statistics, calibration standard uncertainty, internal standard uncertainty, drift correction, and any systematic contributions. Distinct from Uncertainty Level, which states the convention at which the resulting uncertainty is quoted. | Advanced | Editable | Text (free) | 2σ of bracketing standard reproducibility propagated to sample ¦ Monte Carlo simulation through double-spike inversion ¦ ISO GUM quadrature propagation ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-10 | reported property
Normalization / Standards-Based Correction | Post-acquisition normalization applied to output isotope ratio data beyond the primary mass bias correction. Examples: correction of δ values for decay of enriched spike isotopes between calibration and use date (double-spike procedures); age correction of δ238U for secular disequilibrium. Record 'None' if no additional normalization is applied. | Advanced | Editable | Text (free) | Spike isotope decay correction applied ¦ Age correction for U-series disequilibrium ¦ None ¦ N/A ¦ Other: specify |  | 2026-07-24 | reported property
Calibration Factor and Determination Method | An externally-calibrated factor that converts the measured quantity into the reported quantity, how it was determined, and its uncertainty. Applies where the conversion depends on a factor calibrated against a reference of independently known value, rather than on the instrument response alone. Distinct from the fields that name the calibration material and that state which approach applies to which analyte, where the technique has them: this field records the resulting factor itself. | Basic | Editable | Text (free) |  |  | 2026-08-12 | reported property
Procedural Blank Level | The measured level of the analytical blank in the session, and — where the reported quantity is a ratio — its composition, since a blank subtracted from a ratio biases the result unless its own composition is known. Companion to the blank correction method, which is procedure-level: this field records what was actually measured. Follows the criterion-versus-measurement split the library applies wherever a procedure sets a threshold and an analysis reports a value against it. | N/A | Basic | Text (free) |  |  | 2026-08-08 | (none)
Analysis Inclusion and Rejection Criteria | The rules determining which individual analyses contribute to a reported aggregate result, together with the outcome of applying them: how many analyses were acquired, how many were included, and on what grounds any were excluded. Distinct from within-analysis outlier filtering, which removes anomalous points inside a single analysis: this field decides which whole analyses enter the reported value. Criteria and outcome are combined in one field because neither is interpretable without the other, following the precision/accuracy precedent. | Basic | Basic | Text (free) |  |  | 2026-08-08 | (none)
Error Correlation Between Reported Quantities | The correlation coefficient between pairs of reported quantities whose uncertainties are not independent, together with the pair it applies to and how it was obtained. Concordia and isochron regressions cannot be reconstructed without it. | Advanced | Basic | Numeric (dimensionless) / Text | e.g., 'rho(206Pb/238U, 207Pb/235U) = 0.83, from the data reduction software error propagation' ¦ 'rho = 0.31 (double-spike inversion covariance)' |  | 2026-08-11 | pair: reported property
Constants and Reference Values Used | Physical constants and reference values used in data reduction to calculate the final reported quantity (e.g., decay constants for age calculation, standard isotope ratios, or other citable reference values used in a correction or calculation), together with their source. Distinct from the Group 6 reference-material fields, which document accepted values for specific calibration/validation materials rather than universal physical constants. Record "None" if no citable, revisable physical constants feed into this procedure's data reduction. | Basic | Editable | Text (free) | e.g., "λ238U = 1.55125×10⁻¹⁰ yr⁻¹, λ235U = 9.8485×10⁻¹⁰ yr⁻¹ (Jaffey et al. 1971); 238U/235U = 137.818 (Hiess et al. 2012)" ¦ "None" |  | 2026-08-12 | (none)
6. Quality Control & Uncertainty | [8 empty] | N
Primary Calibration Standard Name | Name and reference material identifier of the isotopic reference standard used as the bracketing standard (SSB) or zero-delta anchor. This is an isotopic composition standard (not a concentration standard): it defines the isotopic composition against which all sample δ or ε values are normalized. Must be a pure, homogeneous, internationally distributed material with well-characterized isotopic composition. For double-spike procedures, also used for spike calibration. Include the material name, its source or supplier, and a citation for the accepted values used, since results calibrated against different published values for the same material are not directly comparable. | Basic | Editable | Text (free) | IRMM-014 (Fe isotopes; Taylor et al. 1992) ¦ CRM-112a / SRM-960 (U isotopes) ¦ NIST SRM 981 (Pb isotopes) ¦ JMC Lyon Zn standard ¦ NIST SRM 3102a (Sb) ¦ In-house single-element standard solution: specify ¦ N/A ¦ None ¦ Other: specify |  | 2026-07-24 | (none)
Calibration Standard Measurement Frequency | How often the primary isotopic standard is measured relative to unknown samples within a session. For SSB, typically one standard measurement before and one after each sample (standard–sample–standard triplet). Specify the bracketing scheme. | Basic | Read-Only | Text (free) | SSB: one standard before and one after every sample ¦ One standard every 3 samples ¦ Start and end of session only ¦ N/A ¦ None ¦ Other: specify |  | 2026-06-18 | (none)
Secondary Reference Materials | Reference material(s) measured as unknowns to independently assess analytical accuracy. Specify material name and the isotopic composition reference source (certified value, consensus value, or literature compilation). For isotope ratio procedures, secondary RMs are geological materials with published isotopic compositions (e.g., BHVO-2 for Fe isotopes) or pure isotopic standards of certified composition. | Basic | Editable | Text (free) | BHVO-2 (Fe isotopes; Dauphas & Rouxel 2006 compilation) ¦ BCR-2 ¦ NIST SRM 960 or equivalent U standard ¦ NIST SRM 981 (Pb) ¦ AGV-2 ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-12 | defines: standard
In-Run Isotope Ratio Reproducibility and Assessment Method | Reproducibility of isotope ratio or δ-value measurements within a single analytical session and the method used to assess it. For solution MC-ICP-MS, precision is reported as 2σ (two standard deviations) or 2SE (two standard errors of the mean) on replicate analyses of an isotopic standard or secondary reference material run as an unknown during the session. Specify: (1) the metric (2σ, 2SD, 2SE, 95% CI); (2) the reference material or standard analyzed; (3) the number of replicate measurements; (4) the ratio pair or δ value assessed. Example: 2SD = 0.03‰ on δ56Fe (n = 10 BHVO-2 analyses in session). | Advanced | Basic | Text (free) | 2SD = 0.03‰ on δ56Fe (n=10 BHVO-2); session ¦ 2σ = 0.004‰ on δ238U (n=9 CRM-112a brackets) ¦ N/A ¦ None ¦ Other: specify |  | 2026-06-18 | standard x reported property
Between-Session Reproducibility and Assessment Method | Reproducibility of isotope ratio or δ-value measurements across multiple analytical sessions (long-term external reproducibility) and the method used to assess it. Typically expressed as 2SD of repeated independent measurements of a secondary reference material across sessions over weeks to months. Specify: (1) the metric; (2) the reference material; (3) the number of measurements and sessions; (4) the ratio or δ value assessed. | Advanced | Basic | Text (free) | 2SD = 0.030‰ on δ56Fe (n=23 sessions; BHVO-1; Dauphas et al. 2009) ¦ 2SD = 0.005‰ on δ238U (n > 50 analyses; CRM-112a) ¦ N/A ¦ None ¦ Other: specify |  | 2026-06-18 | standard x reported property
Analytical Accuracy and Assessment Method | Accuracy of isotope ratio or δ-value measurements relative to certified or published consensus values and the method used to assess it. For isotope ratio procedures, accuracy is typically assessed by comparing measured δ values for secondary geological reference materials against published interlaboratory compilations or certified values. | Basic | Basic | Text (free) | % deviation from published δ56Fe for BHVO-1 compilation value ¦ Agreement with certified δ238U of IRMM-3636 spike solution ¦ Z-score relative to interlaboratory consensus ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-11 | standard x reported property
Detection Limit | Elemental detection limits, one per reported concentration variable (one per analyte, these being the same set), applicable when isotope dilution concentrations are also reported. For pure isotope ratio procedures (no concentration output), this field is N/A or reports the minimum sample mass required for a measurement at target precision. Specify units µg/g or µg/L and whether values are procedure-typical estimates or session-measured. Record 'N/A' where the procedure produces isotope ratios only and reports no concentrations. | Advanced | Basic | Numeric + unit / Text |  |  | 2026-08-12 | reported property
Detection Limit Method | Method used to calculate detection limits for each reported concentration variable, where concentration data are reported. | Basic | Read-Only | Controlled list | 3σ blank ¦ 3× blank mean ¦ Poisson statistics ¦ N/A ¦ None ¦ Other: specify |  | 2026-08-12 | reported property
Limit of Quantification (LOQ) Method | Method used to determine the limit of quantification, where applicable. | Advanced | Basic | Controlled list | 10σ blank ¦ 10× DL ¦ S/N ≥ 10 ¦ N/A ¦ None ¦ Other: specify |  | 2026-06-18 | reported property
Goodness-of-Fit or Dispersion Statistic | The statistic reported to show whether scatter among the contributing analyses exceeds what analytical uncertainty alone predicts, together with its value. Answers whether a reported aggregate is defensible as a single population. Procedure-level tier is N/A because the value cannot be known before the analysis; the procedure may still state an acceptance threshold, which belongs with the inclusion criteria. | N/A | Basic | Text (free) |  |  | 2026-08-08 | reported property
Additional Notes | Any procedure- or analysis-specific information not captured by a structured field anywhere in this TAPP — including anomalies, deviations from the registered procedure, instrument modifications, and supplementary context. Scope is the whole document, not Group 6: this is the last field of the TAPP and covers all six groups. Use sparingly; a structured field is preferred for anything that can be formally categorised. | Advanced | Advanced | Text (free) |  |  | 2026-08-11 | (none)

## Legends
Procedure-Level Tier | Description | Analysis-Level Tier | Description
Basic | Mandatory for procedure registration. Must be provided to register a valid procedure. | Read-Only | Directly imported from the registered procedure; cannot be changed by the analyst. Changing this value means running a different procedure. Also used for fields relevant only at procedure level — the value is inherited from the procedure record and shown read-only in the analysis form.
Advanced | Optional for procedure registration. Strongly recommended but not required. | Editable | Imported from the registered procedure but may be adjusted within procedure-defined bounds (e.g., daily tuning, minor software updates). The procedure registers the target or typical value; the analyst confirms or adjusts it. Cannot be left void if the procedure-level tier is Basic.
N/A | Not applicable at procedure level. This field captures analysis-level information only. | Basic | Mandatory user input at analysis time. Value comes from the analysis itself and cannot be pre-specified in the procedure.
|  | Advanced | Optional user input at analysis time. Recommended for complete and reproducible documentation.
Mode Column | Definition
Y | This field applies to this analytical mode and should be reported.
N | This field does not apply to this analytical mode.
Keyed By (Rule 7) | What the field's value repeats over
(none) | Scalar — one value per procedure or analysis. The default.
analyte | One value per chemical species determined, at whatever granularity the procedure determines it.
channel | One value per position on the instrument's selection axis — the address, not the signal. Mass, cup, line + crystal, energy-loss edge, wavenumber.
defines: analyte | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per chemical species determined, at whatever granularity the procedure determines it.
defines: channel per analyte | ENUMERATES the channel domain — the header of the child table, not a column in it — and does so once per analyte. One value per position on the instrument's selection axis — the address, not the signal. Mass, cup, line + crystal, energy-loss edge, wavenumber. One value per chemical species determined, at whatever granularity the procedure determines it.
defines: preparation step | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per stage in sample preparation.
defines: reported property | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
defines: sample | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it.
defines: sampling unit | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per subdivision of the physical sample that carries its own row — grain, spot, aliquot, phase, sub-volume.
defines: standard | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per reference material or reference database entry.
pair: reported property | One value per unordered PAIR. One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
preparation step | One value per stage in sample preparation.
reported property | One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
sample
standard x reported property | Cross-product, read as "for each standard, one value per reported property". One value per reference material or reference database entry. + One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.