## Column F
Field | Disposition | Why | Proposed unified value | Agree? (y/n) | Notes | Q (live v7) | SF (live v7) | MC (live v5)
Sample Aliquot Mass or Volume | WRITE (gap) | Empty in all three. Rule 6.4: a module row is not complete until the consumer supplies Column F.
Number of Digestion Steps | OK (identical) | Already consistent. | 1 ¦ 2 ¦ 3 |  |  | 1 ¦ 2 ¦ 3 | 1 ¦ 2 ¦ 3 | 1 ¦ 2 ¦ 3
Digestion Vessel Type | HARMONISE (drift) | Vessels do not depend on the detector. MC alone lists 'TFE/TFM bomb'. | Savillex beaker ¦ Parr bomb ¦ Carius tube ¦ HPA-S ¦ Open beaker ¦ TFE/TFM bomb ¦ N/A ¦ None ¦ Other: specify |  |  | Savillex beaker ¦ Parr bomb ¦ Carius tube ¦ HPA-S ¦ Open beaker ¦ N/A ¦ None ¦ Other: specify | Savillex beaker ¦ Parr bomb ¦ Carius tube ¦ HPA-S ¦ Open beaker ¦ N/A ¦ None ¦ Other: specify | Savillex beaker ¦ Parr bomb ¦ TFE/TFM bomb ¦ Carius tube ¦ HPA-S ¦ Open beaker ¦ N/A ¦ None ¦ Other: specify
Digestion Acid(s) | HARMONISE (drift) | Acids do not depend on the detector. MC adds HF-HNO3-HClO4 and drops Aqua regia. | HF–HNO3 ¦ HNO3–HCl ¦ HNO3 only ¦ Aqua regia ¦ HF–HNO3–HClO4 ¦ N/A ¦ None ¦ Other: specify |  |  | HF–HNO3 ¦ HNO3–HCl ¦ HNO3 only ¦ Aqua regia ¦ N/A ¦ None ¦ Other: specify | HF–HNO3 ¦ HNO3–HCl ¦ HNO3 only ¦ Aqua regia ¦ N/A ¦ None ¦ Other: specify | HF–HNO3 ¦ HF–HNO3–HClO4 ¦ HNO3–HCl ¦ HNO3 only ¦ N/A ¦ None ¦ Other: specify
Digestion Temperature | WRITE (gap) | Empty in all three. Rule 6.4: a module row is not complete until the consumer supplies Column F.
Digestion Duration | WRITE (gap) | Empty in all three. Rule 6.4: a module row is not complete until the consumer supplies Column F.
Chromatographic Separation Applied | KEEP (earned) | MC needs quantitative element separation (TRU, UTEVA, AG-MP-1, thiol) for isotope-ratio work; Q/SF trace-element runs frequently skip separation entirely. |  |  |  | Anion exchange resin (HCl) ¦ Cation exchange resin ¦ AG1-X8 (HBr) ¦ None ¦ N/A ¦ Other: specify | Anion exchange resin (HCl) ¦ Cation exchange resin ¦ AG1-X8 (HBr) ¦ None ¦ N/A ¦ Other: specify | Anion exchange resin AG1-X8 (HCl) ¦ Cation exchange resin AG50W-X8 ¦ TRU resin ¦ UTEVA resin ¦ AG-MP-1 (HBr) ¦ Thiol resin ¦ Multi-step: specify ¦ None ¦ N/A ¦ Other: specify
Isotope Dilution Spike | KEEP (earned) | MC lists double spikes (IRMM-3636, Sn/Fe) because double-spiking is an MC-ICP-MS mass-bias technique. Q/SF single-element spikes serve a different purpose. |  |  |  | IRMM-016 (Fe) ¦ 67Zn spike ¦ None ¦ N/A ¦ Other: specify | Mixed multi-element spike ¦ Single-element spike ¦ None ¦ N/A ¦ Other: specify | IRMM-3636 (233U–236U double spike) ¦ 117Sn–122Sn double spike (in-house) ¦ 57Fe–58Fe double spike ¦ None ¦ N/A ¦ Other: specify
Final Solution Matrix | KEEP (earned) | PARTLY earned: MC genuinely runs lower acid strengths. But MC states molarity (0.3 M HNO3) where Q/SF state percent — a NOTATION difference that is drift. Recommend harmonising units, keeping the values. |  |  |  | 2% HNO3 ¦ 3% HNO3 ¦ 4% HCl ¦ 2% HNO3 + trace HF ¦ N/A ¦ None ¦ Other: specify | 2% HNO3 ¦ 3% HNO3 ¦ 2% HNO3 + trace HF ¦ 6M HCl (intermediate) ¦ N/A ¦ None ¦ Other: specify | 0.3 M HNO3 ¦ 2% HNO3 ¦ 0.5 M HCl ¦ 3 M HCl ¦ N/A ¦ None ¦ Other: specify
Nebulizer Type | KEEP (earned) | MC routinely runs self-aspirating low-flow and Apex-HF nebulisers; the introduction chain differs with the measurement goal. |  |  |  | PFA MicroFlow ¦ PFA-ST ¦ Concentric glass ¦ Micromist ¦ Meinhard ¦ N/A ¦ None ¦ Other: specify | PFA MicroFlow ¦ Meinhard ¦ Micromist ¦ PFA-ST ¦ N/A ¦ None ¦ Other: specify | PFA self-aspirating MicroFlow (50–100 µL/min) ¦ PFA-ST ¦ Glass concentric ¦ ESI Apex-HF nebulizer ¦ N/A ¦ None ¦ Other: specify
Spray Chamber Type and Cooling Temperature | KEEP (earned) | Follows the nebuliser choice — MC's SIS and downstream-of-Aridus arrangements have no Q/SF counterpart. |  |  |  | Cyclonic quartz at 2°C ¦ Scott double-pass at 7°C ¦ Conical with impact bead (uncooled) ¦ N/A ¦ None ¦ Other: specify | Scott double-pass (glass) ¦ Scott double-pass (PFA) ¦ Cyclonic at 7°C ¦ Cinnabar cyclonic at 7°C ¦ N/A ¦ None ¦ Other: specify | Stable Introduction System (SIS): cyclonic + Scott double-pass ¦ Scott double-pass (glass) at 2°C ¦ Cyclonic at 5°C ¦ Spray chamber downstream of Aridus II ¦ N/A ¦ None ¦ Other: specify
Desolvation System | HARMONISE (drift) | Same 4-5 instruments in different order, with spelling drift ('Apex IR' vs 'Apex-IR'). | ApexQ ¦ Aridus II ¦ Aridus I ¦ Apex IR ¦ Apex-IR ¦ Apex-HF ¦ None ¦ N/A ¦ Other: specify |  |  | ApexQ ¦ Aridus II ¦ Aridus I ¦ Apex IR ¦ None ¦ N/A ¦ Other: specify | Aridus I ¦ Aridus II ¦ Apex IR ¦ ApexQ ¦ None ¦ N/A ¦ Other: specify | Aridus II ¦ Aridus I ¦ Apex-IR ¦ Apex-HF ¦ ApexQ ¦ None ¦ N/A ¦ Other: specify
Sample Uptake Rate | WRITE (gap) | Empty in all three. Rule 6.4: a module row is not complete until the consumer supplies Column F.
Nebulizer Gas Flow Rate | HARMONISE (drift) | 0.8-1.1 / 0.8-1.0 / 0.85-1.05 L/min. No instrument reason for three different ceilings; arbitrary near-identical ranges. | 0.8–1.1 L/min ¦ 0.8–1.0 L/min ¦ 0.85–1.05 L/min |  |  | 0.8–1.1 L/min | 0.8–1.0 L/min | 0.85–1.05 L/min
Wash Time Between Samples | HARMONISE (drift) | Q and SF empty, MC populated. Not a divergence — a gap in two of three. | 120 ¦ 180 ¦ 210 ¦ 300 |  |  |  |  | 120 ¦ 180 ¦ 210 ¦ 300
Internal Standard Concentration | WRITE (gap) | Empty in all three. Rule 6.4: a module row is not complete until the consumer supplies Column F.