## TAPP
Metadata Item | Description | Procedure-Level Tier | Analysis-Level Tier | Data Type | Example/Allowed Content | Comments | Last Update | Q-ICP-MS | SF-ICP-MS | MC-ICP-MS | Literature Assessment
1. Procedure Identification | [7 empty] | N | N | N | N
Procedure Name | A short descriptive name for this analytical procedure, including a version number. Should identify the instrument, the technique, and the scope of what is measured — analyte, material, or feature. | Basic | Read-Only | Text (free) | e.g., 'LA-ICP-MS U-Pb Zircon Geochronology v1' ¦ 'LA-SF-ICP-MS U-Th-Pb Monazite Geochronology v1' | Source: In-house TAPP rule — Group 1 template (mandatory in all TAPPs) | 2026-07-28 | Y | Y | Y
Technique | Top-level analytical technique identifier. | Basic | Read-Only | Controlled list | LA-ICP-MS U-(Th-)Pb Geochronology ¦ LA-Q-ICP-MS U-(Th-)Pb Geochronology ¦ LA-SF-ICP-MS U-(Th-)Pb Geochronology ¦ LA-MC-ICP-MS U-(Th-)Pb Geochronology ¦ Other: specify | Source: In-house TAPP rule — Group 1 template (mandatory in all TAPPs) | 2026-07-28 | Y | Y | Y
Procedure Author | Person(s) or laboratory responsible for developing and registering this procedure. ORCID recommended for individuals. | Basic | Read-Only | Text (free) | e.g., 'M. Horstwood (ORCID: 0000-0002-9720-716X), NERC Isotope Geosciences Laboratory' | Source: In-house TAPP rule — Group 1 template (mandatory in all TAPPs) | 2026-07-28 | Y | Y | Y
Laboratory | Name of the laboratory or institution hosting the instrument. | Basic | Editable | Text (free) | e.g., 'NERC Isotope Geosciences Laboratory, British Geological Survey' | Source: In-house TAPP rule — Group 1 template (mandatory in all TAPPs). Merged 2026-07-27: also covers Table 3's 'Laboratory name' item (paper example: 'Dept of Earth Science, University of Everywhere'). | 2026-07-28 | Y | Y | Y
Laboratory ID | Persistent identifier for the laboratory (e.g., ROR ID). | Advanced | Editable | URI / DOI | ROR ID or equivalent URI (https://ror.org/) | Source: In-house TAPP rule — Group 1 template (mandatory in all TAPPs) | 2026-07-28 | Y | Y | Y
Procedure Start Date | First date this procedure configuration was used in production. | Basic | Read-Only | Date | YYYY-MM-DD, e.g., 2016-06-01 | Source: In-house TAPP rule — Group 1 template (mandatory in all TAPPs) | 2026-07-28 | Y | Y | Y
Funding Source for Procedure Development | Grants and other funding sources that supported instrument acquisition, major upgrades, procedure development, and associated personnel time. Include grant numbers and funding agencies where applicable. | Advanced | Read-Only | Text (free) | e.g., 'NERC IP-1332-0512' ¦ 'None' | Source: In-house TAPP rule — Group 1 template (mandatory in all TAPPs) | 2026-07-28 | Y | Y | Y
Procedure Reference(s) | DOI or URL for peer-reviewed publications or technical reports describing, validating, or benchmarking this procedure. | Advanced | Read-Only | URI / DOI | e.g., doi:10.1111/j.1751-908X.2016.00379.x (Horstwood et al. 2016, community reporting standard) | Source: In-house TAPP rule — Group 1 template (mandatory in all TAPPs) | 2026-07-28 | Y | Y | Y
Procedure DOI | DOI or unique persistent identifier for the registered procedure used in this analysis, generated upon procedure submission to a data repository. This field is mandatory when submitting analytical data, to encourage formal procedure registration and ensure traceability. If a DOI has been applied for but not yet minted, enter "pending". | N/A | Basic | URI / DOI | e.g., http://doi.org/10.60520/IEDA/114187 ¦ pending | Source: In-house TAPP rule — Group 1 template (mandatory in all TAPPs) | 2026-07-28 | Y | Y | Y
Analyst | Name(s) of the analyst(s) who performed the analysis session. ORCID is recommended for persistent identification. | N/A | Basic | Text (free) | e.g., 'J. Smith (ORCID: 0000-0001-2345-6789)' | Source: In-house TAPP rule — Group 1 template (mandatory in all TAPPs) | 2026-07-28 | Y | Y | Y
Analysis Start Date | Date on which the analytical session began. For sessions spanning multiple days, use the date of the first session. Format: YYYY-MM-DD. | N/A | Basic | Date | YYYY-MM-DD | Source: In-house TAPP rule — Group 1 template (mandatory in all TAPPs) | 2026-07-28 | Y | Y | Y
Analysis End Date | Date on which the analytical session ended. May equal Analysis Start Date for single-day sessions. Format: YYYY-MM-DD. | N/A | Basic | Date | YYYY-MM-DD | Source: In-house TAPP rule — Group 1 template (mandatory in all TAPPs) | 2026-07-28 | Y | Y | Y
Funding Source for Analysis | Grants and other funding sources that supported the specific sample analyses and data collection reported here. Include grant numbers and funding agencies. Distinct from procedure development funding, which may have been awarded years earlier. | N/A | Basic | Text (free) | e.g., 'NSF EAR-1524297' ¦ 'None' | Source: In-house TAPP rule — Group 1 template (mandatory in all TAPPs) | 2026-07-28 | Y | Y | Y
Coupled Technique(s) | Other analytical techniques applied to the same sample(s) whose results are intended to be interpreted together with data from this procedure. Document coupling with any technique whose results are functionally linked to this dataset — providing calibration inputs, complementary spatial context, or required companion measurements. Use the same controlled vocabulary as the Technique field. Enter "None" if no coupling is intended. | Advanced | Editable | Controlled list / Text | CA-ID-TIMS ¦ SIMS ¦ EPMA ¦ CL/BSE Imaging ¦ Other: specify ¦ None ¦ N/A | Source: In-house TAPP rule — Group 1 template — standard coupling field (mandatory in all TAPPs) | 2026-07-28 | Y | Y | Y
Coupling Description | Description of how this procedure is coupled with the technique(s) listed above. Include: (1) the functional relationship — what data or context flows between techniques, or how results are combined (e.g. which output from the coupled technique serves as input to data reduction for this technique); and (2) the analytical sequence — which technique is performed first and why (e.g. non-destructive before destructive). Required when Coupled Technique(s) is not "None". | Advanced | Editable | Text (free) | e.g., 'A subset of dated grains were subsequently analysed by CA-ID-TIMS to independently validate the LA-ICP-MS age population; CA-ID-TIMS performed after LA-ICP-MS screening identified target grains' | Source: In-house TAPP rule — Group 1 template — standard coupling field (mandatory in all TAPPs) | 2026-07-28 | Y | Y | Y
Coupled Procedure DOI | Registered procedure DOI for the coupled technique named above. Provides a stable, citable link to the companion method independent of whether a dataset has been deposited. If the coupled procedure has not yet been registered, enter the DOI of a publication describing the coupled method, or "pending". Enter "None" if no coupling is planned. | N/A | Advanced | URI / DOI | e.g., pending ¦ None | Source: In-house TAPP rule — Group 1 template — standard coupling field (mandatory in all TAPPs) | 2026-07-28 | Y | Y | Y
Coupled Dataset or Publication Reference | DOI or other persistent identifier for the co-registered dataset or publication where both datasets are reported together. Accepts: a dedicated dataset DOI (if separately deposited), a shared dataset DOI (if co-submitted in the same package), or a publication DOI. Use "same submission" if the coupled dataset is included in this data package, or "pending" if not yet assigned. If coupling is documented through a shared sample identifier only, that information is already captured in Sample Persistent Identifier (Group 2). | N/A | Advanced | URI / DOI / Text | e.g., 'same submission' ¦ pending ¦ None | Source: In-house TAPP rule — Group 1 template — standard coupling field (mandatory in all TAPPs) | 2026-07-28 | Y | Y | Y
2. Samples | [7 empty] | N | N | N
Imaging | Pre-ablation characterisation imaging (typically cathodoluminescence [CL] or backscattered electron [BSE]) used to reveal internal zoning, inherited cores, and metamict domains for selecting valid spot locations, including the imaging instrument and settings used. | Basic | Editable | Text (free) | CL, FEI Quanta600, 10 nA, 17 mm working distance | Source: Table 3 (Horstwood et al. 2016) | 2026-07-30 | Y | Y | Y
Target Material | General description of the U-(Th-)Pb-bearing mineral(s) this procedure is designed to analyse. Used for discoverability and procedure matching. | Basic | Read-Only | Controlled list / Text | Zircon ¦ Monazite ¦ Titanite ¦ Apatite ¦ Rutile ¦ Xenotime ¦ Baddeleyite ¦ Perovskite ¦ Allanite ¦ Other: specify ¦ N/A ¦ None | Source: In-house TAPP rule — Rule 1 shared Group 2 field (standard name mandatory in all TAPPs). Merged 2026-07-27: also covers Table 3's 'Sample type/mineral' item (paper example: 'Detrital zircons'). | 2026-07-28 | Y | Y | Y
Sample Preparation Method | Standard cross-TAPP description of how samples were prepared for analysis (mounting, polishing, coating, etc.). | Basic | Editable | Text (free) | Conventional mineral separation, 1 inch resin mount, 1 µm polish to finish | Source: In-house TAPP rule — Rule 1 shared Group 2 field (standard name mandatory in all TAPPs). Merged 2026-07-27: replaces Table 3's separate 'Sample preparation' item; example is Table 3's own value. Tier corrected 2026-07-28 to Basic/Editable to match LA-Q/SF-ICP-MS TAPP v4. | 2026-07-28 | Y | Y | Y
Sample Name | Name or identifier of the sample analysed, as used in the laboratory. Should match the identifier used in associated publications or data tables. | N/A | Basic | Text (free) | e.g., 'NWA 7533 zrn-14' ¦ 'Jack Hills detrital zircon JH-22' | Source: In-house TAPP rule — Group 2 universal field (mandatory in all TAPPs) | 2026-07-28 | Y | Y | Y
Sample Persistent Identifier | Globally unique, persistent identifier for the analysed sample. IGSN is the recommended standard for geological and cosmochemical samples. | Advanced | Advanced | URI / IGSN | e.g., IGSN: IEDA-100001 (https://www.igsn.org/) | Source: In-house TAPP rule — Group 2 universal field (mandatory in all TAPPs) | 2026-07-28 | Y | Y | Y
Pre-Ablation Surface Treatment | Procedure applied immediately before each analysis to remove surface contamination or condition the sample surface, distinct from general sample preparation. Pre-ablation pulses are typically discarded before signal acquisition begins. | Advanced | Editable | Text (free) | e.g., '10 pre-ablation pulses discarded before acquisition (25 µm, 5 Hz, 2.5 J cm⁻²)' ¦ 'None' | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
3. Instrument & Software | [7 empty] | N | N | N
Laser Manufacturer & Model | Manufacturer and model of the laser ablation system. | Basic | Read-Only | Text (free) | ESI/New Wave Research, UP193FX | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; field renamed from 'Make, Model and Type (Laser Ablation System)'; wavelength/type content moved to 'Laser Wavelength and Type' below. | 2026-07-28 | Y | Y | Y
Laser Wavelength and Type | Wavelength and type of the laser used for ablation, in nanometres. Table 3 (Horstwood et al. 2016) states wavelength only; laser type (e.g., ArF excimer) is not explicitly stated for this instrument in the source example. | Basic | Read-Only | Text (free) | 193 nm | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; field renamed from Group 4's 'Laser Wavelength (nm)' and relocated to Group 3, matching LA-Q/SF-ICP-MS TAPP's placement of wavelength as a fixed hardware fact. | 2026-07-28 | Y | Y | Y
ICP-MS Manufacturer & Model | Manufacturer and model of the ICP-MS instrument. | Basic | Read-Only | Text (free) | Nu Instruments, Nu Plasma HR | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; field renamed/split from 'Make, Model and Type (ICP-MS Instrument)'; mass-analyzer type moved to 'ICP-MS Type' below. | 2026-07-28 | Y | Y | Y
ICP-MS Type | Type of ICP-MS mass analyser and collector configuration. | Basic | Read-Only | Controlled list | Single-collector quadrupole (Q-ICP-MS) ¦ Single-collector sector-field (SF-ICP-MS) ¦ Multi-collector sector-field (MC-ICP-MS) ¦ Other: specify | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; split out of 'Make, Model and Type (ICP-MS Instrument)' as its own controlled-list field, matching LA-Q/SF-ICP-MS TAPP's split; example instrument in Table 3 (Nu Plasma HR) is MC-ICP-MS. | 2026-07-28 | Y | Y | Y
Ablation Cell Type | Type, design origin, and internal volume of the ablation cell. Cell volume is a primary determinant of aerosol washout time and therefore the achievable time resolution of the data. | Basic | Read-Only | Text (free) | In-house built low volume cell, volume ca. 4 cm³ | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; field renamed from 'Ablation Cell and Volume' for cross-TAPP consistency; content (including volume, which LA-Q/SF-ICP-MS TAPP's field does not explicitly request) unchanged. | 2026-07-30 | Y | Y | Y
Sample Introduction | Configuration by which the ablated aerosol is delivered to the plasma, including any co-aspirated tracer or spike solution (e.g., a Tl solution used for instrumental mass bias correction, or a ²³⁵U spike used for isotope dilution). | Basic | Read-Only | Text (free) | Ablation aerosol combined with co-aspiration of desolvated Tl-235U tracer | Source: Table 3 (Horstwood et al. 2016) | 2026-07-28 | Y | Y | Y
Detector Configuration | Type(s) of detector(s) installed in the mass spectrometer used to measure the U-(Th-)Pb isotope beams. | Basic | Read-Only | Text (free) | mixed Faraday-multiple ion counting array | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; field renamed from 'Detection System' for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP. | 2026-07-28 | Y | Y | Y
Instrument Serial Number or Lab Identifier | Serial number or laboratory-internal identifier for the specific instrument unit. Supports traceability to instrument service records. | Advanced | Editable | Text (free) | e.g., 'Nu Plasma HR S/N 1234' ¦ 'Lab unit: NIGL-MC1' | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Interface Cone Configuration | Geometry and designation of the sampler and skimmer cones installed during analysis. | Advanced | Read-Only | Text (free) | Ni sample cone + standard Ni skimmer ¦ Ni sample cone + Ni H-skimmer ¦ N/A ¦ None ¦ Other: specify | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Sampler and Skimmer Cone Material | Material composition of the sampler and skimmer cones. Nickel (Ni) is standard for most matrices; platinum (Pt) is used for high-TDS or organic-matrix samples. | Advanced | Read-Only | Text (free) | Ni sampler and Ni skimmer ¦ Pt sampler and Pt skimmer ¦ N/A ¦ None ¦ Other: specify | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Torch Depth | Distance between the load coil and the sampling cone tip (mm), also called injector depth or torch position. Affects ion transmission efficiency, oxide formation, and doubly-charged species production. | Advanced | Editable | Text (free) | 5 mm ¦ Standard position (manufacturer default) ¦ N/A ¦ None | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Torch Type | Type of plasma torch used (e.g., standard quartz, high-matrix, low-flow). | Advanced | Read-Only | Text (free) | Standard quartz ¦ High-matrix ¦ Low-flow ¦ N/A ¦ None ¦ Other: specify | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Guard Electrode | Whether the guard electrode (also called extraction lens or Fassel torch shield) was active during analysis. | Advanced | Read-Only | Controlled list | On ¦ Off ¦ Not applicable (instrument does not have guard electrode) | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Mass Resolution Setting | Operating mass resolution of the mass analyser. Fixed at unit resolution for quadrupole instruments; selectable (low/medium/high) for sector-field instruments to balance sensitivity against interference suppression; not applicable in the same sense for multi-collector instruments running static low-resolution collection (see Faraday Cup Array Configuration and Collector Configuration for MC-specific detail). | Advanced | Editable | Text (free) | e.g., Unit resolution — quadrupole (fixed) ¦ Low resolution — SF (~300 m/Δm) ¦ Medium resolution — SF (~3000 m/Δm) ¦ High resolution — SF (~10000 m/Δm) | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 (SF-ICP-MS primarily; Q-ICP-MS answer is a fixed constant) | 2026-07-28 | Y | Y | Y
Collision/Reaction Cell (CRC) Configuration | Whether a collision or reaction cell is installed and its operating mode. Long-standard on quadrupole-based instruments; a small number of modern multi-collector instruments (e.g., Nu Instruments Sapphire, Thermo Scientific Neoma MS/MS) also add a hexapole collision/reaction cell upstream of the collector array, operated in the same STD/KED/DRC modes. Single-collector sector-field instruments have no commercially available collision/reaction cell option and resolve interferences by mass resolution alone. | Advanced | Read-Only | Controlled list | STD (standard mode, no gas) ¦ KED (kinetic energy discrimination, He gas) ¦ DRC (dynamic reaction cell, reactive gas) ¦ Not installed ¦ Not applicable (SF-ICP-MS) ¦ Other: specify | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 (Q-ICP-MS and MC-ICP-MS only) | 2026-07-29 | Y | N | Y
Faraday Cup Array Configuration | Description of the multi-collector Faraday cup array: total number of Faraday cup detectors (fixed and moveable), their labelled positions, and the presence and position of any ion counter (SEM) or Daly detector. The number and span of cups constrains which isotope combinations can be simultaneously collected. | Advanced | Read-Only | Text (free) | e.g., 'Nu Plasma HR: mixed Faraday-multiple ion counting array, [n] Faraday cups + [n] ion counters' ¦ 'N/A (not multi-collector)' | Source: In-house TAPP rule — imported from Solution MC-ICP-MS TAPP v3 (MC-ICP-MS only) | 2026-07-28 | N | N | Y
Faraday Cup Amplifier Resistor Values | Resistance values (Ω) of the feedback resistors in the Faraday cup amplifiers. Standard amplifiers use 10¹¹ Ω resistors; high-gain amplifiers (10¹² or 10¹³ Ω) are fitted to selected cups to improve signal-to-noise for very low-intensity beams (e.g., low-abundance Pb isotopes in young or low-U material). | Advanced | Read-Only | Text (free) | All cups: 10¹¹ Ω ¦ L1 (204Pb): 10¹² Ω; all others: 10¹¹ Ω ¦ N/A (not multi-collector) | Source: In-house TAPP rule — imported from Solution MC-ICP-MS TAPP v3 (MC-ICP-MS only) | 2026-07-28 | N | N | Y
Faraday Cup Gain Calibration Method | Method used to calibrate the relative gain (amplification factor) of each Faraday cup amplifier at the start of or during each analytical session. Relative gain differences between amplifiers cause systematic errors in isotope ratios if uncorrected. For instruments with ion-counting detectors, also describes the SEM-to-Faraday cross-calibration sequence. | Advanced | Read-Only | Text (free) | Instrument internal gain calibration routine; calibrated at session start ¦ SEM-to-Faraday cross-calibration measured per session using a known ratio solution ¦ N/A (not multi-collector) | Source: In-house TAPP rule — imported from Solution MC-ICP-MS TAPP v3 (MC-ICP-MS only) | 2026-07-28 | N | N | Y
Acquisition Software | Instrument control and data acquisition software used to collect raw signal data, including version number. | Basic | Editable | Text (free) | e.g., 'Nu Instruments Nu Plasma TRA 2.x' ¦ 'Thermo Qtegra 2.10' ¦ 'Agilent MassHunter 4.6' | Source: In-house TAPP rule — Rule 1 shared Group 3 software field (standard name mandatory in all TAPPs); Table 3 does not request acquisition software as a distinct item | 2026-07-28 | Y | Y | Y
Data Reduction Software | Software used for signal processing, background subtraction, downhole fractionation correction, and age calculation, including version number. | Basic | Editable | Text (free) | Nu Instruments Nu Plasma TRA software | Source: In-house TAPP rule — Rule 1 shared Group 3 software field (standard name mandatory in all TAPPs); content populated 2026-07-28 by splitting the software name out of Table 3's 'Data Processing Package Used/Correction for LIEF' (Horstwood et al. 2016), leaving the LIEF correction method in the renamed Group 5 field 'Elemental Fractionation Correction' | 2026-07-28 | Y | Y | Y
4. Measurement Information | [7 empty] | N | N | N
Analytical Mode | Procedure-level declaration of the analytical mode(s) this procedure covers. This TAPP is scoped to static spot ablation only, consistent with Table 3 (Horstwood et al. 2016). No mode-flag columns are used in this TAPP because only one mode is in scope. | Basic | Read-Only | Controlled list | Spot Analysis | Source: In-house TAPP rule — Rule 3 — mandatory first field of Group 4 in every TAPP. Merged 2026-07-27: also covers Table 3's 'Sampling mode/pattern' item (paper example: 'Static spot ablation'). | 2026-07-28 | Y | Y | Y
Laser Spot Path / Ablation Mode | Sampling mode or ablation pattern used during analysis, as described in Table 3 of Horstwood et al. (2016). | Basic | Read-Only | Controlled list | Spot (stationary) ¦ Transect (continuous line scan) ¦ Raster area (2D elemental mapping) ¦ Stepped line profile ¦ Other: specify | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; reinstates Table 3's 'static' nuance as its own field, matching LA-Q/SF-ICP-MS TAPP's structure of keeping this field distinct from Analytical Mode. Content: Spot (stationary). | 2026-07-28 | Y | Y | Y
Spot Diameter | Diameter of the laser spot as set in the acquisition method (nominal/target value). | Basic | Editable | Text (free) | 25 µm | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; renamed and narrowed to nominal value only; measured/actual value split out to 'Spot Diameter (Measured)' below. | 2026-07-28 | Y | Y | Y
Spot Diameter (Measured) | Diameter of the laser spot as independently measured on the sample or a test material, distinct from the nominal/target value registered by the procedure. | N/A | Basic | Text (free) | 32 µm | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; split out of the former 'Spot Diameter (µm), Nominal/Actual' field, following the target-vs-measured naming convention used elsewhere in the TAPP library (e.g. Oxide Production). | 2026-07-28 | Y | Y | Y
Laser Pulse Duration | Duration of each individual laser pulse. A fixed hardware property of the laser system that determines the thermal vs. non-thermal ablation regime. | Basic | Read-Only | Numeric (ns) | 4 ns | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; field renamed from 'Pulse Width (ns)' for cross-TAPP consistency. | 2026-07-30 | Y | Y | Y
Laser Fluence (Energy Density) | Laser pulse energy per unit area at the sample surface, the primary control on ablation rate and crater morphology. | Basic | Editable | Numeric (J cm⁻²) | 2.5 J cm⁻² | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; field renamed from 'Fluence (J cm-2)' for cross-TAPP consistency. | 2026-07-28 | Y | Y | Y
Laser Repetition Rate | Laser pulse repetition rate. | Basic | Editable | Numeric (Hz) | 5 Hz | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; field renamed from 'Repetition Rate (Hz)' for cross-TAPP consistency. | 2026-07-28 | Y | Y | Y
Ablation Duration per Spot | Total on-sample ablation time per spot, as set in the acquisition method. | Basic | Editable | Numeric (s) | 40 s | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; field renamed from 'Ablation Duration (s)' for cross-TAPP consistency. | 2026-07-28 | Y | Y | Y
Ablation Pit Depth/Ablation Rate | Depth of the ablation pit produced under the registered laser conditions, the method used to measure it, and the resulting per-pulse ablation rate. Characterises the achievable depth/time resolution of the procedure. | Basic | Editable | Text (free) | 16 µm pit depth, measured using an optical microscope, SEM and interferometry, equivalent to 0.08 µm/pulse | Source: Table 3 (Horstwood et al. 2016) | 2026-07-30 | Y | Y | Y
Laser Beam Energy Profile | Spatial energy distribution of the laser beam at the sample surface, and whether a beam homogenizer is installed. A flat-top (top-hat) profile produces more uniform ablation craters than a Gaussian beam. Fixed hardware property. | Advanced | Read-Only | Controlled list / Text | Flat-top (top-hat) — beam homogenizer installed ¦ Gaussian — no beam homogenizer ¦ Not stated ¦ N/A ¦ None | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Carrier Gas and Flow Rate | Identity and plumbing configuration of the gas used to transport the ablated aerosol from the cell to the torch, including the procedure-registered flow rate(s). | Basic | Editable | Text (free) | 100% He in the cell (0.8 L min⁻¹), Ar make-up gas combined using a Y-piece 50% along the sample transport line to the torch. | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; merged from separate 'Carrier Gas' and 'Cell Carrier Gas Flow (L min⁻¹)' fields for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP's combined field. | 2026-07-28 | Y | Y | Y
RF Power (W) | ICP radiofrequency forward power. Session-tunable within procedure-defined bounds during plasma optimisation; consistent with the analogous, tunable treatment of RF Power across LA-Q/SF-ICP-MS and all three solution ICP-MS TAPPs. | Basic | Editable | Numeric (W) | 1300 W | Source: Table 3 (Horstwood et al. 2016) — tier confirmed 2026-07-28 as Editable (not Read-Only) after cross-checking all four sibling ICP-MS TAPPs, which unanimously use Editable | 2026-07-28 | Y | Y | Y
Make-up Gas Flow (L min⁻¹) | Source and estimated flow rate of any make-up gas added downstream of the ablation cell, including any desolvating nebuliser used for tracer co-aspiration. | Basic | Editable | Text (free) | Sourced from Nu Instruments DSN-100 desolvating nebuliser. Neb pressure 24 psi (estimated at 0.7 L min⁻¹) Ar. | Source: Table 3 (Horstwood et al. 2016) | 2026-07-30 | Y | Y | Y
Coolant (Plasma) Gas Flow Rate | Flow rate of the outer (coolant/plasma) argon gas stream that sustains the ICP plasma. Determines plasma volume and stability. | Advanced | Editable | Numeric (L/min) | 14 L/min ¦ 15 L/min ¦ Other: specify | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Auxiliary Gas Flow Rate | Flow rate of the intermediate (auxiliary) argon gas stream that positions the plasma relative to the load coil. Affects ion extraction efficiency and oxide production rates. | Advanced | Editable | Numeric (L/min) | 0.8 L/min ¦ 1.0 L/min ¦ Other: specify | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Plasma Thermal Mode | Whether the ICP plasma is operated under cool plasma or normal (hot) plasma conditions. Normal plasma is standard for U-(Th-)Pb geochronology; cool plasma would reduce sensitivity for these high-mass analytes and is not typically used. | Advanced | Read-Only | Controlled list | Normal plasma (>1000 W RF) ¦ Cool plasma (≤900 W RF) ¦ Other: specify | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Masses Measured | Mass-to-charge positions (isotope masses) this procedure is designed to measure. Defines the scope of the procedure; the analyst may record the specific subset actually measured in a given session. | Basic | Editable | Text (free) | 202–207, 235, 238 | Source: Table 3 (Horstwood et al. 2016) | 2026-07-28 | Y | Y | Y
Reported Variables and Units | The final reported chemical/isotopic variable(s) and their units — distinct from Masses Measured, which records raw acquisition scope. Covers isotope ratios (both common-Pb corrected and uncorrected, where applicable), elemental/isotopic concentrations, and any derived quantities (age, δ or ε notation) with their reference standard where relevant. For non-mass-spectrometry techniques, this field is generalised to any reported physical or chemical property and its units. | Basic | Read-Only | Text (free) | ²⁰⁶Pb/²³⁸U, ²⁰⁷Pb/²⁰⁶Pb, ²⁰⁷Pb/²³⁵U (corrected and uncorrected for common Pb); ²⁰⁸Pb/²³²Th (where Th measured); U concentration (µg g⁻¹); f206c; %concordance; age (Ma) | Source: In-house TAPP rule — new generic field, informed by Table 4 of Horstwood et al. (2016) (the paper's own analytical data-reporting template, distinct from the Table 3 metadata template this TAPP is otherwise built from) and by the precedent field 'Isotope Ratio Reported' in Solution MC-ICP-MS TAPP v3 | 2026-07-28 | Y | Y | Y
Integration Time per Peak/Dwell Times; Quadrupole Settling Time Between Mass Jumps | Count time (dwell time) per mass position for each measured isotope, and the settling time allowed between mass jumps, as reported in Table 3 of Horstwood et al. (2016). Applies to sequential (Q/SF) mass-hopping acquisition; for simultaneous multi-collector acquisition see Number of Blocks per Measurement, Number of Cycles per Block, and Integration Time per Cycle. | Basic | Editable | Text (free) | 200 ms for each isotope | Source: Table 3 (Horstwood et al. 2016); Q/SF-ICP-MS (sequential acquisition) | 2026-07-28 | Y | Y | N
Total Integration Time per Output Data Point (s) | Total duty-cycle time for one complete mass-scan sweep (sum of all per-isotope dwell times and settling times), which sets the time resolution of the down-hole signal. Applies to sequential (Q/SF) acquisition. | Basic | Editable | Numeric (s) | ~1.2 s (N.B. this should represent the time resolution of the data) | Source: Table 3 (Horstwood et al. 2016); Q/SF-ICP-MS (sequential acquisition) | 2026-07-28 | Y | Y | N
Collector Configuration | Assignment of isotope masses to Faraday cup or ion counter detector positions for this procedure. For static multi-collection, one configuration applies throughout the measurement; for multi-dynamic procedures, lists all configurations and the cycling sequence. Includes interference-monitor masses collected alongside analyte isotopes. | Advanced | Read-Only | Text (free) | e.g., L4=²⁰²Hg, L3=²⁰⁴Pb, L2=²⁰⁵Tl, L1=²⁰⁶Pb, Ax=²⁰⁷Pb, H1=²⁰⁸Pb, H2=²³²Th, H3=²³⁵U, H4=²³⁸U ¦ N/A (not multi-collector) | Source: In-house TAPP rule — imported from Solution MC-ICP-MS TAPP v3 (MC-ICP-MS only; Analyte-Specific) | 2026-07-28 | N | N | Y
Number of Blocks per Measurement | Number of measurement blocks acquired per spot. Each block typically begins with a baseline (on-peak zero) measurement before the analytical cycles. Multiple blocks allow inspection of signal stability and within-measurement drift. | Advanced | Editable | Integer | 1 ¦ 2 ¦ 3 ¦ N/A (not multi-collector) | Source: In-house TAPP rule — imported from Solution MC-ICP-MS TAPP v3 (MC-ICP-MS only) | 2026-07-28 | N | N | Y
Number of Cycles per Block | Number of measurement cycles acquired per block. One cycle corresponds to a single set of simultaneous Faraday cup readings integrated for the duration in Integration Time per Cycle. | Advanced | Editable | Integer | 20 ¦ 40 ¦ N/A (not multi-collector) | Source: In-house TAPP rule — imported from Solution MC-ICP-MS TAPP v3 (MC-ICP-MS only) | 2026-07-28 | N | N | Y
Integration Time per Cycle | Duration of signal integration per measurement cycle. Total integration time per spot = (Number of Cycles per Block) × (Integration Time per Cycle) × (Number of Blocks per Measurement). | Advanced | Editable | Numeric (s) | 4.194 s ¦ 8 s ¦ N/A (not multi-collector) | Source: In-house TAPP rule — imported from Solution MC-ICP-MS TAPP v3 (MC-ICP-MS only) | 2026-07-28 | N | N | Y
'Sensitivity' as Useful Yield (%, element) | Instrument sensitivity expressed as useful yield — the percentage of sampled atoms of a given element that are ultimately detected as ions — for the analytical session, following the method cited. | N/A | Advanced | Text (free) | 0.4% U ((#ions detected/#atoms sampled)*100; Schaltegger et al. 2015) | Source: Table 3 (Horstwood et al. 2016) | 2026-07-28 | Y | Y | Y
IC Dead Time (ns) | Dead time of each ion-counting detector channel, used in dead-time correction of high count rates. | Basic | Editable | Text (free) | 6, 9 and 7 ns IC0, IC1 and IC2 resp. | Source: Table 3 (Horstwood et al. 2016) | 2026-07-30 | Y | Y | Y
Signal Smoothing | Description of any signal smoothing device or approach installed between the ablation cell and the ICP-MS to reduce pulse-to-pulse signal variability. | Advanced | Editable | Text (free) | e.g., 'Squid 2.0 (laurin technik)' ¦ 'None' | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
ICP Tuning | Description of the approach used to optimise ICP plasma conditions prior to analysis, including the reference material used for tuning and the acceptance criteria. | Advanced | Editable | Text (free) | e.g., 'Daily auto-tune on 91500; optimised for max sensitivity on ²³⁸U with ThO⁺/Th⁺ <0.5%' | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Oxide Production Method and Threshold | Method used to quantify plasma oxide production and the acceptance threshold applied before commencing analysis. ThO⁺/Th⁺ (mass 248/232) is the standard proxy for U-Th-Pb geochronology, directly relevant to Th/U ratio accuracy. | Advanced | Read-Only | Text (free) | e.g., 'ThO⁺/Th⁺ (248/232) <0.5%' ¦ 'ThO⁺/Th⁺ <0.3%; UO⁺/U⁺ monitored at 0.95–1.05' | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Oxide Production | Actual oxide production ratio measured during this analytical session, using the method specified in Oxide Production Method and Threshold. | N/A | Basic | Text (free) | e.g., 'ThO⁺/Th⁺ = 0.3% (threshold <0.5% met)' | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Doubly-Charged Species Monitor | The mass ratio monitored to estimate doubly-charged ion (M²⁺) formation during instrument tuning. | Advanced | Editable | Text (free) | ¹³⁸Ba²⁺/¹³⁸Ba⁺ (m/z 69/138) ¦ ²³⁸U²⁺/²³⁸U⁺ (m/z 119/238) | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Doubly-Charged Species Production | Measured percentage of doubly-charged ion production for the monitored species at the time of instrument tuning. | Advanced | Editable | Text (free) | <1% (threshold); 0.4% (measured) | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Instrument Warm-up / Session Duration Limit | Minimum warm-up time required after plasma ignition before analyses begin, and any maximum session duration enforced to maintain stable operating conditions. | Advanced | Read-Only | Text (free) | e.g., '≥30 min warm-up after sample change; sessions ≤8 h after tuning' | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Analysis Sequence | Repeating order of primary calibration standard(s), quality control standard(s), and unknown analyses within a measurement session. | Advanced | Editable | Text (free) | e.g., 'Standard-sample-standard bracketing every 5–10 unknowns' | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Baseline Measurement Approach | Method and timing for measuring the ion beam baseline (background) during acquisition, distinct from the data-processing-level blank subtraction method (see Gas Blank, Group 5). For LA-ICP-MS, baseline is measured with the laser off (gas blank), typically before each ablation or block. | Advanced | Read-Only | Text (free) | e.g., 'Laser-off background measured for [n] cycles before each block' ¦ 'N/A (not multi-collector)' | Source: In-house TAPP rule — imported from Solution MC-ICP-MS TAPP v3 (MC-ICP-MS only; content adapted for laser-off baseline timing instead of solution-introduction beam-deflector/acid-blank methods) | 2026-07-28 | N | N | Y
Analysis Location/Spot Coordinates | Location of the analysis on the sample surface: stage coordinates (X, Y) or a description relative to a named reference feature (e.g., mineral rim, growth zone identified by imaging). | N/A | Basic | Text (free) | e.g., 'X = 12350 µm, Y = 8760 µm (stage coordinates)' ¦ 'Core domain, per CL imaging' | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Number of Replicates | Number of replicate spot analyses performed on the same grain or nominal location in this session. | N/A | Basic | Integer | e.g., 1 (single spot) ¦ 3 (three replicate spots per grain) | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Multi-Run Sequential Analysis Design | Whether the procedure uses a single acquisition pass or multiple sequential runs on the same sample location, each optimised for different analytical objectives. | Advanced | Read-Only | Text (free) | Single run ¦ N/A ¦ Other: specify | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Signal Collection Mode | Mode used to collect ion signal across the monitored masses. In peak hopping mode, the quadrupole jumps sequentially between pre-set mass positions and dwells at each peak. Applicable only to quadrupole-based instruments; this concept does not extend to collision-cell-equipped multi-collector instruments (e.g., Nu Sapphire, Thermo Neoma MS/MS), whose collector array performs genuinely simultaneous multi-collection regardless of whether a collision/reaction cell is present upstream. Some such instruments (e.g., Neoma MS/MS) instead use a pre-cell mass filter of a fundamentally different, non-quadrupole design to pre-select a mass window before the cell — a distinct concept not currently captured by any field in this TAPP. | Advanced | Read-Only | Controlled list | Peak hopping ¦ Scanning ¦ Not applicable (SF/MC-ICP-MS) ¦ Other: specify | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 (Q-ICP-MS only) | 2026-07-29 | Y | N | N
Collision Gas Type | Type of collision gas used in KED mode for polyatomic interference suppression. Applies to quadrupole-based instruments and to collision-cell-equipped multi-collector instruments (e.g., Nu Sapphire, Thermo Neoma MS/MS) operating in KED mode. | Advanced | Read-Only | Controlled list | He ¦ H2 ¦ Not applicable (SF-ICP-MS, or MC-ICP-MS without a collision/reaction cell, or STD mode) ¦ Other: specify | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 (Q-ICP-MS and MC-ICP-MS only) | 2026-07-29 | Y | N | Y
Collision Gas Flow Rate | Flow rate of the collision gas in KED mode. Higher flow rates provide greater interference suppression at the cost of analyte sensitivity. Applies to quadrupole-based instruments and to collision-cell-equipped multi-collector instruments (e.g., Nu Sapphire, Thermo Neoma MS/MS) operating in KED mode. | Advanced | Editable | Numeric (mL/min) | 3.5 mL/min ¦ Not applicable (SF-ICP-MS, or MC-ICP-MS without a collision/reaction cell, or STD mode) | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 (Q-ICP-MS and MC-ICP-MS only) | 2026-07-29 | Y | N | Y
Cell Exit Discrimination Voltage | Kinetic energy discrimination offset voltage applied at the exit of the collision cell in KED mode. Applies to quadrupole-based instruments and to collision-cell-equipped multi-collector instruments (e.g., Nu Sapphire, Thermo Neoma MS/MS) operating in KED mode. | Advanced | Editable | Numeric (V) | −5 V ¦ Not applicable (SF-ICP-MS, or MC-ICP-MS without a collision/reaction cell, or STD mode) | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 (Q-ICP-MS and MC-ICP-MS only) | 2026-07-29 | Y | N | Y
Reaction Gas Type | Type of reactive gas introduced into the dynamic reaction cell (DRC) for interference removal through ion-molecule reactions. Applies to quadrupole-based instruments and to collision-cell-equipped multi-collector instruments (e.g., Nu Sapphire, Thermo Neoma MS/MS) operating in DRC mode. | Advanced | Read-Only | Controlled list | NH3 ¦ O2 ¦ Not applicable (SF-ICP-MS, or MC-ICP-MS without a collision/reaction cell, or DRC mode not used) ¦ Other: specify | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 (Q-ICP-MS and MC-ICP-MS only) | 2026-07-29 | Y | N | Y
Reaction Gas Flow Rate | Flow rate of the reactive gas introduced into the dynamic reaction cell (DRC). Applies to quadrupole-based instruments and to collision-cell-equipped multi-collector instruments (e.g., Nu Sapphire, Thermo Neoma MS/MS) operating in DRC mode. | Advanced | Editable | Numeric (mL/min) | 0.4 mL/min ¦ Not applicable (SF-ICP-MS, or MC-ICP-MS without a collision/reaction cell, or DRC mode not used) | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 (Q-ICP-MS and MC-ICP-MS only) | 2026-07-29 | Y | N | Y
Mass Resolution per Analyte | Mass resolution mode assigned to each analyte in this procedure. The selected resolution determines which polyatomic interferences (e.g., ¹⁸⁶W¹⁶O⁺ on ²⁰²Hg, relevant to common-Pb correction) are physically resolved by the magnetic sector. Applicable only to sector-field instruments. | Advanced | Read-Only | Text (free) | All LR (m/Δm ≈ 300) ¦ LR: Pb, U, Th; MR: Hg-interference-affected masses ¦ N/A (not sector-field) | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 (SF-ICP-MS only; Analyte-Specific) | 2026-07-28 | N | Y | N
E-scan Range | Electric scan range used for peak acquisition, expressed as percentage of the centre mass. Varies the accelerating voltage to cover nearby masses without re-scanning the magnet. Applicable only to sector-field instruments. | Advanced | Read-Only | Numeric (%) | 15% ¦ 10% ¦ N/A (not sector-field) | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 (SF-ICP-MS only) | 2026-07-28 | N | Y | N
Triple Scanning Mode | Whether each mass peak is scanned three times per cycle and the results averaged, to reduce noise from short-term magnetic field instabilities. Applicable only to sector-field instruments with this capability. | Advanced | Read-Only | Controlled list | Y ¦ N ¦ N/A (not sector-field) | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 (SF-ICP-MS only) | 2026-07-28 | N | Y | N
Other Information | Supplementary methodological context not captured by the structured fields above, as reported in Table 3 of Horstwood et al. (2016) — in the source paper's own example, notes on ablation-cell and sample-line configuration for depth-profiling/single-pulse ablation studies, a variant not otherwise in scope for this single-mode (Spot Analysis) TAPP. | Advanced | Editable | Text (free) | Depth profiling and single-pulse ablation studies used a low-volume ablation cell, sample line of 3 m from ablation cell to torch and a 20 s wait time between ablations. Cell washout time was ca. 0.8 s to 1% of peak signal. | Source: Table 3 (Horstwood et al. 2016) | 2026-07-28 | Y | Y | Y
5. Data Processing | [7 empty] | N | N | N
Gas Blank | Duration of the background (gas blank) measurement and how it is applied to correct the ablation signal. | Basic | Read-Only | Text (free) | 30 s on-peak zero subtracted | Source: Table 3 (Horstwood et al. 2016) | 2026-07-28 | Y | Y | Y
Elemental Fractionation Correction | Method used to correct for laser-induced elemental fractionation (LIEF), including any assumption about reference-material/sample behaviour equivalence, as reported in Table 3 of Horstwood et al. (2016). | Basic | Read-Only | Text (free) | LIEF correction assumes reference material and samples behave identically. | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; field renamed from 'Data Processing Package Used/Correction for LIEF' for cross-TAPP consistency; software-name content moved to Group 3's 'Data Reduction Software'. | 2026-07-28 | Y | Y | Y
Mass Bias Correction Strategy | Primary strategy used to correct for instrumental isotopic mass fractionation (mass bias/mass discrimination). | Basic | Read-Only | Controlled list | Internal normalization | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; split out of the former 'Mass Discrimination' field, matching the field structure of Solution MC-ICP-MS TAPP v3. Table 3's Tl-U tracer approach corresponds to internal normalization. | 2026-07-28 | Y | Y | Y
Internal Normalization Element and Isotope Ratio | Element and isotope ratio used as the internal normalization standard for mass bias correction, and how the resulting correction is applied to analyte ratios. | Basic | Read-Only | Text (free) | Tl-U tracer solution used for mass bias correction, with 207Pb/206Pb and 206Pb/238U additionally normalised to reference material | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; split out of the former 'Mass Discrimination' field, matching Solution MC-ICP-MS TAPP v3's 'Internal Normalization Element and Isotope Ratio' field. | 2026-07-28 | N | N | Y
Mass Fractionation Law | Mathematical law used to parameterize instrumental isotope mass fractionation as a function of mass difference. Not stated in Table 3 (Horstwood et al. 2016) for the worked example; the exponential law is most widely used and theoretically best justified for MC-ICP-MS. | Advanced | Read-Only | Controlled list | Exponential ¦ Linear ¦ Power ¦ Not stated | Source: In-house TAPP rule — imported from Solution MC-ICP-MS TAPP v3 (primarily relevant to MC-ICP-MS mass bias correction) | 2026-07-28 | N | N | Y
Initial/Common-Daughter Correction, Composition and Uncertainty | Whether a correction was applied for a daughter-isotope component present in the sample at or after formation that is not attributable to in-situ radioactive decay of the parent since that time, and if so, the assumed or measured composition and its uncertainty used in the correction. This concept is known by different names across geochronometer systems: common-Pb correction (U-Pb, typically using a model composition such as Stacey and Kramers 1975, or a measured non-radiogenic proxy isotope such as ²⁰⁴Pb); initial-Sr or initial ⁸⁷Sr/⁸⁶Sr correction (Rb-Sr); initial-Nd or εNd correction relative to a reference reservoir such as CHUR (Sm-Nd); initial-Hf or εHf correction relative to CHUR (Lu-Hf); detrital-Th or initial ²³⁰Th correction, typically using ²³²Th as a non-radiogenic proxy isotope (U-series / Th-U disequilibrium dating); common-Os correction (Re-Os). Record 'Not applicable' if the daughter-isotope component from formation is negligible relative to radiogenic ingrowth for this procedure, or if it is solved for internally via isochron regression rather than applied as an explicit standalone correction (see Age/Model-Age Calculation Method). | Basic | Read-Only | Text (free) | No common-Pb correction applied to the data. ¦ Common-Pb corrected using Stacey and Kramers (1975) model composition at estimated age. ¦ Initial ⁸⁷Sr/⁸⁶Sr = 0.7045, assumed from a coexisting low-Rb/Sr mineral. ¦ εNd(t) calculated relative to CHUR (Jacobsen and Wasserburg 1980). ¦ Detrital-Th corrected using an assumed initial ²³⁰Th/²³²Th activity ratio of 0.8 ± 0.4 (bulk-Earth estimate). ¦ Not applicable — solved internally via isochron regression. | Source: Generalized 2026-07-30 from Horstwood TAPP v7's U-Pb-specific 'Common-Pb Correction, Composition and Uncertainty' field, to cover the structurally analogous correction required by other parent-daughter geochronometer systems (Rb-Sr, Sm-Nd, Lu-Hf, U-series, Re-Os). | 2026-07-30 | Y | Y | Y
Age/Model-Age Calculation Method | The equation or statistical approach used to convert corrected isotope ratio(s) into the final reported age. Common approaches for U-(Th-)Pb geochronology include: a single-ratio age from one measured ratio (e.g. ²⁰⁶Pb/²³⁸U or ²⁰⁷Pb/²⁰⁶Pb); a weighted mean age across multiple analyses of the same population; a Concordia (Wetherill) upper- or lower-intercept age; a Tera-Wasserburg lower-intercept age, with the common-Pb composition either fixed (anchored) or solved for as part of the regression. Distinct from Common-Pb Correction, Composition and Uncertainty, which addresses the isotope-ratio correction step that typically precedes this calculation, and from Uncertainty Level and Propagation, which addresses how component uncertainties are combined into the reported age uncertainty. | Advanced | Read-Only | Controlled list | Single-ratio age (²⁰⁶Pb/²³⁸U) ¦ Single-ratio age (²⁰⁷Pb/²⁰⁶Pb) ¦ Weighted mean age ¦ Concordia (Wetherill) upper intercept ¦ Concordia (Wetherill) lower intercept ¦ Tera-Wasserburg lower intercept (anchored common-Pb) ¦ Tera-Wasserburg lower intercept (free regression) ¦ Other: specify | Source: In-house TAPP rule — new field, no Table 3 precedent | 2026-07-30 | Y | Y | Y
Constants and Reference Values Used | Physical constants and reference values used in data reduction to calculate the final reported quantity (e.g., decay constants for age calculation, the 238U/235U ratio used to convert measured 235U-based data to 207Pb/235U, or other citable reference values), together with their source. Distinct from Reference Material Information (Group 6), which documents accepted values for specific calibration/validation materials rather than universal physical constants. Critical for traceability: community-recommended values for decay constants and isotope ratios are periodically revised (e.g., the 238U/235U ratio revision by Hiess et al. 2012), and a reported age or ratio can only be correctly reinterpreted against a future revision if the constants originally used are known. | Basic | Editable | Text (free) | λ238U = 1.55125×10⁻¹⁰ yr⁻¹, λ235U = 9.8485×10⁻¹⁰ yr⁻¹, λ232Th = 4.9475×10⁻¹¹ yr⁻¹ (Jaffey et al. 1971); 238U/235U = 137.818 (Hiess et al. 2012) | Source: In-house TAPP rule — new field motivated by Table 4 of Horstwood et al. (2016), whose footnote states 'Decay constants of Jaffey et al. (1971) used' — information Table 3's own metadata template never captures as a structured item. Generalizes beyond geochronology to any procedure whose final reported quantity depends on citable, potentially-revisable physical constants. | 2026-07-29 | Y | Y | Y
Uncertainty Level and Propagation | Uncertainty level at which ages/ratios are quoted (e.g., 1σ, 2σ) and the method used to propagate component uncertainties (e.g., quadratic addition) into the final reported value. | Basic | Editable | Text (free) | Ages are quoted at 2s absolute, propagation is by quadratic addition. Reproducibility and age uncertainty of reference material and common-Pb composition uncertainty are propagated where appropriate. | Source: Table 3 (Horstwood et al. 2016) | 2026-07-30 | Y | Y | Y
Per-Analyte Calibration Strategy | Documents cases where different analytes are calibrated using different strategies within the same session. If a single strategy applies uniformly, record that here. | Advanced | Read-Only | Controlled list | Single strategy: matrix-matched primary reference material for all analytes ¦ N/A ¦ Other: specify | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Matrix Offset Correction (LIEF) | Whether an empirical correction was applied to account for systematic differences in laser-induced elemental fractionation (LIEF) patterns between the external calibration standard and the sample matrix. | Advanced | Read-Only | Text (free) | e.g., 'Not applicable (matrix-matched mineral standard used)' ¦ 'Applied: specify' | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Signal Integration Interval Method | Approach used to select the stable signal interval for integration during data reduction, and to exclude anomalous signals from inclusions, cracks, or mixed domains. Central to interpreting excess scatter/MSWD in the resulting age population. | Advanced | Read-Only | Text (free) | Manual visual inspection per analysis ¦ Automated threshold (specify software and threshold) ¦ Fixed time window applied uniformly | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Signal Integration Time | Actual integration time used for the ablation signal in this analysis. Determined at analysis time by applying the Signal Integration Interval Method to the time-resolved signal. | N/A | Basic | Numeric (s) | e.g., 40 s ¦ 35 s | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | N
Spike / Outlier Filtering Approach | Method used to identify and remove anomalous signal spikes arising from micronuggets, inclusions, cracks, or instrument artifacts during time-resolved signal processing. | Advanced | Editable | Text (free) | e.g., 'Significant single spikes visually identified and removed' ¦ 'No filtering applied' | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Isobaric Interference Corrections Applied | Whether isobaric interference corrections were applied for any measured isotope in this procedure. A procedure-level Boolean; element-specific detail is in Interfering Species and Interference Correction Method. | Advanced | Read-Only | Boolean | Yes ¦ No | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Interfering Species | Elemental or molecular species overlapping with the measured isotope, relevant to U-Th-Pb geochronology (e.g., ²⁰⁴Hg on ²⁰⁴Pb, relevant to the Common-Pb Correction field above). | Advanced | Read-Only | Text (free) | e.g., '²⁰⁴Hg on ²⁰⁴Pb' ¦ 'None' | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Interference Correction Method | Equation or procedure used to correct for isobaric interferences, including the reference material or monitor isotope used to characterise it. | Advanced | Read-Only | Text (free) | e.g., '²⁰²Hg monitored to correct ²⁰⁴Hg on ²⁰⁴Pb via natural ²⁰⁴Hg/²⁰²Hg' ¦ 'None' | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Isotope Dilution Data Reduction Method | Mass balance approach used to calculate sample mass fractions from spike-sample isotope ratio measurements, relevant when a concentration spike (e.g., ²³⁵U, as in Table 3's own worked example) is co-aspirated. Record 'None' if isotope dilution is not used. | Advanced | Read-Only | Text (free) | e.g., 'IUPAC isotope dilution equations' ¦ 'None' | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Double-Spike Isotope Pair | The pair of enriched isotopes used to construct a double spike for combined mass bias and/or common-Pb correction. Both isotopes must be naturally rare and free from isobaric interferences. Record 'N/A' if double-spike method is not used. | Advanced | Read-Only | Text (free) | e.g., ²⁰⁴Pb–²⁰⁷Pb ¦ N/A | Source: In-house TAPP rule — imported from Solution MC-ICP-MS TAPP v3 (MC-ICP-MS only, double-spike procedures) | 2026-07-28 | N | N | Y
Double-Spike Mixing Ratio | Target proportion of double-spike signal relative to total analyte signal in the spiked mixture. An optimal mixing ratio minimises error propagation through the double-spike inversion. Record 'N/A' if double-spike method is not used. | Advanced | Editable | Text (free) | e.g., ~0.4 (spike fraction) ¦ N/A | Source: In-house TAPP rule — imported from Solution MC-ICP-MS TAPP v3 (MC-ICP-MS only, double-spike procedures) | 2026-07-28 | N | N | Y
Double-Spike Inversion Algorithm | Mathematical algorithm used to simultaneously solve for the true sample isotope ratio and the mass fractionation factor from the measured spike-sample mixture. Record 'N/A' if double-spike method is not used. | Advanced | Read-Only | Text (free) | e.g., Rudge et al. (2009) Double Spike Toolbox ¦ N/A | Source: In-house TAPP rule — imported from Solution MC-ICP-MS TAPP v3 (MC-ICP-MS only, double-spike procedures) | 2026-07-28 | N | N | Y
Memory Effect Mitigation | Procedure applied to identify and minimise memory effects from high-concentration elements in the previous sample or standard that may contaminate subsequent analyses. | Advanced | Editable | Text (free) | e.g., 'Extended washout between high-U standards and unknowns' ¦ 'None required' | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Normalization / Standards-Based Correction | Any post-acquisition normalization applied to correct for systematic biases identified from secondary reference materials. | Advanced | Editable | Text (free) | e.g., 'None' ¦ 'Applied: specify' | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
6. Quality Control & Uncertainty | [7 empty] | N | N | N
Primary Calibration Standard Name | Primary reference material used to calibrate the instrument and convert raw signal intensities to ages/isotope ratios. Include citation for the accepted values used. | Basic | Editable | Text (free) | 91500 (Wiedenbeck et al. 1995) | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; split out of the former 'Calibration Strategy' + 'Reference Material Information' fields, re-partitioned along LA-Q/SF-ICP-MS TAPP's primary/secondary axis. | 2026-07-28 | Y | Y | Y
Secondary Reference Materials | Quality-control reference material(s) analysed as unknowns alongside samples in the same session to assess accuracy and monitor drift. Include citation/accepted values for each. | Basic | Editable | Text (free) | Plesovice (Slama et al. 2008); GJ1, 206Pb/238U 0.097877 ± 0.07%, 207Pb/206Pb 0.060171 ± 0.08% (in-house CA-TIMS, this study) | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; split out of the former 'Calibration Strategy' + 'Reference Material Information' fields, re-partitioned along LA-Q/SF-ICP-MS TAPP's primary/secondary axis. | 2026-07-28 | Y | Y | Y
Within-Session Analytical Precision and Assessment Method | Reproducibility of repeated measurements within a single analytical session. Assessment method specifies the reference material used, number of replicates n, and the statistic reported. | Basic | Basic | Text (free) | Plesovice — Wtd ave 206Pb/238U age = 338 ± 3 (2s, MSWD = 0.9, n = 8); GJ-1 — Wtd ave 206Pb/238U age = 602 ± 5 (2s, MSWD = 1.1, n = 7) | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; split out of the former 'Quality Control/Validation' field, which reported within-session validation-material weighted-average ages, MSWD and n; this destination inherits that content directly. | 2026-07-28 | Y | Y | Y
Between-Session (Long-Term) Analytical Precision and Assessment Method | Reproducibility of measurements across multiple analytical sessions over weeks to months. Not addressed in Table 3 of Horstwood et al. (2016), which reports only within-session validation results. | Advanced | Basic | Text (free) | e.g., '2σ RSD = X% for 206Pb/238U age on [RM] across N sessions over M months' | Source: In-house TAPP rule — new field with no Table 3 precedent, added alongside the split of 'Quality Control/Validation' to complete the precision/accuracy structure used in LA-Q/SF-ICP-MS TAPP and Solution MC-ICP-MS TAPP; Advanced (not Basic) because no paper content exists to inherit, unlike its two sibling fields from the same split | 2026-07-28 | Y | Y | Y
Analytical Accuracy and Assessment Method | Offset between measured and accepted reference values for secondary reference materials. Table 3's validation weighted-average ages, compared against each material's independently accepted age (cited in Secondary Reference Materials above), serve as the accuracy statement for this worked example. | Basic | Basic | Text (free) | Plesovice measured weighted-average age (338 ± 3 Ma) compared against accepted age (337.13 ± 0.37 Ma, Slama et al. 2008); GJ-1 measured weighted-average age (602 ± 5 Ma) compared against accepted age from Reference Material Information | Source: Table 3 (Horstwood et al. 2016) — restructured/split for cross-TAPP consistency with LA-Q/SF-ICP-MS TAPP; split out of the former 'Quality Control/Validation' field; reframed as a bias-vs-accepted-value statement using the accepted ages already cited in Secondary Reference Materials, since Table 3's compact example reports precision and accuracy together in one value. | 2026-07-28 | Y | Y | Y
Calibration Standard Measurement Frequency | How often the primary calibration standard is measured relative to unknown samples within a session, i.e. the bracketing interval used to monitor and correct for instrumental drift. Implied but not stated explicitly in Table 3 (Horstwood et al. 2016) via the validation-material n counts. | Advanced | Read-Only | Text (free) | e.g., 'Before and after every 10 unknowns' ¦ 'Start and end of session only' | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y
Detection Limit | Session detection limit for U, Th, and/or Pb, demonstrating the reliability of reported concentrations near the detection limit. Not requested by Table 3 of Horstwood et al. (2016), which focuses on age uncertainty rather than elemental detection limits. | Advanced | Basic | Numeric (unit) | e.g., '0.05 ppm U (3σ background)' ¦ 'Range: 0.02–0.4 ppm across U, Th, Pb' | Source: In-house TAPP rule — Rule 1 shared Group 6 QC term (gap vs. Table 3) | 2026-07-28 | Y | Y | Y
Detection Limit Method | Reference or description of the method used to calculate the session detection limit reported above. | Advanced | Basic | Text / URI | e.g., '3σ of background ion counts per Longerich et al. (1996)' | Source: In-house TAPP rule — Rule 1 shared Group 6 QC term (gap vs. Table 3) | 2026-07-28 | Y | Y | Y
Limit of Quantification (LOQ) Method | Reference or description of the method used to calculate the limit of quantification: the lowest concentration reliably measurable with acceptable precision and accuracy. | Advanced | Basic | Controlled list | 10σ blank ¦ 10× DL ¦ S/N ≥ 10 ¦ Other: specify | Source: In-house TAPP rule — imported from LA-Q/SF-ICP-MS TAPP v4 | 2026-07-28 | Y | Y | Y

## Legends
Procedure-Level Tier | Description | Analysis-Level Tier | Description
Basic | Mandatory for procedure registration. Must be provided to register a valid procedure. | Read-Only | Directly imported from the registered procedure; cannot be changed by the analyst. Changing this value means running a different procedure. Also used for fields relevant only at procedure level — the value is inherited from the procedure record and shown read-only in the analysis form.
Advanced | Optional for procedure registration. Strongly recommended but not required. | Editable | Imported from the registered procedure but may be adjusted within procedure-defined bounds (e.g., daily tuning, minor software updates). The procedure registers the target or typical value; the analyst confirms or adjusts it. Cannot be left void if the procedure-level tier is Basic.
N/A | Not applicable at procedure level. This field captures analysis-level information only. | Basic | Mandatory user input at analysis time. Value comes from the analysis itself and cannot be pre-specified in the procedure.
|  | Advanced | Optional user input at analysis time. Recommended for complete and reproducible documentation.
Mode Column | Definition
Q-ICP-MS
SF-ICP-MS
MC-ICP-MS
Y | This field applies to this analytical mode and should be reported.
N | This field does not apply to this analytical mode.