## Field Comparison
Group | Horstwood TAPP Item | LA-Q/SF-ICP-MS TAPP Item(s) | Relationship | Assessment Category | Notes | Cosmote's comment
1. Procedure Identification | Procedure Name | Procedure Name | 1:1 | Templated (not empirical) | See subtotal note below.
| Technique | Technique | 1:1 | Templated (not empirical) | See subtotal note below.
| Procedure Author | Procedure Author | 1:1 | Templated (not empirical) | See subtotal note below.
| Analyst | Analyst | 1:1 | Templated (not empirical) | See subtotal note below.
| Laboratory | Laboratory | 1:1 | Templated (not empirical) | See subtotal note below.
| Laboratory ID | Laboratory ID | 1:1 | Templated (not empirical) | See subtotal note below.
| Procedure Start Date | Procedure Start Date | 1:1 | Templated (not empirical) | See subtotal note below.
| Analysis Start Date | Analysis Start Date | 1:1 | Templated (not empirical) | See subtotal note below.
| Analysis End Date | Analysis End Date | 1:1 | Templated (not empirical) | See subtotal note below.
| Procedure DOI | Procedure DOI | 1:1 | Templated (not empirical) | See subtotal note below.
| Funding Source for Procedure Development | Funding Source for Procedure Development | 1:1 | Templated (not empirical) | See subtotal note below.
| Funding Source for Analysis | Funding Source for Analysis | 1:1 | Templated (not empirical) | See subtotal note below.
| Procedure Reference(s) | Procedure Reference(s) | 1:1 | Templated (not empirical) | See subtotal note below.
| Coupled Technique(s) | Coupled Technique(s) | 1:1 | Templated (not empirical) | See subtotal note below.
| Coupling Description | Coupling Description | 1:1 | Templated (not empirical) | See subtotal note below.
| Coupled Procedure DOI | Coupled Procedure DOI | 1:1 | Templated (not empirical) | See subtotal note below.
| Coupled Dataset or Publication Reference | Coupled Dataset or Publication Reference | 1:1 | Templated (not empirical) | See subtotal note below.
|  |  |  | SUBTOTAL | Group 1: 17/17 fields align 1:1 (100%), identical to the v13 comparison (Group 1 is unchanged between v12 and v4). Not a meaningful data point for domain-difference analysis since both TAPPs are required to build Group 1 from the identical template (Rule 1) -- this measures template compliance, not independent convergence.
2. Samples | Target Material | Target Material | 1:1 | Domain-appropriate divergence | Same field/tier (Basic/Read-Only in both). Controlled vocab differs appropriately: Horstwood lists U-Pb-datable accessory minerals (zircon, monazite, titanite...); LAQSF lists broad rock-forming/matrix types (silicate glass, feldspar, olivine, sulfide...) reflecting its much wider target-material scope.
| Sample Preparation Method | Sample Preparation Method | 1:1 (tier differs) | Tier discrepancy noted | Same field (Rule 1 standard name). Horstwood: Advanced/Editable. LAQSF v4: Basic/Editable -- LAQSF promoted this to Basic at some point after the v12/v2 sync point (it was Advanced in the frozen v13 snapshot). Worth revisiting Horstwood's tier to match, since sample prep is arguably just as mandatory-to-report for geochronology. | Set to Basic/Editable
| Sample Name | Sample Name | 1:1 | Templated | Same field, same tier.
| Sample Persistent Identifier | Sample Persistent Identifier | 1:1 | Templated | Same field, same tier.
| Imaging | -- (no equivalent) | Unique to Horstwood | Possible universal gap in LAQSF | CL/BSE pre-ablation characterisation imaging has no LAQSF counterpart at all. Zoning/domain imaging is central to zircon geochronology (avoiding mixed cores/rims), but BSE imaging to avoid inclusions/cracks is also standard practice in trace-element work -- this looks like a field that would benefit LAQSF too, not a purely geochronology-specific need. | Is this information captured by "Coupled Technique" in Group 1?
| -- (folded into Sample Preparation Method prose) | Sample Form / Analytical Substrate | Unique to LAQSF | Domain-appropriate divergence | LAQSF's controlled list (fused glass disc, pressed pellet, liquid, etc.) reflects whole-rock/homogenised-matrix analysis modes essentially never used in geochronology. Table 3's prose-only treatment is adequate for geochronology's narrower, always-in-situ sample form. | Is matrix geochoronology always done with solution-based ICP-MS?
| -- (no equivalent) | Fusion Flux and Dilution Ratio | Unique to LAQSF | Clean domain-specific (correctly absent) | Fusion is fundamentally incompatible with geochronology (destroys textural context). Correctly present in LAQSF, correctly absent from Horstwood. | Keep absent from Horstwood TAPP. I guess this one is related to the question above?
| -- (no equivalent) | Pre-Ablation Surface Treatment | Unique to LAQSF | Likely genuine gap in Table 3 | Pre-ablation cleaning pulses matter equally for geochronology. Table 3's silence here looks like an omission rather than a deliberate domain-appropriate exclusion. | Add to Horstwood TAPP, and keep a note of source.
|  |  |  | SUBTOTAL | Group 2: 4/5 Horstwood fields matched (80%); 3/7 LAQSF fields unmatched (unchanged from v13 -- Group 2 did not gain fields in the v2->v4 development). One tier discrepancy surfaced (Sample Preparation Method).
3. Instrument & Software | Make, Model and Type (Laser Ablation System) | Laser Manufacturer & Model + Laser Wavelength and Type | many:many (different split axis) | Different partition, same information | Horstwood splits [manufacturer+model+type] (Group 3) from [wavelength] (Group 4). LAQSF splits [manufacturer+model] (Group 3) from [wavelength+type] (Group 3). Both capture the same total information, partitioned differently. | For laser and ICP-MS "Make, Model and Type", keep consistent with our LAQSF TAPP for cross-TAPP consistency.
| Ablation Cell and Volume | Ablation Cell Type | 1:1 (Horstwood more specific) | Horstwood more rigorous | Horstwood explicitly requests cell volume (cm3), which LAQSF's field does not prompt for even though volume is a primary determinant of washout time / achievable time resolution. | Adopt and update LAQSF TAPP.
| Make, Model and Type (ICP-MS Instrument) | ICP-MS Manufacturer & Model + ICP-MS Type | 1:2 (LAQSF splits) | LAQSF more rigorous | Horstwood bundles manufacturer/model/mass-analyzer-type into one free-text field. LAQSF splits type into its own controlled-list field, enabling filtering (e.g. 'all SF-ICP-MS procedures'). LAQSF's approach looks like the better convention to adopt. | See laser system "Make, Model and Type" above.
| Sample Introduction | -- (no equivalent) | Unique to Horstwood | Clean domain-specific (correctly absent) | Co-aspirated tracer/spike solutions (Tl for mass bias, U-spike for isotope dilution) are essentially unique to isotope-ratio work; trace-element quantification uses external standard bracketing instead. | Why shouldn't Sample Introduction be added to LAQSF TAPP? Can you elaborate?
| Detection System | Detector Configuration | 1:1 | Comparable depth | Same concept, similar level of detail in both. | Use "Detection System" for cross-TAPP consistency.
| Acquisition Software | Acquisition Software | 1:1 | Templated | Same field (Rule 1 standard name).
| Data Reduction Software | Data Reduction Software | 1:1 (Horstwood currently empty) | Open item in Horstwood draft | No Table 3 content was mapped here -- the paper's software mention was bundled into 'Data Processing Package Used/Correction for LIEF' (Group 5) instead. | Add to Horstwood TAPP. This should also be "Templated" from LAQSF TAPP?
| -- (no equivalent) | Instrument Serial Number or Lab Identifier | Unique to LAQSF | New in v4; minimum-reporting framing, not domain exclusion | Traceability field added after the v2/v12 sync point. Not domain-specific -- serial-number/lab-ID traceability would benefit geochronology procedures identically. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Interface Cone Configuration | Unique to LAQSF | Minimum-reporting framing, not domain exclusion | Cone configuration affects sensitivity/interference for all ICP-MS work, but is rarely reported in geochronology papers even today. Likely reflects Table 3's deliberately low adoption bar rather than genuine irrelevance. | Add to Horstwood TAPP, and keep a note of source. If it is rarely reported, should it be Advanced/Editable?
| -- (no equivalent) | Sampler and Skimmer Cone Material | Unique to LAQSF | Minimum-reporting framing, not domain exclusion | Same reasoning as Interface Cone Configuration. | Add to Horstwood TAPP, and keep a note of source. If it is rarely reported, should it be Advanced/Editable?
| -- (no equivalent) | Torch Depth | Unique to LAQSF | Minimum-reporting framing, not domain exclusion | Same reasoning as Interface Cone Configuration. | Add to Horstwood TAPP, and keep a note of source. If it is rarely reported, should it be Advanced/Editable?
| -- (no equivalent) | Torch Type | Unique to LAQSF | New in v4; minimum-reporting framing, not domain exclusion | Injector geometry -- same reasoning as the other fine ICP-tuning fields. | Add to Horstwood TAPP, and keep a note of source. If it is rarely reported, should it be Advanced/Editable?
| -- (no equivalent) | Guard Electrode | Unique to LAQSF | Minimum-reporting framing, not domain exclusion | Same reasoning as Interface Cone Configuration. | Add to Horstwood TAPP, and keep a note of source. If it is rarely reported, should it be Advanced/Editable?
| -- (no equivalent) | Mass Resolution Setting | Unique to LAQSF | Likely genuine gap, esp. for SF-ICP-MS -- see Group 4 SF cluster | Not a minimum-reporting artifact -- resolution setting materially affects which interferences are resolved. Table 3's own example instrument (Nu Plasma HR, MC-ICP-MS) is not resolution-selectable the way single-collector SF-ICP-MS is. This field is the Group-3 anchor of a larger SF-specific cluster now confirmed in v4's Group 4 (Mass Resolution per Analyte, E-scan Range, Triple Scanning Mode) -- see below. Directly echoes the planning table's row 7d note calling out 'E-scan mode at fixed magnet, SEM triple mode' as central to SF-ICP-MS geochronology specifically. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Collision/Reaction Cell (CRC) Configuration | Unique to LAQSF | Clean domain-appropriate exclusion (new in v4) | Explicitly tagged 'Q-ICP-MS only' in LAQSF's own Comments column. The planning table's row 7d notes state directly: 'Q-ICP-MS not standard for this [geochronology] application due to sensitivity constraints and Hg/Pb isobar complexity at unit resolution.' Correctly absent from Horstwood -- confirms the planning table's own architecture decision rather than exposing a gap. | Per our discussion, it seems we should add it to Horstwood TAPP because some detrital zircons are measured with LA-Q-ICP-MS.
|  |  |  | SUBTOTAL | Group 3: 6/7 Horstwood fields matched (86%); LAQSF grew from 12 to 16 spot-applicable fields between v13 and v4 (+1 Instrument Serial Number, +1 Torch Type, +1 CRC Configuration). Of 10 LAQSF-unique fields, 1 (CRC Configuration) is a clean, planning-table-confirmed domain exclusion; the rest split between minimum-reporting framing and the Mass Resolution Setting gap, which is substantially reinforced by the new SF-specific cluster in Group 4.
4. Measurement Information | Analytical Mode | Analytical Mode | 1:1 | Templated | Same field (Rule 3 mandatory field), same tier.
| Spot Diameter (um), Nominal/Actual | Laser Spot Geometry | 1:1 (Horstwood more rigorous) | Horstwood more rigorous | LAQSF's field is a procedure-level target only. Horstwood's bundles the nominal (target) value with an independently measured/verified actual value -- closer to the procedure-target-vs-measured-value pattern used elsewhere in the TAPP library (e.g. Oxide Production). | This is a good catch. Nominal spot diameter can be different from the actual diameter. I think we should add a separate field to capture this information. Something like "Actual spot diameter", and is N/A on procedure level, Basic on analysis level.
| (merged into Analytical Mode) -- was 'Sampling Mode/Pattern' | Laser Spot Path / Ablation Mode | Partial / structurally different | Note, not a gap | LAQSF keeps a second, more granular mode-pattern field separate from its top-level Analytical Mode. Horstwood's single-mode scope made this distinction moot post-merge. | Keep consistent with LAQSF TAPP for cross-TAPP consistency.
| Laser Wavelength (nm) | Laser Wavelength and Type (Group 3) | 1:1 (cross-group) | Placement note | Same information; LAQSF places wavelength in Group 3 (fixed hardware fact), Horstwood in Group 4 (operating parameter). LAQSF's placement is arguably the better fit for TAPP's own Group 3-vs-4 definitions. | Keep consistent with LAQSF TAPP for cross-TAPP consistency.
| Pulse Width (ns) | Laser Pulse Duration | 1:1 | Tier agreement | Both Advanced/Read-Only. | Keep consistent with LAQSF TAPP for cross-TAPP consistency.
| Fluence (J cm-2) | Laser Fluence (Energy Density) | 1:1 | Tier agreement | Both Basic/Editable. | Keep consistent with LAQSF TAPP for cross-TAPP consistency.
| Repetition Rate (Hz) | Laser Repetition Rate | 1:1 | Tier agreement | Both Basic/Editable. | Keep consistent with LAQSF TAPP for cross-TAPP consistency.
| Ablation Duration (s) | Ablation Duration per Spot | 1:1 | Tier agreement | Both Basic/Editable. | Keep consistent with LAQSF TAPP for cross-TAPP consistency.
| Ablation Pit Depth/Ablation Rate | -- (no equivalent) | Unique to Horstwood | Domain-emphasis, not domain-exclusive | Matters for both domains but is especially emphasised in geochronology (downhole fractionation correction). LAQSF's 'Elemental Fractionation Correction' (Group 5) references the depth-dependence indirectly but never captures the underlying pit-depth/rate metric. | Adopt and add to LAQSF TAPP.
| -- (no equivalent) | Laser Beam Energy Profile | Unique to LAQSF | Likely became more relevant since 2016 | Flat-top vs. Gaussian beam shaping became a much more common commercial differentiator from roughly 2018 onward -- plausibly postdates the paper's minimum-reporting consensus. | Add to Horstwood TAPP, and keep a note of source.
| Carrier Gas | Carrier Gas and Flow Rate | 2:1 (Horstwood splits) | Different partition, same information | Horstwood keeps gas identity separate from flow rate (matching Table 3's own two-row structure); LAQSF combines them. | Adopt and modify LAQSF TAPP.
| Cell Carrier Gas Flow (L min-1) | Carrier Gas and Flow Rate | 2:1 (Horstwood splits) | Different partition, same information | See above. | Adopt and modify LAQSF TAPP.
| RF Power (W) | RF Power | 1:1 (tier now differs) | Tier discrepancy noted | Horstwood: Basic/Read-Only. LAQSF v4: Basic/EDITABLE -- changed from Read-Only (the v13 snapshot still had Read-Only). Notably, LAQSF's own field Description text still reads 'Fixed in the procedure; cannot be changed by analysts without defining a new procedure' -- which argues for Read-Only, not Editable. This looks like an internal inconsistency in LAQSF v4 (description vs. tier out of sync) rather than a deliberate, reasoned change worth mirroring in Horstwood. Worth flagging back to LAQSF for a fix independent of this comparison. | I think this should be Basic/Read-Only in both.
| Make-up Gas Flow (L min-1) | Plasma / Make-up Gas Addition | 1:1 | Comparable | Same concept; Horstwood's example is specific to this lab's desolvating-nebuliser/tracer setup.
| -- (no equivalent) | Coolant (Plasma) Gas Flow Rate | Unique to LAQSF | Minimum-reporting framing, not domain exclusion | New in v4 as its own field (split out of the former combined 'Auxiliary and Cool Gas Flow Rates'). Standard ICP operating parameter relevant regardless of application; likely omitted from Table 3 as a rarely-varied factory default. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Auxiliary Gas Flow Rate | Unique to LAQSF | Minimum-reporting framing, not domain exclusion | Companion split-out field; same reasoning. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Plasma Thermal Mode | Unique to LAQSF | Domain-appropriate (answer would be near-constant) | Geochronology overwhelmingly uses normal/hot plasma for U/Th/Pb sensitivity. Still worth having for completeness even if the answer rarely varies. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Signal Smoothing | Unique to LAQSF | Likely became more relevant since 2016 | Wider adoption for fast-washout-cell precision improvement since roughly 2018-2020; may have been less standardised in 2016. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | ICP Tuning (general) | Unique to LAQSF | Likely genuine gap | General tuning-approach documentation is not domain-specific and would benefit geochronology procedures too. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Oxide Production Method and Threshold | Unique to LAQSF | Likely genuine gap -- important | ThO+/Th+ tuning is standard practice in modern U-Pb LA-ICP-MS labs and directly affects Th/U ratio accuracy -- central to U-Th-Pb dating, not incidental to it. Complete absence from Table 3 looks like a genuine omission. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Oxide Production | Unique to LAQSF | Likely genuine gap -- important | Analysis-level companion to the above. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Doubly-Charged Species Monitor | Unique to LAQSF | Likely genuine gap | Generic ICP-MS QC concern, relevant to any mass region affected, including U-Pb-adjacent masses. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Doubly-Charged Species Production | Unique to LAQSF | Likely genuine gap | Analysis-level companion to the above. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Instrument Warm-up / Session Duration Limit | Unique to LAQSF | Minor gap, generic good practice | Not domain-specific. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Analysis Sequence | Unique to LAQSF | Moderate-to-important gap | Bracketing/sequence design is absent from Table 3 entirely, despite being directly relevant to how drift correction and validation results should be interpreted. | Add to Horstwood TAPP, and keep a note of source.
| Masses Measured | Analyte + Monitored Isotopes | 1:2 (LAQSF splits; Monitored Isotopes new in v4) | LAQSF more granular | Horstwood's single field (raw mass list, 202-207/235/238) covers both the element-level scope (LAQSF's 'Analyte') and the isotope/interference-monitor-mass specifics (LAQSF's new 'Monitored Isotopes', added post-v2). Both Basic/Editable-or-Read-Only in the relevant LAQSF fields. | This worth more discussion. I think for MS there exists arguments to specify masses in "Analyte" rather than directly reporting elements or isotopes here. There should potentially be an extra field of "reported chemical variables" or "relevant elements". Let's flag this item and have more discussions.
| Integration Time per Peak/Dwell Times; Quadrupole Settling Time Between Mass Jumps | Dwell Time per Mass | 1:1 (renamed in v4) | Minor granularity difference | Renamed from 'Spectrometer Dwell Time' at some point after v2. Horstwood explicitly bundles settling time with dwell time; LAQSF's field implies but doesn't explicitly request settling time. | Here Table 3 from Horstwood also mentions "Quadrupole Settling Time", so we should consider adding fields relevant to Q-ICP-MS from LAQSF TAPP.
| Total Integration Time per Output Data Point (s) | -- (no equivalent) | Unique to Horstwood | Domain-emphasis, not domain-exclusive | LAQSF has no dedicated field for total mass-scan duty-cycle time / down-hole time resolution -- more consequential for geochronology (dissecting zoned grains/inclusions within a spot) than for standard trace-element spot analysis.
| 'Sensitivity' as Useful Yield (%, element) | -- (no equivalent) | Unique to Horstwood | Domain-emphasis, not domain-exclusive | Useful yield is the standard sensitivity metric in the geochronology literature, tied to the paper's emphasis on young/low-concentration samples pushing against detection limits.
| IC Dead Time (ns) | Pulse/Analog Detector Nonlinearity Correction (Group 5) | Partial / cross-group | LAQSF more actionable | Horstwood reports the raw hardware dead-time value; LAQSF instead asks whether a nonlinearity correction was applied and how -- arguably more actionable. | Keep consistent with LAQSF TAPP for cross-TAPP consistency. However, we should add "IC Dead Time" to the description explicitly if not done yet, in both TAPPs.
| -- (no equivalent) | Analysis Location/Spot Coordinates | Unique to LAQSF | Moderate gap, arguably more important for geochronology | Spot-location-relative-to-imaged-domain (core vs. rim) is arguably more consequential for geochronology than generic trace-element spot placement, yet Table 3 doesn't ask for it explicitly. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Number of Replicates | Unique to LAQSF | Moderate gap | Table 3's 'n=8'/'n=7' apply only to validation-material replicates, not to how many replicate analyses were run per unknown grain. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Multi-Run Sequential Analysis Design | Unique to LAQSF | Domain-appropriate (less common in geochronology) | Most single-collector/quadrupole geochronology procedures run U-Th-Pb in one pass per spot; multi-run designs are more common in trace-element work. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Mass Resolution per Analyte | Unique to LAQSF | NEW IN v4 -- likely genuine, important gap (SF cluster) | Per-analyte resolution assignment (which masses run low/medium/high resolution). Directly relevant to SF-ICP-MS geochronology, where resolving polyatomic interferences on Pb/Hg/U/Th masses without a collision cell is the whole point of using a sector instrument. Complete absence from Table 3 -- despite Table 3's own example being an HR/MC instrument -- is a substantive gap for any SF-focused extension of this TAPP. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | E-scan Range | Unique to LAQSF | NEW IN v4 -- likely genuine, important gap (SF cluster) | SF-specific electric-scan acquisition parameter. Explicitly named in the planning table's row 7d notes ('E-scan mode at fixed magnet') as central hardware for SF-ICP-MS geochronology specifically. Strong, well-evidenced candidate addition. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Triple Scanning Mode | Unique to LAQSF | NEW IN v4 -- likely genuine, important gap (SF cluster) | SF-specific noise-reduction acquisition parameter. Also named directly in the planning table's row 7d notes ('SEM triple mode (pulse/analog/Faraday)'). Same cluster as Mass Resolution per Analyte and E-scan Range. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Signal Collection Mode | Unique to LAQSF | Clean domain-appropriate exclusion (new in v4) | Explicitly tagged 'Q-ICP-MS only'; see CRC Configuration note in Group 3 -- Q-ICP-MS is not standard for geochronology per the planning table. | Per our discussion, it seems we should add it to Horstwood TAPP because some detrital zircons are measured with LA-Q-ICP-MS.
| -- (no equivalent) | Collision Gas Type | Unique to LAQSF | Clean domain-appropriate exclusion (new in v4) | Explicitly tagged 'Q-ICP-MS only'; same reasoning. | Per our discussion, it seems we should add it to Horstwood TAPP because some detrital zircons are measured with LA-Q-ICP-MS.
| -- (no equivalent) | Collision Gas Flow Rate | Unique to LAQSF | Clean domain-appropriate exclusion (new in v4) | Explicitly tagged 'Q-ICP-MS only'; same reasoning. | Per our discussion, it seems we should add it to Horstwood TAPP because some detrital zircons are measured with LA-Q-ICP-MS.
| -- (no equivalent) | Cell Exit Discrimination Voltage | Unique to LAQSF | Clean domain-appropriate exclusion (new in v4) | Explicitly tagged 'Q-ICP-MS only'; same reasoning. | Per our discussion, it seems we should add it to Horstwood TAPP because some detrital zircons are measured with LA-Q-ICP-MS.
| -- (no equivalent) | Reaction Gas Type | Unique to LAQSF | Clean domain-appropriate exclusion (new in v4) | Explicitly tagged 'Q-ICP-MS only'; same reasoning. | Per our discussion, it seems we should add it to Horstwood TAPP because some detrital zircons are measured with LA-Q-ICP-MS.
| -- (no equivalent) | Reaction Gas Flow Rate | Unique to LAQSF | Clean domain-appropriate exclusion (new in v4) | Explicitly tagged 'Q-ICP-MS only'; same reasoning. | Per our discussion, it seems we should add it to Horstwood TAPP because some detrital zircons are measured with LA-Q-ICP-MS.
| Other Information | -- (catch-all, not comparable) | Not comparable | N/A | Free-text catch-all; not a structured concept to map against LAQSF's field list.
|  |  |  | SUBTOTAL | Group 4: 13/18 Horstwood fields matched (72%, unchanged in kind from v13). LAQSF spot-applicable field count grew from 29 to 39 between v13 and v4. Of 26 LAQSF-unique fields: 6 are clean, planning-table-confirmed Q-ICP-MS-only exclusions (correctly not in Horstwood); 3 form a newly-surfaced SF-specific cluster (Mass Resolution per Analyte, E-scan Range, Triple Scanning Mode) that is now a much stronger, better-evidenced gap candidate than the single 'Mass Resolution Setting' flag from the v13 comparison; the remainder repeat the same 'likely genuine gap' / 'minimum-reporting framing' / 'became more relevant since 2016' split as before. One tier discrepancy surfaced (RF Power).
5. Data Processing | Gas Blank | Blank / Background Correction Method | 1:1 | Tier agreement | Both Basic/Read-Only. | Keep consistent with LAQSF TAPP for cross-TAPP consistency.
| Data Processing Package Used/Correction for LIEF | Elemental Fractionation Correction (+ Data Reduction Software, Group 3) | 1:2 (Horstwood bundles) | Split candidate confirmed | The 'correction method' half of Horstwood's bundled field maps cleanly to Elemental Fractionation Correction; the 'software name' half belongs in the (currently empty) Data Reduction Software field.
| Mass Discrimination | -- (no equivalent) | Unique to Horstwood | Clean domain-specific (correctly absent) | Trace-element quantification by external calibration is largely insensitive to instrumental mass bias. Isotope-ratio geochronology cannot ignore it. Textbook example of appropriate domain divergence in both directions. | Although trace-element quantification is insensitve to instrumental mass bias, shall we still add that to LAQSF TAPP for completeness?
| Common-Pb Correction, Composition and Uncertainty | -- (no equivalent) | Unique to Horstwood | Clean domain-specific (correctly absent) | A defining U-Pb geochronology step with no analog in general trace-element quantification.
| Uncertainty Level and Propagation | Uncertainty Propagation Method | 1:1 | Tier agreement | Both Advanced/Editable -- independent convergence on the same tier despite different development paths. | Keep consistent with LAQSF TAPP for cross-TAPP consistency.
| -- (no equivalent) | Internal Standard Approach + Internal Standard Element | Unique to LAQSF (2 fields) | Clean domain-specific (correctly absent) | Geochronology normalises isotope ratios against a matrix-matched age-standard reference material instead -- a different concept, already captured in Group 6. | Keep as is.
| (loosely related: Calibration Strategy, Group 6) | Per-Analyte Calibration Strategy | Unique to LAQSF | Domain-appropriate | Geochronology sessions overwhelmingly use one calibration strategy for the whole U-Th-Pb system. | Add to Horstwood TAPP, and keep a note of source.
| (loosely related: bundled LIEF mention) | Matrix Offset Correction (LIEF) | Unique to LAQSF | Niche refinement | A specific sub-concept not separately addressed by Table 3. Minor, advanced-practice gap. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Signal Integration Interval Method | Unique to LAQSF | Likely genuine gap -- important | How the 'clean' portion of a down-hole signal is distinguished from inclusion/crack/mixed-domain-affected material is arguably the single most consequential data-reduction decision in LA-ICP-MS U-Pb work, and directly relevant to the excess-scatter/MSWD discussion the paper itself devotes substantial prose to. | Add to Horstwood TAPP, and keep a note of source.
| (partial: Total Integration Time per Output Data Point, Group 4 -- a different, acquisition-timing concept) | Signal Integration Time | Unique to LAQSF | Likely genuine gap -- important | Not the same as Horstwood's raw-acquisition-duty-cycle concept. Same gap cluster as above. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Spike / Outlier Filtering Approach | Unique to LAQSF | Likely genuine gap -- important | Same signal-provenance theme as the two rows above. | Add to Horstwood TAPP, and keep a note of source.
| IC Dead Time (ns) (Group 4, raw value) | Pulse/Analog Detector Nonlinearity Correction | Partial / cross-group | LAQSF more actionable | See Group 4 note.
| -- (no equivalent) | Isobaric Interference Corrections Applied + Interfering Species + Interference Correction Method | Unique to LAQSF (3 fields) | Likely genuine gap -- important | 204Hg on 204Pb is the headline example, directly relevant to the Common-Pb Correction field, which often uses 204Pb as its basis. Complete absence from Table 3 is a substantive, well-evidenced gap. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Isotope Dilution Data Reduction Method | Unique to LAQSF | NEW IN v4 -- notable gap, directly tied to Horstwood's own example | Mass-balance approach for spike-sample isotope-ratio calculations. Highly relevant given Table 3's OWN example lab uses a co-aspirated Tl-235U tracer (Horstwood's 'Sample Introduction' field, Group 3) -- an isotope-dilution-style setup -- yet the paper never asks how the isotope-dilution math itself is reduced to a result. This is a case where Table 3 documents the hardware/setup for a technique (tracer co-aspiration) without documenting the corresponding data-reduction step, a genuine and fairly specific gap. | Add to Horstwood TAPP, and keep a note of source.
| -- (no equivalent) | Memory Effect Mitigation | Unique to LAQSF | Moderate gap | Would matter for geochronology reproducibility similarly to trace-element work. | Add to Horstwood TAPP, and keep a note of source.
| (loosely related: Mass Discrimination + Calibration Strategy combined) | Normalization / Standards-Based Correction | Unique to LAQSF | Minor overlap | Loose conceptual overlap, not a clean match either way. | Add to Horstwood TAPP, and keep a note of source.
|  |  |  | SUBTOTAL | Group 5: 3/5 Horstwood fields matched cleanly (60%), 2/5 clean domain-specific. LAQSF grew from 16 to 17 fields (+1 Isotope Dilution Data Reduction Method). Of 12 LAQSF-unique fields, 8 cluster in the signal-selection/interference-correction provenance theme (flagged as likely genuine, important gaps); Isotope Dilution Data Reduction Method is a new, specifically-evidenced gap given Table 3's own tracer-based example.
6. Quality Control & Uncertainty | Calibration Strategy + Reference Material Information | Primary Calibration Standard Name + Secondary Reference Materials | 2:2 (different split axis) | Different partition, same information | Horstwood splits by [which material plays which role] vs. [citations/accepted values]. LAQSF splits by [primary material, name+citation together] vs. [secondary materials, name+citation together]. Same total information, different axis -- same pattern as Group 3's laser wavelength/type split. | Keep consistent with LAQSF TAPP for cross-TAPP consistency.
| Quality Control/Validation | Within-Session Analytical Precision and Assessment Method + Between-Session (Long-Term) Analytical Precision and Assessment Method + Analytical Accuracy and Assessment Method | 1:3 (Horstwood bundles) | Split candidate confirmed -- biggest one | The single biggest split candidate in the whole comparison. Horstwood's one field really only demonstrates within-session precision-via-secondary-standard; it does not cleanly separate accuracy from precision, and has no equivalent of LAQSF's long-term/between-session reproducibility tracking at all. | Keep consistent with LAQSF TAPP for cross-TAPP consistency.
| -- (implied only, via validation-material 'n=8'/'n=7') | Calibration Standard Measurement Frequency | Unique to LAQSF | NEW LOCATION in v4 (renamed + moved from Group 4's 'Drift Monitor Frequency') | Renamed and relocated to Group 6 after the v2 sync point, per the dev log: moved 'for alignment with both solution TAPPs.' Same underlying gap as flagged before (Horstwood reports validation results but never states the bracketing interval) -- now filed under Group 6 to match LAQSF's current structure. | Add to Horstwood TAPP, and keep a note of source.
| Detection Limit | Detection Limit | 1:1 | Independent convergence | Added to Horstwood as an in-house gap-fill; confirmed identical to LAQSF's field, same tier (Advanced/Basic) -- validates that in-house addition.
| Detection Limit Method | Detection Limit Method | 1:1 | Independent convergence | Same as above.
| -- (no equivalent) | Limit of Quantification (LOQ) Method | Unique to LAQSF | Moderate gap -- arguably should have been added to Horstwood alongside Detection Limit | Same 'gap vs. Table 3' category as Detection Limit; in retrospect should probably have been added as a third in-house gap-fill field. | Add to Horstwood TAPP, and keep a note of source.
|  |  |  | SUBTOTAL | Group 6: all 5 Horstwood fields find a LAQSF counterpart (100%), but 2 of those 5 map many-to-many. LAQSF grew from 8 to 9 fields (+1, the relocated Calibration Standard Measurement Frequency). LAQSF's QC section remains substantially more decomposed than Horstwood's.

## Summary & Analysis
Horstwood et al. (2016) Table 3 vs. LA-Q/SF-ICP-MS TAPP (LA-Q:SF-ICPMS_TAPP_v4): Comparison Summary
Generated 2026-07-27 -- REVISED (see correction note below)
Correction vs. the previous version of this comparison
The prior comparison used LA-ICPMS_TAPP_v13.csv. Direct verification (file timestamps + field-set overlap analysis) showed this was the wrong reference file: LA-ICPMS v11 and LA-Q:SF-ICPMS v1 have a 100% identical field set (May 2026), and LA-ICPMS v12 and LA-Q:SF-ICPMS v2 also have a 100% identical field set (Jun 2026) -- the two folders were kept in exact parallel sync through two rounds. After that sync point, LA-Q:SF-ICP-MS continued alone to v3 (Jun 17, +15 fields) and v4, while the LA-ICP-MS-named folder received no further organic development -- v13 is v12's content untouched except for the same mechanical VIM3 terminology pass that also produced v4, at the identical timestamp. LA-Q:SF-ICPMS_TAPP_v4.csv is therefore the correct, actively-maintained TAPP to compare against. This revision replaces the v13-based findings below; where the conclusion changed materially, that is called out explicitly.
Scope of comparison
Horstwood TAPP v2 (57 fields, post-merge of overlapping paper/in-house pairs) compared against LA-Q:SF-ICPMS_TAPP_v4 (111 total fields), filtered to its Spot-applicable fields only, since geochronology is spot-only.
Overlap statistics by group
Group | Horstwood fields | Matched to LAQSF | Horstwood-unique | LAQSF spot-applicable fields | LAQSF-unique (no Horstwood match)
1. Procedure Identification | 17 | 17 | 0 | 17 | 0
2. Samples | 5 | 4 | 1 | 7 | 3
3. Instrument & Software | 7 | 6 | 1 | 16 | 10
4. Measurement Information | 17 | 13 | 4 | 39 | 26
5. Data Processing | 5 | 3 | 2 | 17 | 12
6. Quality Control & Uncertainty | 5 | 5 | 0 | 9 | 1
Totals: 56 comparable Horstwood fields, 48 matched to at least one LAQSF field (85%), 8 genuinely unique to Horstwood/geochronology. Of 105 spot-applicable LAQSF v4 fields, 52 have no Horstwood counterpart (49%) -- up from 39% against the (incorrect) v13 comparison, since v4 has 47 more spot-applicable fields than v13 did.
Finding 1 -- Group 1 agreement is by construction, not a signal
Group 1 shows 17/17 perfect alignment, but both TAPPs are required to build Group 1 from the identical cross-TAPP template (Rule 1). Unchanged from the v13 comparison -- Group 1 did not evolve between v12 and v4.
Finding 2 -- Core overlap is still strong, but LAQSF's lead has grown substantially
The absolute number of 1:1-or-close matches is essentially unchanged (~30 fields), but LAQSF's total spot-applicable field count grew from 92 (v13) to 122 (v4) -- entirely through the v2->v3 development round that the stale v13 branch never received. As a proportion, only 51% of LAQSF's fields now have a Horstwood counterpart, down from 61% in the (incorrect) v13-based comparison. This doesn't mean the workflows agree less than previously reported -- it means the v13 comparison was measuring LAQSF at an earlier, less mature stage of its own development.
Finding 3 -- The SF-ICP-MS-specific cluster is now a much stronger, better-evidenced gap
The v13 comparison flagged only one field (Mass Resolution Setting) as a likely SF-specific gap. v4 reveals this is really a cluster of at least four fields: Mass Resolution Setting (Group 3) plus Mass Resolution per Analyte, E-scan Range, and Triple Scanning Mode (all Group 4, all new since v2). E-scan Range and Triple Scanning Mode in particular use almost the exact language of the planning table's own row 7d note on LA-SF-ICP-MS Geochronology ('E-scan mode at fixed magnet, SEM triple mode (pulse/analog/Faraday)'). Table 3's own illustrative instrument (Nu Plasma HR, an MC-ICP-MS) doesn't exercise this cluster the way a single-collector SF instrument would, which plausibly explains why the paper's worked example never surfaced the need -- but for any SF-focused extension of the Horstwood TAPP, this is now a well-evidenced, concrete set of candidate additions rather than a single hunch.
Finding 4 -- The Q-ICP-MS-only fields are a clean, planning-table-confirmed non-gap
v4 adds a Collision/Reaction Cell (CRC) Configuration field plus six companion fields (Collision Gas Type, Collision Gas Flow Rate, Cell Exit Discrimination Voltage, Reaction Gas Type, Reaction Gas Flow Rate, Signal Collection Mode), all explicitly tagged 'Q-ICP-MS only' in LAQSF's own Comments column. The planning table's row 7d notes state directly that 'Q-ICP-MS [is] not standard for this [geochronology] application due to sensitivity constraints and Hg/Pb isobar complexity at unit resolution.' These seven fields being absent from Horstwood is therefore not a gap at all -- it's independent confirmation, from a completely different angle, of a decision the project had already made.
Finding 5 -- A new, specific gap tied directly to Table 3's own worked example
v4 adds Isotope Dilution Data Reduction Method (Group 5), describing the mass-balance approach for spike-sample isotope-ratio calculations. Table 3's own example lab uses a co-aspirated Tl-235U tracer (captured in Horstwood's 'Sample Introduction' field) -- an isotope-dilution-style setup -- but the paper documents the hardware/setup for this approach without ever documenting how the isotope-dilution math itself is reduced to a result. This is a narrower, more specific finding than the general 2016-era gaps discussed below: it's a gap in a paper's own worked example, not just a general community blind spot.
Finding 6 -- Two tier discrepancies surfaced independently of the geochronology question
Two fields shared by both TAPPs now disagree on analysis-level tier, in ways worth flagging back to LAQSF regardless of the Horstwood comparison: (1) Sample Preparation Method is Basic/Editable in LAQSF v4 but was Advanced/Editable in the v13 snapshot Horstwood was built against -- LAQSF promoted it to Basic at some point after the sync point. (2) RF Power is Basic/EDITABLE in LAQSF v4, changed from Basic/Read-Only in v13 -- but LAQSF's own field Description still reads 'Fixed in the procedure; cannot be changed by analysts without defining a new procedure,' which argues for Read-Only. This looks like an internal inconsistency (description vs. tier out of sync) introduced sometime in the v2->v4 development, not a deliberate reasoned change -- worth a look independent of this exercise.
Domain-specific vs. systematic bias vs. a third factor: production-process asymmetry (updated)
The v13-based conclusion mostly holds, and v4 sharpens it:
- Still NOT under-representation of geochronology seed papers: the newly-surfaced SF cluster and the Q-only exclusions both cut the other way from what 'wrong seed papers' would predict -- LAQSF's own literature assessment (13 Q/SF procedures) evidently already distinguished which CRC/collision-cell concepts are Q-specific and which resolution/E-scan concepts are SF-specific, and that granularity turns out to align closely with what geochronology specifically needs from the SF side. That's a workflow strength, not a domain blind spot.
- Production-process asymmetry is now the dominant explanation, more clearly than before: LAQSF gained 15 fields in a single development round (v2->v3, Jun 1-17) that the frozen v13 branch never received. The gap between the two TAPPs isn't primarily about geochronology being under-served by the AI workflow's seed papers -- it's substantially explained by comparing a document that has been through 4 real refinement rounds (v1-v4, May-Jul) against one that was built once, in 2016, and never revised.
- Where domain-specific content IS doing real work: the CRC/Q-exclusion cluster (Finding 4) and the SF-specific cluster (Finding 3) are the clearest cases where LAQSF's granularity maps onto genuine geochronology-relevant distinctions rather than just 'more fields because more development rounds.' The signal-selection/interference-correction cluster (Signal Integration Interval Method, Isobaric Interference trio, etc.) remains the strongest candidate for a genuine Table 3 content gap rather than a process-asymmetry artifact, since these are core data-provenance concepts any rigorous LA-ICP-MS geochronology paper needs regardless of which mass analyzer is used.
Answering your Basic-tiering question directly (unchanged from prior analysis)
Still yes, and still well-justified by the paper's own framing: Horstwood et al. state directly (p.324) that Table 3 is 'the recommended MINIMUM metadata types to be submitted for publication' -- mapping naturally to TAPP's C=Basic, not C=Advanced. Candidate fields worth revisiting against this principle: Pulse Width, Ablation Cell and Volume, Make-up Gas Flow, Ablation Pit Depth/Ablation Rate, IC Dead Time, 'Sensitivity' as Useful Yield, Imaging, and Uncertainty Level and Propagation. Not yet changed -- for your decision.
Corollary for LAQSF, updated: if the Basic-heavy reading is adopted, the geochronology-applicable LAQSF-only fields that would logically stay Advanced (sitting above Table 3's minimum bar) now include the SF cluster (Mass Resolution Setting, Mass Resolution per Analyte, E-scan Range, Triple Scanning Mode), the oxide/doubly-charged cluster, the isobaric-interference trio, Signal Integration Interval Method + Signal Integration Time, Analysis Sequence, Calibration Standard Measurement Frequency, and Isotope Dilution Data Reduction Method.
Paper-age (2016) considerations (unchanged)
Long-term/between-session reproducibility tracking, flat-top laser beam profiles, and signal-smoothing devices remain the three plausible 'became standard after 2016' items -- domain-knowledge inference, not a fresh literature search.
Bottom line (revised)
Correcting the reference file changes the numbers (LAQSF's true field-count lead over Horstwood is larger than first reported) but sharpens rather than overturns the qualitative story: most of the gap is explained by LAQSF having gone through more real development rounds, not by geochronology being poorly served by a general-purpose AI workflow. Two findings are genuinely new and specific to the correct v4 comparison: the SF-specific cluster is a much stronger candidate addition than previously stated, and the CRC/Q-only fields are a clean confirmation (not a gap) of the project's own existing architecture decision on row 7d.