## NMMP
Metadata Item | Description / Purpose | Basic/Advanced | Data Type | Example/Allowed Content | Last update: 04/09/2026
1. Method Identification
Method Name | A short descriptive name for this analytical method, with version number. Should include instrument, technique, analyte, and measured chemical variable. | Basic | Text (free) | e.g., 'LA-ICP-MS Trace Elements Glass v1'
Technique | Top-level analytical technique identifier. | Basic | Controlled list | LA-ICP-MS ¦ LA-ICP-OES ¦ LA-MC-ICP-MS ¦ Others: specify
Method Author | Author(s) of the method. Can be individual names or lab names. | Basic | Text (free) | e.g., individual names or lab names
Laboratory | Name of the laboratory or institution hosting the instrument. | Basic | Text (free) | e.g., 'Geo-/Cosmochemistry Lab, University of Cologne'
Laboratory ID | Persistent identifier for the laboratory, e.g., ROR ID. | Advanced | URI | ROR ID or equivalent URI (https://ror.org/)
Method Start Date | First date this method configuration was used in production. | Basic | Date | YYYY-MM-DD, e.g., 2022-04-22
Method DOI | DOI or unique persistent identifier for this method, generated after submission to a data repository. | Advanced | URI / DOI | e.g., http://doi.org/10.60520/IEDA/114187
Funding Source | Grants and other funding that supported instrument purchase, major upgrades, and personnel. Include grant IDs where applicable. | Advanced | Text (free) | e.g., 'DFG INST 216/1019-1 FUGG no. 665508'. Can also report "None"
Method Reference(s) | DOI or URL for publications or reports describing or evaluating this method. | Advanced | URI / DOI | e.g., doi:10.1016/j.oregeorev.2011.11.001
2. Samples
Target Material | General description of the material(s) analysed. | Basic | Controlled list / Text | Silicate glass ¦ Feldspar ¦ Pyroxene ¦ Olivine ¦ Oxide ¦ Sulfide ¦ Fluid inclusion ¦ Melt inclusion ¦ Other
Sample Preparation Method | Description of how samples were prepared for analysis (mounting, polishing, coating, etc.). | Advanced | Text (free) | e.g., 'Polished thin section, uncoated' ¦ '25 mm epoxy mount, polished to 0.25 µm diamond'
3. Instrument & Software
Laser Manufacturer & Model | Manufacturer and model of the laser ablation system. | Basic | Text (free) | e.g., 'ESI imageGEO193' ¦ 'Coherent GeoLas Pro' ¦ 'RESOlution M-50'
ICP-MS Manufacturer & Model | Manufacturer and model of the ICP-MS instrument. | Basic | Text (free) | e.g., 'Thermo Fischer iCAP Q' ¦ 'Agilent 7900' ¦ 'PerkinElmer NexION 2000'
ICP-MS Type | Type of ICP-MS detector configuration. | Basic | Controlled list | Single-collector (quadrupole) ¦ Multi-collector (MC-ICP-MS) ¦ Sector-field (SF-ICP-MS) ¦ Time-of-flight (TOF-ICP-MS) ¦ Other
Laser Wavelength | Wavelength of the laser used for ablation in nanometres. | Basic | Numeric (nm) | e.g., 193 (ArF excimer) ¦ 213 (Nd:YAG) ¦ 266 ¦ 532
Data Reduction Software | Software used for data reduction and concentration calculation, including version number. | Basic | Text (free) | e.g., 'Iolite 4' ¦ 'GLITTER 4.4' ¦ 'SILLS' ¦ 'LAtools'
Ablation Cell Type | Type and model of ablation cell used. | Advanced | Text (free) | e.g., 'Two-volume HelEx II' ¦ 'Single-volume standard cell'
4. Measurement Information
Laser Spot Geometry | Shape and dimensions of the laser ablation spot: width and length in micrometres. | Basic | Numeric pair (µm) | e.g., '15 µm circular' ¦ '20×20 µm square' ¦ '15; 20 µm (multiple sizes used)'
Laser Spot Path / Ablation Mode | Sampling mode or ablation pattern used during analysis. | Basic | Controlled list | Point ¦ Transect (line scan) ¦ Raster area ¦ Other
Laser Energy | Laser pulse energy in millijoules as set at the laser output. | Basic | Numeric (mJ) | e.g., 0.001 mJ ¦ 3.5 mJ
Laser Repetition Rate | Laser pulse repetition rate in hertz. | Basic | Numeric (Hz) | e.g., 5 ¦ 10 ¦ 20
Laser Pulse Duration | Duration of each laser pulse. | Advanced | Numeric + unit | e.g., 7 ns ¦ 25 fs
Transect Rate or Step Size | For line scans: continuous drive rate in µm/s; for stepped ablation: step size. | Advanced | Text (free) | e.g., 5 µm/s continuous ¦ 20 µm step
Carrier Gases and Flow Rates | Gas(es) used to transport ablated material from the ablation cell to the ICP-MS, with flow rates. | Basic | Text (free) | e.g., 'He (0.9 l min⁻¹) mixed with Ar (0.8 l min⁻¹)'
RF Power | ICP radiofrequency forward power in watts. | Basic | Numeric (W) | e.g., 1200 W ¦ 1400 W
Signal Smoothing | Description of any signal smoothing device or approach used between ablation cell and ICP-MS. | Advanced | Text (free) | e.g., 'Glass smoothing device' ¦ 'Squid 2.0' ¦ 'None'
ICP Tuning Procedure | Description of the method used to tune the ICP plasma conditions. | Advanced | Text (free) | e.g., 'Line scan ablating NIST612 using auto-tune function of Thermo iCAP Q'
Oxide Production (ThO/Th) | ThO/Th ratio (mass 248/232) or equivalent ratio used to monitor and minimize oxide production in the plasma. | Basic | Numeric (%) | e.g., ca. 0.7% ¦ <0.5% ¦ <1%
Drift Monitor Frequency | Frequency at which primary reference materials were analysed to monitor and correct for instrument drift. | Advanced | Text (free) | e.g., 'Every 10 unknowns' ¦ 'Not applicable' ¦ 'Start and end of session'
Analysis Sequence | Repeating order of calibration standard, quality control standard, and unknown analyses. | Basic | Text (free) | e.g., '2 primary std / 4 QC / 15 unknowns' ¦ '1/1/10'
Spectrometer Dwell Time | Count time per mass (dwell time) for each measured isotope in seconds. | Basic | Numeric (s) | e.g., 1×10⁻⁸ s per mass (quadrupole); element specific | Element specific
Analysis Count Time | Total integration time per analysis (on-peak) in seconds. | Basic | Numeric (s) | e.g., 40 s; element specific | Element specific
Background Count Time | Total time spent measuring gas blank (background) before each ablation signal in seconds. | Basic | Numeric (s) | e.g., 30 s; element specific | Element specific
5. Data Processing
Internal Standard Element | Element used as the internal standard to correct for variations in ablation yield and plasma conditions, and how its concentration value was determined. | Basic | Text (free) | e.g., 'Si (by EPMA)' ¦ 'Ca (by EDS)' ¦ '⁴³Ca; sample-average EPMA value'
Elemental Fractionation Correction | Method used to correct for elemental and isotopic fractionation during laser ablation, including reference material used. | Basic | Text (free) | e.g., 'Normalised to NIST 612; no additional downhole fractionation correction' ¦ 'NIST 612 with exponential downhole correction in Iolite'
Blank / Background Correction Method | Method used to subtract the gas blank (background signal) from the ablation signal. | Basic | Text (free) | e.g., 'Mean of 30 s pre-ablation gas blank subtracted per isotope'
Isobaric Interference Corrections Applied | Whether isobaric interference corrections were applied for any measured isotope. | Basic | Boolean | Yes ¦ No; element specific | Element specific
Interfering Isotopes | List of isotopes or molecular interferences overlapping with the measured isotope, if any. | Advanced | Text (free) | e.g., '⁸⁷Rb on ⁸⁷Sr' ¦ 'BaO on Eu' ¦ 'None' | Element specific
Interference Correction Method | Description of the method or equation used to correct for isobaric interferences. | Advanced | Text (free) | e.g., 'Measured ⁸⁶Sr and assumed natural ⁸⁷Rb/⁸⁵Rb to subtract ⁸⁷Rb contribution' | Element specific
Normalization / Standards-Based Correction | Any post-acquisition normalization applied using secondary reference materials. | Advanced | Text (free) | e.g., 'Normalized to NIST 610 measured in same session' ¦ 'None'
6. Quality Control & Uncertainty
Primary Calibration Standard Name | Primary reference material(s) used to calibrate and standardise the instrument. Include material source and accepted concentration reference used. | Basic | Text (free) | e.g., 'NIST SRM 612 (Pearce et al., 1997)' ¦ 'USGS BCR-2G' | Element specific
Secondary Reference Materials | Quality-control reference materials (secondary standards) routinely analysed alongside unknowns in the same session. Corresponding secondary standard data should be reported alongside sample data. | Basic | Text (free) | e.g., 'NIST SRM 610, ATHO-G, STHS-G' | Element specific | Include also in session metadata
Typical Detection Limit | Method detection limit at 99% confidence (3σ) for each measured isotope. Report value and method used. | Advanced | Numeric (ppm or wt%) | e.g., sample-individual LOD per Pettke et al., 2012; element specific | Element specific | Include also in session metadata
Detection Limit Method | Reference or description of the method used to calculate detection limits. | Advanced | Text / URI | e.g., 'Sample-individual LOD calculation per Pettke et al. (2012), doi:10.1016/j.oregeorev.2011.11.001'
Typical Analytical Precision | Reproducibility of repeated measurements on the same or equivalent reference material, expressed as 1σ relative standard deviation (%). | Advanced | Text (free) | e.g., '2% RSD (1σ, n=10 on NIST 612 for Ba)' | Element specific | Include also in session metadata
Typical Analytical Accuracy | Offset between measured and accepted reference values for secondary standards, expressed as % relative bias. | Advanced | Text (free) | e.g., '+1.5% relative to GeoReM preferred value for NIST 610 Ba' | Element specific | Include also in session metadata
Additional Notes | Any additional context about this method not captured by the structured fields above. | Advanced | Text (free)