## TAPP
Metadata Item | Description | Procedure-Level Tier | Analysis-Level Tier | Data Type | Example / Allowed Content | Comments | Last Update | Keyed By | Purpose | EDS Point Analysis | EDS Mapping | WDS Point Analysis | WDS Mapping | Literature Assessment | Ma+2015 ¦ Caltech GPS ¦ WDS Point Analysis (JEOL 8200) | Hu+2020 ¦ IGGCAS ¦ WDS Point Analysis (JEOL JXA-8100) | Liu+2016_UT ¦ Cameca SX100 ¦ WDS Mapping (U.Tennessee) | Liu+2016_Cal ¦ JEOL JXA-8200 ¦ WDS Point Analysis (Caltech GPS) | Ma+2017 ¦ JEOL 8200 ¦ WDS Point Analysis (Caltech GPS Analytical Facility) | Frank+2023 ¦ Cameca SX100 ¦ WDS Point Analysis (ARES JSC) | Broussard+2026 ¦ JEOL JXA-8200 ¦ WDS Mapping (WashU) | Seifert+2026 ¦ JEOL 8530 ¦ WDS Point Analysis (ARES JSC) | Pang+2016 ¦ JEOL JXA-8100 ¦ WDS Point Analysis (Nanjing U.) | McCoy+2025_SI ¦ JEOL 8530F+ ¦ WDS Point Analysis (Smithsonian) | McCoy+2025_UA ¦ Cameca SX-100 ¦ WDS Point Analysis (K-ALFAA U.Arizona) | Zega+2025 ¦ Cameca SX-100 Ultra ¦ WDS Point Analysis (K-ALFAA U.Arizona) | Barnes+2025 ¦ JEOL JXA-8230 ¦ WDS Point Analysis (CRPG Nancy) | Barnes+2025 ¦ Cameca SX100 ¦ WDS Point Analysis (NHM London) | Neuman+2025 ¦ WashU St. Louis ¦ WDS Mapping (JEOL JXA-8200)
1. Procedure Identification | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Procedure Name | A short descriptive name for this analytical procedure, including a version number. Should identify the instrument, the technique, and the scope of what is measured — target species, material, or feature. | Basic | Read-Only | Text (free) | e.g., 'EPMA-WDS Major Element Silicate Minerals' ¦ 'EPMA-WDS+EDS Minor Element Phosphates' | Source: Core module | 2026-09-01 | (none) |  | Y | Y | Y | Y |  | EPMA-WDS Major Element Silicates/Oxides, Tissint Mars Meteorite (Caltech GPS, JEOL 8200) | EPMA Major Element Silicates/Oxides, NWA 8657 Shergottite (IGGCAS, JEOL JXA-8100) | EPMA Major Element Silicates/Oxides+Mapping, Tissint (U. Tennessee, Cameca SX100) | EPMA Major Element Silicates/Oxides, Tissint (Caltech GPS, JEOL JXA-8200) | EPMA-WDS Major Element Silicates/Glasses, Zagami (Caltech GPS Analytical Facility, JEOL 8200) | EPMA Major/Minor Element Silicates+Oxides+Sulfides, CI Chondrite (ARES JSC, Cameca SX100) | EPMA-WDS Quantitative Mapping+Analysis, CI Chondrite Minerals (WashU, JEOL JXA-8200) | EPMA Major Element Apatite incl. Halogens, Bennu (ARES JSC, JEOL 8530 EMPA) | EPMA-WDS Major Element Silicates/Oxides, NWA 8003 Eucrite (Nanjing U., JEOL JXA-8100) | EPMA Carbonate, Magnetite and Olivine Composition, Bennu (Smithsonian, JEOL 8530F+ Hyperprobe) | EPMA-WDS Phosphate+Carbonate Composition, Bennu (K-ALFAA U.Arizona, Cameca SX-100) | EPMA Major Element Silicates/Sulfides/Oxides/Phosphates/Carbonates, Bennu (K-ALFAA, Cameca SX-100 Ultra) | EPMA Major Element Silicates/Oxides/Carbonates, Bennu Anhydrous Minerals (CRPG Nancy, JEOL JXA-8230) | EPMA Major/Minor Element Anhydrous Silicates, Bennu (NHM London, Cameca SX100) | EPMA-WDS quantitative compositional mapping, Apollo 17 core 73001 continuous thin sections (Washington University in St. Louis)
Technique | Top-level analytical technique identifier. | Basic | Read-Only | Controlled list | EPMA-WDS ¦ EPMA-EDS ¦ EPMA-WDS+EDS | Source: Core module | 2026-08-30 | (none) |  | Y | Y | Y | Y |  | EPMA-WDS | N | N | N | EPMA-WDS | N | EPMA-WDS | N | EPMA-WDS | N | EPMA-WDS | N | N | N | EPMA-WDS
Procedure Author | Person(s) or laboratory responsible for developing and registering this procedure. ORCID recommended for individuals. | Basic | Read-Only | Text (free) | e.g., 'Chi Ma, Caltech' ¦ 'ARES Geochemistry Lab, NASA JSC' | Source: Core module | 2026-07-24 | (none) |  | Y | Y | Y | Y |  | Chi Ma | Sen Hu | N | N | Chi Ma | N | N | Logan B. Seifert | N | T. J. McCoy | T. J. Zega | N | N | N | N (no individual operator named for the EPMA work)
Laboratory | Name of the laboratory or institution hosting the instrument. | Basic | Editable | Text (free) | e.g., 'Caltech GPS Division Analytical Facility' ¦ 'ARES, NASA Johnson Space Center' | Source: Core module | 2026-05-11 | (none) |  | Y | Y | Y | Y |  | Caltech GPS Division Analytical Facility | Institute of Geology and Geophysics, Chinese Academy of Sciences (IGGCAS), Beijing | Department of Earth and Planetary Sciences, University of Tennessee, Knoxville | Division of Geological and Planetary Sciences, Caltech | Division of Geological and Planetary Sciences Analytical Facility, Caltech | ARES, NASA Johnson Space Center | Washington University in St. Louis | ARES, NASA Johnson Space Center | State Key Laboratory for Mineral Deposits Research, Nanjing University | Smithsonian Institution, National Museum of Natural History | Kuiper-Arizona Laboratory for Astromaterials Analysis (K-ALFAA), University of Arizona | Kuiper-Arizona Laboratory for Astromaterials Analysis (K-ALFAA), University of Arizona | Centre de Recherches Pétrographiques et Géochimiques (CRPG), Nancy, France | Natural History Museum (NHM), London, UK | Washington University in St. Louis
Laboratory ID | Persistent identifier for the laboratory (e.g., ROR ID). | Advanced | Editable | URI / DOI | e.g., 'https://ror.org/05dxps055' (Caltech ROR) | Source: Core module | 2026-05-11 | (none) |  | Y | Y | Y | Y | [13 empty] | N | N | N
Procedure Start Date | First date this procedure configuration was used in production. | Basic | Read-Only | Date | YYYY-MM-DD | Source: Core module | 2026-07-24 | (none) |  | Y | Y | Y | Y | [13 empty] | N | N | N
Funding Source for Procedure Development | Grants and other funding sources that supported instrument acquisition, major upgrades, procedure development, and associated personnel time. Include grant numbers and funding agencies where applicable. | Advanced | Read-Only | Text (free) | e.g., 'NSF EAR-1234567' ¦ 'NASA 80NSSC20K0371' ¦ 'None' | Source: Core module | 2026-07-24 | (none) |  | Y | Y | Y | Y |  | NSF EAR-0318518; NSF DMR-0080065 — 'SEM, EBSD and EPMA analyses were carried out at the Caltech GPS Division Analytical Facility, which is supported, in part, by NSF Grants EAR-0318518 and DMR-0080065' | N | N | N | NSF EAR-0318518; NSF EAR-1322082; NSF DMR-0080065 — 'SEM, EBSD, EPMA, and Raman measurements were carried out at the Geological and Planetary Science Division Analytical Facility at Caltech, which is supported in part by NSF grants EAR-0318518, EAR-1322082, and DMR-0080065' | N | N | N | N | N | N | NASA Planetary Science Enabling Facilities 80NSSC23K0327; NASA Planetary Major Equipment NNX12AL47G and NNX15AJ22G; NASA Early Career Award 80NSSC20K1087; NSF Major Research Instrumentation 1531243 and 0619599; Gordon and Betty Moore Foundation; State of Arizona Technology and Research Initiative Fund — acknowledged 'for supporting K-ALFAA operations' and 'for supporting the instrumentation in K-ALFAA'; which award bought the microprobe is not stated | N | N | N
Procedure Reference(s) | DOI or URL for peer-reviewed publications or technical reports describing, validating, or benchmarking this procedure. | Advanced | Read-Only | URI / DOI | e.g., 'https://doi.org/10.1234/example' | Source: Core module | 2026-07-24 | (none) |  | Y | Y | Y | Y |  | Ma et al. 2015, Earth Planet. Sci. Lett. 422:194-205; doi:10.1016/j.epsl.2015.03.057 | Hu et al. 2020, Geochim. Cosmochim. Acta 278:185-198; doi:10.1016/j.gca.2019.06.012 | Liu et al. 2016, Meteorit. Planet. Sci.; doi:10.1111/maps.12726 | Liu et al. 2016, Meteorit. Planet. Sci.; doi:10.1111/maps.12726 | Ma et al. 2018, Meteorit. Planet. Sci. 53:50-61; doi:10.1111/maps.13000 (file dated 2017) | Frank et al. 2023, Meteorit. Planet. Sci. 58:1495-1511; doi:10.1111/maps.14083 | Broussard et al. 2026, Meteorit. Planet. Sci.; doi:10.1111/maps.70138 | Seifert et al. 2026, Meteorit. Planet. Sci.; doi:10.1111/maps.70167 | Pang et al. 2016, Sci. Rep. 6:26063; doi:10.1038/srep26063 | McCoy et al. 2025, Nature 637:320-325; doi:10.1038/s41586-024-08495-6 | McCoy et al. 2025, Nature 637:320-325; doi:10.1038/s41586-024-08495-6 | Zega et al. 2025, Nat. Geosci.; doi:10.1038/s41561-025-01741-0 | Barnes et al. 2025, Nat. Astron.; doi:10.1038/s41550-025-02631-6 | Barnes et al. 2025, Nat. Astron.; doi:10.1038/s41550-025-02631-6 | Neuman et al. 2025, J. Geophys. Res. Planets 130, e2024JE008556; doi:10.1029/2024JE008556 (section 2.6)
Procedure DOI | DOI or unique persistent identifier for the registered procedure used in this analysis, generated upon procedure submission to a data repository. If a DOI has been applied for but not yet minted, enter "pending". | N/A | Basic | URI / DOI | e.g., 'https://doi.org/10.xxxxx/procedure-id' ¦ 'pending' | Source: Core module | 2026-08-25 | (none) | Encourages formal procedure registration and ensures traceability. | Y | Y | Y | Y | [13 empty] | N | N | N
Session Identifier | Identifier for the analytical session this record describes — the laboratory's own run, sequence or batch identifier as generated by the instrument or acquisition software. Distinct from any persistent identifier a repository mints on submission. | N/A | Basic | Text (free) |  | Source: Core module | 2026-08-25 | (none) | The analysis record corresponds to one session, which may cover several samples, and this is the link back to the raw instrument files. | Y | Y | Y | Y | [15 empty] | N
Analyst | Name(s) of the analyst(s) who performed the analysis session. ORCID is recommended for persistent identification. | N/A | Basic | Text (free) | e.g., 'Sen Hu' | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y |  | Chi Ma | Sen Hu | N | N | Chi Ma | David R. Frank | N | Logan B. Seifert | N | N | N | N | N | N | N
Analysis Start Date | Date on which the analytical session began. For sessions spanning multiple days, use the date of the first session. Format: YYYY-MM-DD. | N/A | Basic | Date | YYYY-MM-DD | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y | [13 empty] | N | N | N
Analysis End Date | Date on which the analytical session ended. May equal Analysis Start Date for single-day sessions. Format: YYYY-MM-DD. | N/A | Basic | Date | YYYY-MM-DD | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y | [13 empty] | N | N | N
Funding Source for Analysis | Grants and other funding sources that supported the specific sample analyses and data collection reported here. Include grant numbers and funding agencies. Distinct from procedure development funding, which may have been awarded years earlier. | N/A | Basic | Text (free) | e.g., 'NSF EAR-2345678' ¦ 'NASA 80NSSC23K0415' ¦ 'None' | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y |  | DOE Cooperative Agreement DE-NA0001982; NSF EAR 322082; NASA NNN13D465T; NASA Cosmochemistry NNX11AG58G and NNX12AH63G; NSF EAR 1344942 and 1440005 — acknowledged author by author for the study as a whole; none is attributed to the microprobe work | NSFC 41573057, 41430105 and 41973062; China Scholarship Council 201804910284; IGGCAS key research program IGGCAS-201905 — 'This study was financially supported by ...' | NASA Cosmochemistry NNX11AG58G and NNN13D465T; NSF EAR-1226270; NSF EAR-1019770 — 'We acknowledge partial support by ...' | NASA Cosmochemistry NNX11AG58G and NNN13D465T; NSF EAR-1226270; NSF EAR-1019770 — 'We acknowledge partial support by ...' | NNSA Stewardship Science Academic Alliances, DOE Cooperative Agreements DE-NA0001982 and DESC0005278; NASA NNX12AJ01G — 'This work was supported in part by ...' | NASA NNX11AG78G, 13-COS13-0026, NNX14AI19G and 80NSSC18K0586; NASA Cosmochemistry and Emerging Worlds Programs — 'This work was supported by ...' | NASA 80NSSC22K1689; NASA 80NSSC24K1284; McDonnell Center for the Space Sciences — 'Individual funding was provided by ...' | NASA NNH09ZDA007O under contract NNM10AA11C (New Frontiers Program); NASA Postdoctoral Program — acknowledgements | NSFC 41373065; State Key Laboratory for Mineral Deposits Research ZZKT-201322; Fundamental Research Funds for the Central Universities — 'This study was supported by grants from ...' | NASA NNH09ZDA007O under contract NNM10AA11C (New Frontiers Program) — 'This material is based on work supported by ...'; individual grants are also listed by person (80NSSC22K1692, MR/T020261/1, ST/V000675/1, 22EXPOSITO) | NASA NNH09ZDA007O under contract NNM10AA11C (New Frontiers Program) — 'This material is based on work supported by ...'; individual grants are also listed by person (80NSSC22K1692, MR/T020261/1, ST/V000675/1, 22EXPOSITO) | NASA NNH09ZDA007O under contract NNM10AA11C (New Frontiers Program) — 'This material is based upon work supported by ...'; the K-ALFAA facility and instrumentation support is under Funding Source for Procedure Development | NASA NNH09ZDA007O under contract NNM10AA11C (New Frontiers Program) — acknowledged with the named authors; further grants are listed by person across many laboratories, and none is attributed to the microprobe work | NASA NNH09ZDA007O under contract NNM10AA11C (New Frontiers Program) — acknowledged with the named authors; further grants are listed by person across many laboratories, and none is attributed to the microprobe work | NASA ANGSA program 80NSSC19K0958 — 'We thank NASA for support of the ANGSA program (Grant 80NSSC19K0958)'
Coupled Technique(s) | Other analytical techniques applied to the same sample(s) whose results are intended to be interpreted together with data from this procedure. Document coupling with any technique whose results are functionally linked to this dataset — providing calibration inputs, complementary spatial context, or required companion measurements. Name the instrument, laboratory or purpose alongside each technique where known, and join multiple techniques with '; '. Enter "None" if no coupling is intended. | Advanced | Editable | Text (free) | e.g., 'EPMA (Brown University CAMECA SX-100)' ¦ 'FIB-SEM (Crossbeam 540) for 3D tomography; EDS for mineral analysis' ¦ 'NanoSIMS; STXM; Raman' ¦ 'None' ¦ 'N/A' | Source: Core module | 2026-08-31 | (none) |  | Y | Y | Y | Y |  | SEM (Carl Zeiss 1550VP FE-SEM, BSE imaging); EBSD (HKL system on ZEISS 1550VP); synchrotron XRD; micro-Raman | SEM-EDS (FEI Nova NanoSEM 450); Raman spectroscopy | SEM (BSE imaging); petrographic microscopy; LA-ICP-MS (Agilent 7500ce, Virginia Tech) | SEM (BSE imaging); petrographic microscopy; LA-ICP-MS (Agilent 7500ce, Virginia Tech) | SEM (Carl Zeiss 1550VP FE-SEM, BSE imaging); EBSD; synchrotron XRD; micro-Raman | Petrographic microscopy; SEM-BSE (JEOL 5900LV); EPMA X-ray mapping; Cameca ims1280 ion microprobe (O isotopes; 26Al-26Mg); FIB-TEM | Powder XRD (Rigaku MiniFlex 600); ICP-MS (Thermo Fisher iCAP Qc, WashU); K isotope MC-ICP-MS (Neptune Plus, WashU); CO2 laser-fluorination O isotope MS (U. New Mexico); AMS (PRIME Lab, Purdue) | SEM-EDS; SIMS (Cameca ims 1280); TEM-EDS | SEM (BSE imaging); petrographic microscopy | SEM-EDS (NHM London; Smithsonian; JSC); TEM-EDS/EELS; FIB-SEM; ToF-SIMS; XRD; XANES | SEM-EDS (NHM London; Smithsonian; JSC); TEM-EDS/EELS; FIB-SEM; ToF-SIMS; XRD; XANES | SEM-EDS; TEM-EDS/EELS; FIB-SEM; XRD; XANES | SEM-BSE (JEOL JSM-6510, 15 kV, 3 nA); SEM-EDS (multi-element mapping); SIMS (CAMECA IMS 1270 E7, CRPG); NanoSIMS (K-ALFAA); ICP-MS; MC-ICP-MS; noble gas MS | SEM-EDS (Zeiss EVO 15LS + Oxford X-Max80, 20 kV, 3 nA); NanoSIMS (OU); SIMS (CAMECA ims-1280-HR, Hokkaido); laser fluorination O isotopes (OU) | BSE mosaic imaging (same JEOL JXA-8200); QEMSCAN (FEI QUANTA 650 FEG-SEM, Univ. Manchester); optical microscopy (Keyence VHX 7000, NASA JSC); micro-XCT (custom NSI instrument, UTCT)
Coupling Description | Description of how this procedure is coupled with the technique(s) listed above. Include: (1) the functional relationship — what data or context flows between techniques, or how results are combined (e.g. which output from the coupled technique serves as input to data reduction for this technique); and (2) the analytical sequence — which technique is performed first and why (e.g. non-destructive before destructive). Required when Coupled Technique(s) is not "None". | Advanced | Editable | Text (free) | e.g., 'EPMA WDS major element concentrations used as internal standard for LA-ICP-MS trace element data reduction. EPMA performed first because LA-ICP-MS ablation is destructive.' | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | Single-element and RGB composite X-ray maps are compared with BSE and optical image mosaics (reflected light, plane polarized, crossed polars) to discriminate crystalline from glassy phases; an Al-Mg-Fe RGB composite X-ray map is used to discriminate feldspathic (red) from ferromagnesian (green and blue) phases
Coupled Procedure DOI | Registered procedure DOI for the coupled technique named above. If the coupled procedure has not yet been registered, enter the DOI of a publication describing the coupled method, or "pending". Enter "None" if no coupling is planned. | N/A | Advanced | URI / DOI | e.g., 'https://doi.org/10.xxxxx/coupled-procedure' ¦ 'pending' ¦ 'None' | Source: Core module | 2026-08-25 | (none) | Provides a stable, citable link to the companion method independent of whether a dataset has been deposited. | Y | Y | Y | Y | [13 empty] | N | N | N
Coupled Dataset or Publication Reference | DOI or other persistent identifier for the co-registered dataset or publication where both datasets are reported together. Accepts: a dedicated dataset DOI (if separately deposited), a shared dataset DOI (if co-submitted in the same package), or a publication DOI. Use "same submission" if the coupled dataset is included in this data package, or "pending" if not yet assigned. If coupling is documented through a shared sample identifier only, that information is already captured in Sample Persistent Identifier (Group 2). | N/A | Advanced | URI / DOI / Text | e.g., 'https://doi.org/10.xxxxx/dataset' ¦ 'same submission' ¦ 'pending' ¦ 'None' | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y |  | N | https://doi.org/10.1016/j.gca.2020.07.021 — 'Supplementary data to this article can be found online at ...' | N | N | N | same submission — 'All data are included within the manuscript' | N — 'The data that support the findings of this study are available from the corresponding author upon reasonable request' | astromat.org, at the DOIs in Table S1 — 'Instrument data products that support the findings of this study will be available via astromat.org at the DOIs given in Table S1' | Supplementary Information at www.nature.com/srep — 'Supplementary information accompanies this paper at http://www.nature.com/srep'; the EPMA data are in Supplementary Tables 1–4 | astromat.org, at the DOIs in Extended Data Table 3 — 'Instrument data supporting the experimental results from the samples analysed in this study will be available from Astromat (astromat.org) at the DOIs listed in Extended Data Table 3' | astromat.org, at the DOIs in Extended Data Table 3 — as for the Smithsonian procedure | N | N | N | Ogliore (2025) — 'The quantitative EPMA image data and coregistered image sets are available in Ogliore (2025)'
2. Samples | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Target Material | General description of the material type(s) this procedure is designed to analyse. | Basic | Read-Only | Controlled list / Text | Silicate mineral ¦ Silicate glass ¦ Oxide ¦ Sulfide ¦ Carbonate ¦ Phosphate ¦ Metal or alloy ¦ N/A ¦ None | Source: Core module | 2026-09-28 | defines: target material | Used for discoverability and procedure matching, and because the material type constrains sample preparation, calibration and matrix-matching requirements. | Y | Y | Y | Y |  | tissintite; maskelynite; pigeonite; fayalite — Table 1: wormy and rimming tissintite, associated maskelynite, pigeonite and fayalite surrounding tissintite | maskelynite; melt inclusion glasses; silica glasses; coesite aggregates; mesostasis — 'Quantitative analyses of maskelynite, melt inclusion glasses, silica glasses, coesite aggregates, and mesostasis were conducted by EPMA' (p.2) | olivine; pyroxene; Fe-Ti-Cr oxides; maskelynite; phosphate; sulfide; glass — the analytical conditions are given for these groups (p.3) | olivine; pyroxene; Fe-Ti-Cr oxides; maskelynite; phosphate; sulfide; glass — the analytical conditions are given for these groups (p.3) | liebermannite; lingunite; maskelynite — Table 1: 'Analytical data for liebermannite and associated lingunite and maskelynite' | melilite; spinel; grossmanite — Table 1: 'Chemical compositions for the major minerals in the Ivuna CAI' | phyllosilicate (matrix); oxide (magnetite, ilmenite); sulfide (pyrrhotite, pentlandite); carbonate (dolomite, magnesite); phosphate (Ca phosphate, Na-Mg hydrous phosphate) — phases analysed and mapped by EPMA (results) | Phosphate (apatite) | plagioclase and its polymorphs — 'Measurements of most minerals were performed with a focused beam ... whereas measurements of plagioclase and its polymorphs were performed with a defocused beam'; the other minerals are not listed (EPMA data in Supplementary Table 4) | carbonate; magnetite; olivine — 'Carbonate analyses were run at 15 kV and 10 nA ...; analyses of magnetite and olivine were run under the same conditions' (p.7) | "Mg,Na phosphate"; carbonate — p.7 | silicates; sulfides; oxides; phosphates; carbonates — p.9 | carbonates — 'Quantitative chemical analyses ... to determine the chemical compositions of minerals'; carbonates are the only material named with conditions of their own | olivine; pyroxene — 'Olivine and pyroxene grains were identified and characterized at the NHM' (p.13) | lunar regolith (Apollo 17 double drive tube, lower section 73001) — as continuous thin sections
Sample Preparation Method | The form in which the sample is presented to the instrument, and the preparation that brought it to that form — for example mounting, sectioning, polishing, coating, crushing, fusion, or extraction of an electron-transparent section. Record 'None' where the material is analysed as received. | Basic | Editable | Controlled list / Text | Polished thin section ¦ Polished thick section ¦ Polished block or epoxy mount ¦ Grain mount ¦ Fragments or particles on conductive tape ¦ Bulk specimen, polished face ¦ Acid-dissolution residue ¦ N/A ¦ None | Source: Core module | 2026-09-28 | sample |  | Y | Y | Y | Y |  | Polished thin section; carbon coating N | Polished thick section (NWA 8657); carbon coating N | Polished thin sections (coating type N) | Polished thin sections (coating type N) | Polished thin section USNM 7619 (coating type N) | N | Fragments mounted and dry-polished in a petrographic thin section; carbon coating N | Fragments embedded in epoxy; dry-polished with diamond powder; one mount ion-polished before carbon coating | Polished thin section; carbon coating N | Ir-coated specimens — 'Electron microprobe analysis was conducted on Ir-coated specimens'; the mounting for the microprobe work is not stated | Polished section; 20 nm carbon coat | N | Aggregate particles (<1 mm) mounted in epoxy at Université Côte d'Azur; polished; carbon coated (thickness N) | Mounted in resin blocks; polished at NHM London; fragmented during polishing (P1, P2); initial carbon coat for SEM/EPMA (thickness N); additional coat added after for SIMS (total ~30 nm) [N5,N6] | Core extruded and dissected in 0.5 cm depth intervals; the remaining material impregnated with epoxy to create a continuous thin section set of the entire core; sections are 50 x 25 mm. Carbon coating N
Sample Name | Name or identifier of each sample analysed in this session, as used in the laboratory — a sample mount, section or aliquot counts as one entry where that is what the laboratory tracks. Should match the identifier used in associated publications or data tables. | N/A | Basic | Text (free) | e.g., 'NWA 8657 polished thin section 1' ¦ 'Tissint UT1' ¦ 'Ivuna MZ2' | Source: Core module | 2026-08-25 | defines: sample | The analysis record corresponds to one session and may cover several samples. | Y | Y | Y | Y |  | Tissint Mars meteorite | NWA 8657 shergottite | Tissint thin sections UT1, UT2, UT3; Tata-1-C1 to C3; Tata-2-C1 to C3; Tata-3-C1 to C3; Tissint-B | Tissint thin sections UT1, UT2, UT3; Tata-1-C1 to C3; Tata-2-C1 to C3; Tata-3-C1 to C3; Tissint-B | Zagami USNM 7619 | Ivuna CI chondrite, section MZ2 | OC002 LAB24-2 (10-11 fragments of Oued Chebeika 002) | OREX-803079-0 and OREX-803080-0 (Bennu OSIRIS-REx samples) | NWA 8003 eucrite | Bennu OSIRIS-REx samples (OREX-8#####-###) | Bennu OSIRIS-REx samples (OREX-8#####-###) | Bennu OSIRIS-REx samples (OREX-5/8#####-###) | OREX-800045-103 and OREX-800045-107 | OREX-501054-0 and OREX-501059-0 (particles P1, P2) | 73001,6014-73001,6021
Sampling Unit Type | The kind of physical subdivision of the sample to which one row of reported values corresponds — the unit that is analysed and reported, as distinct from the sample as a whole. Where units nest (e.g. confined tracks within grains), state both levels. The units themselves are listed in Sampling Unit Name. | Basic | Read-Only | Controlled list / Text | Whole sample ¦ Aliquot ¦ Grain ¦ Spot ¦ Analysis point ¦ Phase ¦ Sub-volume ¦ Region of interest ¦ N/A ¦ None | Source: Core module | 2026-09-15 | (none) |  | Y | Y | Y | Y |  | Phase > Analysis point — Table 1 reports one column per phase and textural setting ("Wormy type tissintite", "Maskelynite away from melt pockets" …), each the mean of n = 5–17 focused-beam point analyses (p.5) | Phase > Analysis point — "Quantitative analyses of maskelynite, melt inclusion glasses, silica glasses, coesite aggregates, and mesostasis were conducted by electron probe microanalysis" (p.2); points are grouped by phase, not reported individually in the archived PDF | Whole sample (thin section) > Phase — "elemental X-ray maps (Ca Ka, Al Ka, Fe Ka, and Mg Ka) of four sections" (p.3); the reported quantity is a modal fraction, "the number of pixels attributed to each mineral ... divided by the total number of pixels in the whole section" (p.3) | Phase > Analysis point — "Major and minor element compositions of selected minerals" by phase (Table 2, p.6); glass compositions are means, "EMP avg (n = 73)" and "avg (n = 14)" (Table 3, p.9) | Grain > Analysis point — Table 1 reports one column per occurrence ("Type liebermannite", "The second liebermannite", "The third liebermannite", "Lingunite next to type liebermannite" …), each the mean of n = 2–6 points (p.3) | Phase > Analysis point — "Representative electron-microprobe measurements of melilite, grossmanite, and spinel are given in Table 1" (p.5), all within the single Ivuna CAI | Region of interest > Phase — "wavelength-dispersive quantitative compositional mapping and analysis" (p.3) of whole fragments; phases are identified within a map ("Round Phy1 phyllosilicate clast", "Lithic clast (LC1)", Fig. 3, p.6) and abundances given per section ("sulfides which make up 2.3 areal% of the section", p.6) | Grain > Analysis point — Table 1 reports one column per apatite grain ("Ap. #1" … "Ap. #5") within each of the three particles (p.7) | Phase > Analysis point — compositions are reported as per-phase means, "The average compositions (Supplementary Table 1) of orthopyroxene" (p.2) and "41 ± 8 mol% on average; based on 13 analyses" (p.4); the individual points are in a supplement not in the archived PDF | Phase > Analysis point — "Electron microprobe analysis was conducted on Ir-coated specimens" (p.7); results are reported by phase, e.g. "the calcite has near-end member composition (4 mol.% or less MgCO3 and FeCO3)" (p.2) | Phase > Analysis point — "EMPA analyses were carried out using a Cameca SX-100 electron microprobe located at K-ALFAA" (p.7); results are reported by phase, not per named point | Grain > Phase — sulfide compositions are reported per grain within named particles ("Pyrrhotite compositions measured by EMPA", p.2), and "phase mapping via electron microprobe analysis (EMPA)" gives modal abundances of carbonates, sulfides and magnetite per particle (p.2) | Grain > Analysis point — "Aggregate particles (<1 mm) were mounted in epoxy" and mapped "of the different grains"; "Quantitative analyses were performed with ... beam diameter of 1 µm", the beam rastered "over 5 × 5 µm2" for carbonates (p.11) | Grain > Analysis point — "Olivine and pyroxene grains were identified and characterized at the NHM" in particles P1 and P2, and "Analyses were performed at 20 kV, using a focused 1-μm beam" (p.13) | Analysis point - each map pixel is a fully quantitative analysis (1,024 x 1,024 per stage map; five stage maps per thin section; 20 x 10^6 analyses across all slides)
Sampling Unit Name | The name or label of each sampling unit analysed in this session, as the laboratory records it, together with the sample it belongs to — e.g. a spot number, a grain label, a map or region-of-interest name, or an aliquot identifier. Where units are too numerous to name individually, such as map pixels or reconstructed voxels, name the acquisition area they belong to instead. | N/A | Basic | Text (free) | e.g., NWA 8657-1: spots 1–24 ¦ Zircon mount ZM-3: grains Z1–Z30 ¦ Core 73001: map areas A1–A5 ¦ Bennu aggregate OREX-803015-100: aliquot for Ti | Source: Core module | 2026-09-15 | defines: sample > sampling unit | Gives every field keyed by sampling unit a row to attach its values to, and ties each unit to the sample it belongs to; without it those fields collapse to one value per session. | Y | Y | Y | Y |  | Labelled at section level only: tissintite "was identified in sections UT1, UT2 and UT3" (p.2); Table 1 groups the point analyses by phase and setting ("Wormy type tissintite", "Rimming tissintite", "Maskelynite away from melt pockets" …) with counts (n = 6, 9, 17 …), not labels (p.5) | Sample name only — "a polished thick section of the martian meteorite NWA 8657" (p.2), not otherwise labelled; analyses are grouped by phase ("maskelynite, melt inclusion glasses, silica glasses, coesite aggregates, and mesostasis", p.2) | Labelled at section level: X-ray maps are placed by thin section, e.g. an olivine megacryst "in Tata-2-C3" whose "Red box outlines the area of X-ray maps" (Fig. 3 caption, p.7); the sections are "UT1 to UT3" and "Tata-1-C1 to C3, Tata-2-C1 to C3, and Tata-3-C1 to C3" (pp.2–3). Map areas carry no labels of their own | Labelled: thin sections "UT1 to UT3" and "Tata-1-C1 to C3, Tata-2-C1 to C3, and Tata-3-C1 to C3" (pp.2–3), and pyroxene grains within a section, e.g. "T-3-C2 px1" to "T-3-C2 px5" (Fig. 8, p.12); individual point analyses are not labelled | Labelled at section level only: "The Zagami thin section (USNM 7619)" (p.3); occurrences are identified by figure panel ("First occurrence of liebermannite", Fig. 1, p.2), not by label | Labelled at section level: the CAI "in Ivuna section MZ2" (p.3), in "polished mount 'Ivuna MZ2'" (p.14); the single CAI and its analysis points carry no labels | Labelled: "OC002 LAB24-2 fragment 1" and "fragment 2" (Figs 3–4, pp.6–7), and within fragment 1 "lithic clast LC1" (Figure S3 caption, p.17); map areas are otherwise unlabelled | Labelled: apatite grains numbered within each particle — "OREX-803166-0 … Ap. #1 Ap. #2 Ap. #3 Ap. #4" (Table 1, p.7) — in particles "OREX-803166-0, OREX-803169-0, and OREX-803173-0" (p.3) from the mounts "OREX-803079-0 and OREX-803080-0" (p.2) | Sample name only — NWA 8003; grains are identified by setting ("garnet grains within the eclogitic mineral assemblage", p.4), and the point analyses are in "Supplementary Table 4" (p.4), not in the archived PDF | N — the Smithsonian microprobe passage names no specimens ("conducted on Ir-coated specimens", p.7); the paper's OREX numbers identify figure images, not microprobe analyses | N — the Arizona microprobe passage names no specimen ("the section was coated with a thin film of carbon", p.7); the paper's OREX numbers identify figure images, not microprobe analyses | Labelled by curation number per particle: "EMPA data from Bennu particles" (Fig. 1) cites "OREX-803095-0" and "OREX-803096-0", and sulfide compositions "in samples OREX-803095-0, OREX-803096-0, OREX-803066-0, OREX-803067-0 and OREX-803070-0" (p.2); points within particles are not labelled | Labelled at split level: "Samples OREX-800045-103 and OREX-800045-107 ... Aggregate particles (<1 mm) were mounted in epoxy" (p.11); particles and points are not labelled, and the data are "compiled in Supplementary Table 14" (p.11), not in the archived PDF | Labelled: "The samples OREX-501054-0 and OREX-501059-0 ... fragmented into particles, identified as P1 and P2" (p.13); the olivine and pyroxene grains within them are not labelled | Labelled at thin-section level: "The 73001,6014–73001,6021 samples are 50 × 25 mm continuous thin sections" (p.6), e.g. the "73001,6019" X-ray map (p.11); every map pixel is an analysis, so the section — the acquisition area — is the unit named
Target Material of Sampling Unit | The entry in Target Material that the sampling unit belongs to, which links the unit to the point-analysis conditions registered for that material. | N/A | Basic | Text (free) | e.g., 'Silicate mineral' for points P1-01–P1-12; 'Phosphate' for points P3-01–P3-06 |  | 2026-09-28 | sample > sampling unit | Without it a consumer cannot tell which of the conditions registered per target material a given analysis point was acquired under. | Y | N | Y | N |  | each unit is named by its material: tissintite (wormy, rimming), maskelynite (three settings), pigeonite, fayalite — Table 1 | N | N | N | liebermannite (type, second and third occurrences); lingunite (next to type liebermannite); maskelynite (near type and near second liebermannite) — Table 1 | N | N | Phosphate (apatite) for every grain, Ap. #1 onward — Table 1 | N | N | N | N | N | N | N
Sample Persistent Identifier | Globally unique, persistent identifier for each sample listed in Sample Name. IGSN (International Geo Sample Number) is the recommended standard for geological and cosmochemical samples, as used by Astromat, EarthChem and SESAR. Where a sample and its sub-samples are separately registered, record the identifier at the level actually analysed. | N/A | Advanced | URI / IGSN | e.g., 'IGSN:AU1234567' | Source: Core module | 2026-09-15 | sample |  | Y | Y | Y | Y |  | N | N | N | N | USNM 7619 — Smithsonian catalogue number: 'The Zagami thin section (USNM 7619)'; no IGSN | N | N | OREX-803079-0; OREX-803080-0 — NASA curation numbers of the mounts, not IGSN | N | OREX-8#####-### (NASA curatorial ID, not IGSN) | OREX-8#####-### (NASA curatorial ID, not IGSN) | OREX-5/8#####-### (NASA curatorial ID, not IGSN) | OREX-800045-103; OREX-800045-107 (NASA curatorial IDs, not IGSN) | OREX-501054-0; OREX-501059-0 (NASA curatorial IDs, not IGSN) | N (NASA curation sample numbers given as names; not stated as persistent identifiers)
Sampling Unit Selection Criteria | The rules governing which sampling unit(s) within a sample are selected for analysis, and why. Covers the criteria applied when choosing grains, aliquots, spots, or a region of interest — size, morphology, clarity, freedom from inclusions or alteration, phase identity, or spatial position. Distinct from Target Material, which states the material type the procedure is designed for, and from Sampling Unit, which names the kind of subdivision one row of reported values corresponds to: this field states how, within such a sample, the unit actually analysed is picked out. | Basic | Editable | Text (free) |  | Source: Sampling Unit Selection module | 2026-09-01 | (none) |  | Y | Y | Y | Y |  | Textural position relative to the shock-melt pockets — the analyses are grouped as "Wormy type tissintite", "Rimming tissintite", "Maskelynite associated with wormy tissintite" and "Maskelynite away from melt pockets" (Table 1, p.5), the phase itself occurring "only in maskelynite less than ~25 μm of a shock melt pocket" (p.1) | N — the paper names the phases it analysed (see `Sampling Unit Type`) but states no rule for choosing the individual units | N — the map areas are shown rather than specified: a "Red box outlines the area of X-ray maps" on an olivine megacryst (Fig. 3 caption, p.7), with no stated rule for placing them | N — the paper names the phases it analysed (see `Sampling Unit Type`) but states no rule for choosing the individual units | Occurrence of the target phase — the units analysed are the observed occurrences of the new mineral: "A first occurrence of liebermannite was observed with lingunite, silica, ilmenite, and baddeleyite ... (this is the type material). A second occurrence ... A third occurrence is shown in Fig. 1c, close to the type occurrence" (p.4) | Opportunistic — the object analysed "was found by David Frank in Ivuna section MZ2 during a study of the minor-element compositions of matrix olivine and pyroxene in types 1, 2, and 3 chondrites" (p.3); no rule was applied to choose it, and it is the section's only CAI | N — the paper names the phases it analysed (see `Sampling Unit Type`) but states no rule for choosing the individual units | Identification by EDS, then by CL — "Apatite grains in OREX-803079-0 and OREX-803080-0 were identified via EDS mapping and point analysis", after which "CL images were obtained for each apatite grain to search for zoning or internal structures not resolvable in EDS maps" (p.2) | Spatial position within the shock assemblage — garnet is analysed "within the eclogitic mineral assemblage of zoned veins" and contrasted with grains "in either thin melt veins or the edge zones of zoned melt veins" (p.4); the individual grains are not otherwise chosen by a stated rule | N — the paper names the phases it analysed (see `Sampling Unit Type`) but states no rule for choosing the individual units | N — the paper names the phases it analysed (see `Sampling Unit Type`) but states no rule for choosing the individual units | N — beam conditions are given per phase ("Quantitative analyses of silicates, sulfides and oxides were run using a focused beam"; "A 2-μm defocused beam size, lower beam currents and shorter count times were used for phosphate and carbonate analyses", p.9), but that is a setting per phase, not a rule for choosing which grains were analysed | Size — "Aggregate particles (<1 mm) were mounted in epoxy, polished and were subsequently carbon coated" (p.11); the grains analysed within them are not otherwise chosen by a stated rule | Phase identity — "Olivine and pyroxene grains were identified and characterized at the NHM" (p.13), the paper's target being the anhydrous silicates in particles P1 and P2 | N - continuous thin sections of the whole core; five stage maps per section, no selection rule stated
Pre-Analysis Imaging and Screening | Imaging or other characterisation performed before the measurement in order to select or locate the sampling unit to be analysed, including the technique, instrument and settings used, and how individual analyses are linked back to the images. Distinct from any imaging the procedure performs as its own measurement. Where the imaging is performed on a separate instrument, it should also be recorded in the Group 1 coupling fields. | Advanced | Editable | Text (free) |  | Source: Sampling Unit Selection module | 2026-09-01 | sample |  | Y | Y | Y | Y |  | SEM BSE imaging on a ZEISS 1550VP field-emission SEM, which locates the occurrences the probe then analyses — "SEM BSE image showing tissintite in a shock melt pocket, in Tissint section UT2" (Fig. 1 caption, p.2). The paper lists EPMA, SEM, EBSD, synchrotron XRD and micro-Raman as one suite (p.2) without stating the order | SEM imaging and EDS mapping on three instruments before the probe — "Scanning electron microscopy (SEM) imaging and energy dispersive X-ray spectroscopy (EDS) mapping were conducted" on a Nova NanoSEM 450 at IGGCAS, a SUPRA55 at NAOC and a JEOL JSM-7100F at NIPR, after which "Quantitative analyses ... were conducted by electron probe microanalysis (EPMA) with the JEOL JXA-8100 at IGGCAS" (p.2) | Petrographic microscopy and SEM — "The petrography of these sections was examined using a petrographic microscope and a scanning electron microscope" before the microprobe work (p.3) | Petrographic microscopy and SEM — "The petrography of these sections was examined using a petrographic microscope and a scanning electron microscope" (p.3); the same sections then went to both microprobes | SEM BSE imaging on a ZEISS 1550VP field-emission SEM — "Backscattered electron (BSE) imaging was performed using a Carl Zeiss, LLC 1550VP field emission SEM" (p.2), locating the three occurrences of the new mineral in the Zagami thin section (Fig. 1, p.2). The paper lists EPMA, SEM, EBSD, synchrotron XRD and micro-Raman as one suite (p.2) without stating the order | Petrographic microscopy and SEM — "The CAI was characterized by petrographic microscope, scanning electron microscope, electron microprobe, and X-ray mapping before being measured for oxygen isotopes" (p.3); the object itself had been found during an earlier survey of matrix compositions (p.3) | Optical microscopy of the same thin section — "Eleven OC002 LAB24-2 fragments were mounted and dry-polished in a petrographic thin section used for optical microscopy and electron probe microanalyses" (p.3) | SEM EDS mapping, then CL imaging, on the JEOL 7900F at JSC — "Apatite grains in OREX-803079-0 and OREX-803080-0 were identified via EDS mapping and point analysis", after which "CL images were obtained for each apatite grain to search for zoning or internal structures not resolvable in EDS maps", collected "at 5 kV with beam currents ranging from 1 to 1.5 nA" (p.2); the numbered grains (Ap. #1 …) tie the probe analyses back to those images | SEM petrography — "The petrographic texture of NWA 8003 was observed using a JEOL 7000F field emission gun scanning electron microscope (FEG-SEM) at Hokkaido University" (p.7); phase identifications also rest on Raman spectra and EBSD patterns (p.4) | SEM mapping at the Smithsonian — particles 'analysed at 15 kV and around 0.5 nA in high vacuum using a Thermo Fisher Quattro FE-SEM'; 'Maps of loose grains were investigated, then used to inform sectioning' (p.7) | N — the paper describes SEM characterisation at JSC, the NHM, the Smithsonian and Curtin, not for the K-ALFAA microprobe work | SEM imaging and EDS mapping before the probe — particles were characterized by SE and BSE imaging and by EDS mapping of "The compositional heterogeneity of the particles" (p.9), and the paper's Fig. 1 pairs those BSE images with the EMPA data (p.2); the Methods list SEM before electron microprobe analysis without stating an explicit order | SEM imaging and multi-element EDS mapping — "SEM observations were performed on the samples using a JEOL JSM-6510 with 3-nA primary beam at 15 kV. We also performed multi-element EDS mapping (Mg, Si, Fe, Ni, S, Na, Ca and Al) of the different grains", after which "Quantitative chemical analyses were performed using a JEOL JXA-8230 electron microprobe" (p.11) | SEM characterisation at the NHM — "Olivine and pyroxene grains were identified and characterized at the NHM" and "Following characterization by SEM/EPMA, an additional carbon coat was added for a total thickness of ~30 nm" (p.13); additional quantitative data came from a Zeiss EVO 15LS analytical SEM with an Oxford X-Max80 EDS (p.13) | BSE mosaic: approximately 325 backscattered-electron images collected with the JEOL guide-net mapping software at 15 kV, 2 nA probe current and 70x magnification, stitched with the ImageJ Fiji grid-collection stitching plug-in (Donovan et al., 2021) into a 20k x 5k pixel mosaic at ~1.5 um pixel resolution
3. Instrument & Software | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Instrument Manufacturer | Manufacturer of the instrument that performs the measurement, recorded as a controlled value. Where a procedure couples a sample-introduction system to an analysing instrument, this records the analysing instrument. Instrument Model gives the specific designation. | Basic | Read-Only | Controlled list | JEOL ¦ Cameca ¦ Unknown ¦ N/A ¦ None | Source: Core module | 2026-08-30 | (none) | Recording it as a controlled value lets procedures be found by vendor. | Y | Y | Y | Y |  | JEOL | JEOL | Cameca | JEOL | JEOL | Cameca | JEOL | JEOL | JEOL | JEOL | Cameca | Cameca | JEOL | Cameca | JEOL
Instrument Model | Model designation of the instrument that performs the measurement, including any generation or configuration suffix. Conventionally written with the manufacturer name included; Instrument Manufacturer records the vendor separately, as a controlled value, so that procedures remain findable by vendor. | Basic | Read-Only | Text (free) | e.g., JXA-8530F Plus ¦ JXA-8230 ¦ JXA-8200 ¦ JXA-8100 ¦ SXFiveTactis ¦ SXFive ¦ SX100 ¦ iHP200F ¦ Unknown ¦ N/A ¦ None | Source: Core module | 2026-08-31 | (none) |  | Y | Y | Y | Y |  | JEOL 8200 (stated as "JEOL 8200 electron microprobe"; no JXA prefix stated) | JXA-8100 (stated as "JEOL JXA-8100") | SX100 (stated as "Cameca SX100") | JXA-8200 (stated as "JEOL JXA-8200") | JEOL 8200 (stated as "JEOL 8200 electron microprobe"; no JXA prefix in paper) | SX100 (stated as "Cameca SX100") | JXA-8200 (stated as "JEOL JXA-8200 electron microprobe") | JEOL 8530 EMPA (stated as "JEOL 8530 EMPA at NASA JSC"; no F suffix or JXA prefix stated) | JXA-8100 (stated as "JEOL 8100") | JXA-8530F Plus (stated as "JEOL 8530 F+ Hyperprobe Field Emission Electron Probe Microanalyzer") | SX-100 (stated as "Cameca SX-100 electron microprobe located at K-ALFAA") | SX-100 Ultra (stated as "SX-100 Ultra electron microprobe in the K-ALFAA") | JXA-8230 (stated as "JEOL JXA-8230 electron microprobe analyser (EPMA)") | SX100 (stated as "CAMECA SX100 electron microprobe") | JEOL JXA-8200
Electron Source | Type of electron gun used in the instrument. | Basic | Read-Only | Controlled list | Cold-FEG ¦ Schottky FEG (X-FEG) ¦ Schottky FEG (standard) ¦ Field emission gun (FEG) — subtype not specified ¦ LaB6 / CeB6 ¦ Tungsten (W) ¦ Unknown ¦ N/A ¦ None |  | 2026-08-30 | (none) |  | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | Field Emission (FEG) | N | N | N | N | N
Acquisition Software | Instrument control, automation and data acquisition software used to collect the raw data, including version number. Distinct from Data Processing Software(s), which covers everything applied to the data after acquisition. | Basic | Editable | Text (free) | e.g., 'Probe for EPMA v12.9.5' ¦ 'JEOL PC-EPMA v3.0' ¦ 'Cameca PeakSight 6.2' | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y |  | Probe for EPMA (Probe Software, Inc.) | N | N | N | Probe for EPMA (Probe Software, Inc.) | N | Probe for EPMA microanalysis software | N | N | N | N | N | N | N | JEOL guide-net mapping software (BSE mosaic); N for the WDS stage maps
Data Processing Software(s) | All software applied to the data after acquisition in order to produce the reported quantities, including version numbers. List every package used. Distinct from Acquisition Software, which controls the instrument and collects the raw data. | Basic | Editable | Text (free) | e.g., 'Probe for EPMA (built-in)' ¦ 'CalcZAF v12' ¦ 'CITZAF (Armstrong 1995)' ¦ 'XMapTools 4.1 (for mapping)' | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y |  | N — the 'CITZAF correction procedure' is recorded under Matrix Correction Method; the only software named is Probe for EPMA | N — 'The Bence-Albee method was used' is recorded under Matrix Correction Method; no software is named | N | N | N — the 'CITZAF correction procedure' is recorded under Matrix Correction Method; the only software named is Probe for EPMA | N | Probe for EPMA (CITZAF matrix correction, Armstrong 1995); CalcImage and Quantitative Microanalysis Explorer web-based tool (for stage mapping) | N | N | N | N | XMapTools (for phase maps) | N | N | Probe Software CalcImage and Probe for EPMA (full Phi(rho-z) correction at each pixel); MATLAB routines generating 32-bit floating point .tiff quantitative maps; Fiji and MATLAB for stitching; ENVI input image stacks
WDS Spectrometer Configuration | Number, type, and crystal range of WDS spectrometers on the instrument. Include manufacturer, model, and crystal range. For SEM-WDS configurations (third-party WDS on a non-EPMA platform), include WDS manufacturer and model. | Advanced | Read-Only | Text (free) | e.g., '5 Rowland-circle WDS spectrometers (JEOL), crystal range LDE1-LIFH' ¦ '4 WDS spectrometers (Cameca)' |  | 2026-08-12 | (none) |  | N | N | Y | Y |  | N | N | N | N | N | N | 5 wavelength-dispersive spectrometers (JEOL) | N | N | N — the crystals (LIFL, TAPL, PETL) are recorded under Diffracting Crystal; the spectrometer configuration and the X-ray lines are not stated | N | N | 5 wavelength-dispersive spectrometers (JEOL) | N | Fixed wavelength-dispersive spectrometers; count not stated (five elements collected per pass)
EDS Detector Configuration | EDS detector type, manufacturer, number of detector elements, active area and solid angle, window type, and geometry (take-off angle, position). List multiple detectors separately. Record 'N/A' where the procedure has no EDS detector. | Advanced | Read-Only | Text (free) | e.g., 'Oxford Ultim Max 100 SDD, 100 mm2, polymer window, 35 deg take-off angle' ¦ 'JEOL (e2v/Gresham) SDD, Be window' |  | 2026-08-27 | (none) |  | Y | Y | N | N |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A — WDS mapping procedure
4. Measurement Information | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Analytical Mode | Primary analytical mode(s) executed under this procedure. Records one value for a single-mode procedure and all applicable modes for a multi-mode one. Applies even where a procedure defines no mode flag columns. | Basic | Read-Only | Controlled list | EDS Point Analysis ¦ EDS Mapping ¦ WDS Point Analysis ¦ WDS Mapping | Source: Core module | 2026-09-09 | (none) | Serves as the procedure-level declaration of measurement type, distinct from the mode flag columns, which state per-field applicability. Ensures every procedure record is self-describing and comparable across the library. | Y | Y | Y | Y |  | WDS Point Analysis — 'WDS: 15 kV; 5 nA; beam in focused mode'; point analyses only | N — quantitative point analyses ('Quantitative analyses of maskelynite ... were conducted by electron probe microanalysis'); WDS or EDS is not stated, and the list has no value without one | N — quantitative point analyses and elemental X-ray maps; WDS or EDS is not stated, and the list has no value without one | N — quantitative point analyses; WDS or EDS is not stated, and the list has no value without one | WDS Point Analysis — 'WDS: 15 kV, 5 nA'; the paper also mentions 'K-mapping by EPMA', with no detector or conditions | N — point analyses and X-ray mapping ('electron microprobe, and X-ray mapping'); WDS or EDS is not stated, and the list has no value without one | WDS Point Analysis; WDS Mapping — 'wavelength-dispersive quantitative compositional mapping and analysis' | N — 14 quantitative point analyses; WDS or EDS is not stated, and the list has no value without one | WDS Point Analysis — 'Electron Probe Micro-Analyzer (EPMA) with wavelength dispersive spectrometers (WDS)' | N — point analyses with named crystals (LIFL, TAPL, PETL); WDS or EDS is not stated, and the list has no value without one | WDS Point Analysis — 'Wavelength-dispersive X-ray spectroscopy analyses of Mg,Na phosphate' | N — quantitative point analyses and X-ray maps ('BSE images, element maps and quantitative compositional analyses'); WDS or EDS is not stated, and the list has no value without one | N — quantitative point analyses on an instrument 'equipped with five wavelength-dispersive spectrometers and one silicon drift detector energy dispersive spectrometer'; which detector measured which element is not stated | N — point analyses ('Analyses were performed at 20 kV, using a focused 1-μm beam'); WDS or EDS is not stated, and the list has no value without one | WDS Mapping — 'five EPMA stage maps were acquired using fixed wavelength-dispersive spectrometers (WDS)'
Beam Mode | Whether the electron beam was operated as a stationary focused spot, defocused to a specified diameter, or rastered over a small area during a single-point analysis. Must be consistent with the Beam Diameter and Beam Raster Dimensions fields. For mapping, beam scanning is controlled by Step Size / Pixel Size and Stage Scan vs. Beam Scan instead. | Basic | Read-Only | Controlled list / Text | Focused ¦ Defocused ¦ Rastered ¦ N/A ¦ None |  | 2026-09-28 | target material |  | Y | N | Y | N |  | all: Focused — 'WDS: 15 kV; 5 nA; beam in focused mode' (p.3) | N — only 15 kV and 10 nA are given | maskelynite, phosphate, sulfide, glass: Defocused; other: N — 'Maskelynite, phosphate, sulfide, and glass were analyzed using a defocused beam of 5–10 µm size'; for olivine, pyroxene and Fe-Ti-Cr oxides only a '1–2 µm beam diameter' is given, not a mode | maskelynite, phosphate, sulfide, glass: Defocused; other: N — 'Maskelynite, phosphate, sulfide, and glass were analyzed using a defocused beam of 5–10 µm size'; for olivine, pyroxene and Fe-Ti-Cr oxides only a '1–2 µm beam diameter' is given, not a mode | all: Focused — 'WDS: 15 kV, 5 nA, beam in focused mode' (p.2) | all: Focused — 'Analyses were performed at 20 kV and 20 nA using a focused beam of 1 μm' (p.3) | N | N — 'a 2μm probe size' is recorded under Beam Diameter; no mode is named | plagioclase and its polymorphs: Defocused; other: Focused — Methods | N — only spot sizes are given | N — only a '1-µm beam size' is given, for the phosphate analyses | silicates, sulfides, oxides: Focused; phosphates, carbonates: Defocused — 'Quantitative analyses of silicates, sulfides and oxides were run using a focused beam ... A 2-μm defocused beam size ... for phosphate and carbonate analyses' | carbonates: Rastered; other: N — 'For carbonates, we rastered the beam over 5×5 µm2' | all: Focused — 'Analyses were performed at 20 kV, using a focused 1-μm beam' | N/A — mapping-only procedure; map conditions under Mapping Beam Mode
Accelerating Voltage | Electron beam accelerating voltage in kilovolts (kV). Justify any deviation from the standard operating voltage. | Basic | Editable | Numeric (kV) | e.g., 15 ¦ 20 ¦ 10 ¦ 7.5 |  | 2026-08-27 | (none) |  | Y | Y | Y | Y |  | 15 kV | 15 kV | 15 kV | 15 kV | 15 kV | 20 kV | 15 kV | 15 kV | 15 kV | 15 kV | 15 kV | 15 kV | 20 kV | 20 kV | 15 kV (stage maps and BSE mosaic)
Beam Current | Probe current in nanoamperes (nA). Often varies by phase type or target species; record the procedure-standard value(s). | Basic | Editable | Numeric (nA) | e.g., 5 ¦ 10 ¦ 20 ¦ 50 ¦ 100 |  | 2026-09-28 | target material |  | Y | N | Y | N |  | all: 5 nA — p.3 | all: 10 nA — 'The accelerating voltage was 15 kV, and the beam current was 10 nA' | olivine, pyroxene, Fe-Ti-Cr oxides: 20 nA; maskelynite, phosphate, sulfide, glass: 10 nA — p.3 | olivine, pyroxene, Fe-Ti-Cr oxides: 20 nA; maskelynite, phosphate, sulfide, glass: 10 nA — p.3 | all: 5 nA — p.2 | all: 20 nA — p.3 | all: 25 nA — '15 kV accelerating potential and 25 nA probe current for point analysis' | all: 20 nA — '14 analyses were performed at 15kV, 20nA, using a 2μm probe size' | all: 20 nA — 'a focused beam of 20nA ... at the same beam current' | all: 10 nA — p.7 | "Mg,Na phosphate": 8 nA; other: N — carbonate conditions are not stated | silicates, sulfides, oxides: 20 nA; phosphates: 8 nA; carbonates: 4 nA — p.9 | all: 10 nA — 'an accelerating voltage of 20 kV, a probe current of 10 nA and beam diameter of 1 µm' | N | N/A — mapping-only procedure; map current under Mapping Beam Current; the 2 nA BSE mosaic is recorded under Pre-Analysis Imaging and Screening
Beam Diameter | Diameter of the electron beam in micrometers. 0 indicates a fully focused beam. Document defocused diameter when used to minimize beam damage or improve spatial averaging for beam-sensitive phases. | Basic | Editable | Numeric (µm) / Text | e.g., 0 (focused) ¦ 1 ¦ 2 ¦ 5 ¦ 10 ¦ 20 |  | 2026-09-28 | target material |  | Y | N | Y | N |  | N — 'beam in focused mode' is recorded under Beam Mode; no diameter is given | N — only 15 kV and 10 nA are given | olivine, pyroxene, Fe-Ti-Cr oxides: 1–2 µm; maskelynite, phosphate, sulfide, glass: 5–10 µm — p.3 | olivine, pyroxene, Fe-Ti-Cr oxides: 1–2 µm; maskelynite, phosphate, sulfide, glass: 5–10 µm — p.3 | N — 'beam in focused mode' is recorded under Beam Mode; no diameter is given | all: 1 µm — p.3 | N | all: 2 µm — as above | plagioclase and its polymorphs: 2–5 µm; other: N — 'a defocused beam (2–5 μm in diameter)' | carbonate: 5 µm; magnetite, olivine: 1 µm — 'with an analytical spot size of 5 µm' (carbonates); 'Analyses were conducted at 15kV and 10nA, with an analytical spot size of 1µm' (the magnetite and olivine sentence) | "Mg,Na phosphate": 1 µm; other: N — 'using a 1-µm beam size' | phosphates, carbonates: 2 µm; other: N — p.9 | all: 1 µm — as above | all: 1 µm — as above | N/A — mapping-only procedure; map beam under Mapping Beam Diameter
Beam Raster Dimensions | Dimensions of the small area over which the beam is rastered at a single analysis point, reported as width × height in µm. Applicable when Beam Mode = Rastered; defines the effective spatial footprint of the measurement. Not applicable when mapping. | Advanced | Editable | Numeric pair (µm x µm) | e.g., 5x5 ¦ 10x10 ¦ None |  | 2026-09-28 | target material | Averages over a coarse-grained or beam-sensitive phase by distributing dose across a larger area. | Y | N | Y | N |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | carbonates: 5 × 5 µm²; other: N/A — as above | N/A | N/A — stage scan, not beam scan
Beam Damage Minimization | Measures taken to minimize beam damage, particularly volatilization or migration of Na, K, F, and Cl in hydrous minerals, glasses, feldspars, phosphates, and carbonates. Document approach, beam conditions used, and phases for which it was applied. | Advanced | Editable | Text (free) | e.g., 'Na measured first with 10 um defocused beam at 5 nA' ¦ 'Time-resolved acquisition extrapolated to t=0 for Na and K' ¦ 'None required' |  | 2026-09-29 | target material |  | Y | N | Y | N |  | N | N | N — a defocused 5–10 µm beam at 10 nA is used for maskelynite, phosphate, sulfide and glass, but the paper gives no reason for it | N — a defocused 5–10 µm beam at 10 nA is used for maskelynite, phosphate, sulfide and glass, but the paper gives no reason for it | all: low beam current (5 nA) — 'Although a low beam current of 5 nA was used for the analysis, the low cation sum (0.92) for the K-site is likely due to diffusion of Na away from the electron beam' (p.3) | N | N | all: Durango apatite compared at 10 μm and 3 μm spot sizes, with no significant volatile loss — 'in order to assess volatilization of halogens using our beam conditions' | N | N | N | phosphates, carbonates: 2 µm defocused beam, lower beam currents and shorter count times; other: N — 'to minimize possible beam damage effects' | N | N | N/A
Mapping Beam Mode | Whether the electron beam was focused or defocused to a stated diameter during X-ray mapping. How the beam or stage moves across the mapped area is recorded by Stage Scan vs. Beam Scan and Step Size / Pixel Size. | Basic | Read-Only | Controlled list / Text | Focused ¦ Defocused ¦ N/A ¦ None |  | 2026-09-28 | sample > sampling unit |  | N | Y | N | Y |  | N/A | N/A | Focused (section maps and olivine megacryst maps) — 'a 20 nA focused beam'; 'a focused beam of 15 kV voltage and 200 nA current' | N/A | N/A | N | N | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Fixed 10 µm beam (stated 'a fixed 10 um electron beam')
Mapping Beam Current | Probe current in nanoamperes (nA) used during X-ray mapping. | Basic | Editable | Numeric (nA) | e.g., 50 ¦ 100 ¦ 200 |  | 2026-09-28 | sample > sampling unit |  | N | Y | N | Y |  | N/A | N/A | 20 nA (section maps: Ca, Al, Fe, Mg Kα); 200 nA (olivine megacryst maps) — p.3 | N/A | N/A | N | N | N/A | N/A | N/A | N/A | 20 nA — 'X-ray maps and BSE images were run at 15kV and 20nA' | N/A | N/A | 100 nA probe current (stage maps)
Mapping Beam Diameter | Diameter of the electron beam in micrometres during X-ray mapping. 0 indicates a fully focused beam. | Basic | Editable | Numeric (µm) / Text | e.g., 0 (focused) ¦ 1 ¦ 10 |  | 2026-09-28 | sample > sampling unit |  | N | Y | N | Y |  | N/A | N/A | N | N/A | N/A | N | N | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 10 µm (fixed)
Drift Correction | Method used to monitor and correct for instrument drift (beam current drift, spectrometer drift) during the analytical session. | Advanced | Editable | Text (free) | e.g., 'Primary standards re-measured at start, midpoint, and end of session; calibration linearly interpolated' ¦ 'Faraday cup beam current verified every 30 min' ¦ 'None' |  | 2026-05-11 | (none) |  | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Target Species | The chemical species this procedure is designed to determine, recorded at whatever resolution the chemistry is resolved — element(s) for this technique; valence species where a procedure resolves oxidation state; compounds where it resolves molecules. Isotopes are not target species: isotopes of an element are the same chemical species. The subset actually measured may be narrower than the registered suite where species were excluded due to interferences or scope reduction. | Basic | Editable | Text (free) | e.g., 'Si, Ti, Al, Fe, Mn, Mg, Ca, Na, K, Cr, Ni, P, S, Cl, F' ¦ 'Si, Al, Fe, Mg, Ca, Na, K' | Source: Target Species module | 2026-09-01 | defines: target species |  | Y | Y | Y | Y |  | Si, Ti, Al, Cr, Fe, Mn, Mg, Ca, Na, K — from the standards sentence (p.3) | Si, Ti, Al, Cr, Fe, Mn, Mg, Ca, Na, K | Si, Ti, Al, Mg, Ca, Fe, Mn, Cr, Ni, Na, K, P — read from the detection-limit sentence, which names their oxides (p.3) | Si, Ti, Al, Mg, Ca, Fe, Mn, Cr, Ni, Na, K, P — read from the detection-limit sentence, which names their oxides (p.3) | Si, Al, K, Na, Ca, Fe, Mg, Ti, Cr, Mn (stated via standards sentence and detection limit list) | Si, Al, Ti, K, Na, Fe, Mg, Ca, S, Mn, Cr, Ni, P, V | F, O, CO2, H2O — F measured; 'Oxygen was calculated by elemental stoichiometry with water estimated by inspection of the analytical total, and for carbonates, CO2 was calculated by stoichiometry'; the other analysed elements are not named | P, F, Cl, Ca, Mn, Fe, Na, Mg, Si, S, OH — the ten elements analysed (p.3); OH 'calculated by difference based on 1–F–Cl=OH' (p.4) | N — no element is named in the paper; the analysed elements are in Supplementary Table 4, which is not in the archived PDF | Fe, Mn, Mg, Ca — the carbonate analyses, named with their crystals; the elements of the magnetite and olivine analyses are not named | F, P, Ca, Si, Mg, Fe, Al, S, K, Cl, Na, Mn — phosphate standards name F, P, Ca, Si, Mg, Fe, Al, S, K, Cl; carbonate standards name Na, Si, Mg, Ca, Mn, P, S, Fe | N | Al, Ti, Ca, Cr, Mn, Ni, Mg, Fe, Si, Na, K — '(1) Al, Ti, Ca, Cr, Mn, Ni, Mg, Fe and Si (session 1) and (2) Na, K, Al, Ti, Ca, Cr, Mn, Ni, Mg, Fe and Si (session 2)' (p.11) | N | Mg, Al, Fe, Ca, Ti, Na, Si, Mn, K, Cr (ten elements, collected in two passes)
Monitored Elements | Specific elements monitored in this procedure, grouped by the target species they serve where they serve one. Includes elements monitored only to correct an interference, which serve no target species and so have no parent. The target species list is given by the Target Species field and is never inferred from the elements appearing here. A target species determined by stoichiometry or by difference, rather than measured, has no monitored element. The X-ray line, diffracting crystal, spectrometer assignment and counting times used for each monitored element are recorded in their own fields, keyed to this one. | Basic | Read-Only | Text (free) | e.g., 'Si, Al, K, Na, Ca, Fe, Mg, Ti, Cr, Mn' ¦ 'Fe, Mg, Si, Ca (determined); Cr (monitored to correct the Kbeta overlap on Mn)' ¦ N/A ¦ None |  | 2026-09-28 | defines: monitored property per target species |  | Y | Y | Y | Y |  | Si, Al, Ca, Na, Fe, Mg, Mn, Ti, Cr, K — all determined; no monitor-only element. "Standards for analysis were anorthite (SiKα, AlKα, CaKα); albite (NaKα); fayalite (FeKα); forsterite (MgKα); Mn2SiO4 (MnKα); TiO2 (TiKα); Cr2O3 (CrKα); and microcline (KKα)" (p.3) | Si, Mg, Fe, Na, Al, Ca, Mn, K, Ti, Cr — all determined. "The EPMA standards were natural and synthetic minerals: natural kaersutite for Si, Mg and Fe, jadeite for Na and Al, bustamite for Ca and Mn, and K-feldspar for K, synthetic rutile for Ti and Cr2O3 for Cr". Cr also carries the interference correction — "X-ray interference of the Kα line of Mn by the Kβ line of Cr was corrected" — but is itself determined, so it is not an orphan | Si, Ti, Al, Mg, Ca, Fe, Mn, Cr, Ni, Na, K, P — point analyses on both instruments (the detection-limit sentence names their oxides); maps: 'Ca Ka, Al Ka, Fe Ka, and Mg Ka' (four sections) and 'Fe Ka, P Ka, Al Ka, Ca Ka, and Cr Ka' (two olivine megacrysts, collected on 'the EMP' — the instrument is not named) | Si, Ti, Al, Mg, Ca, Fe, Mn, Cr, Ni, Na, K, P — all determined. "Detection limits are typically <0.03 wt% for SiO2, TiO2, Al2O3, MgO, and CaO; <0.05–0.1 wt% for FeO, MnO, Cr2O3, NiO, Na2O, K2O, and P2O5" (p.4) | Si, Al, K, Ca, Na, Fe, Mg, Ti, Cr, Mn — all determined. "Standards for analysis were Asbestos microcline (SiKa, AlKa, KKa), synthetic anorthite (CaKa), Amelia albite (NaKa), synthetic fayalite (FeKa), synthetic forsterite (MgKa), synthetic TiO2 (TiKa), synthetic Cr2O3 (CrKa), and synthetic Mn-olivine (MnKa)", with detection limits quoted for the same ten (p.2) | Si, Al, Ti, K, Na, Fe, Mg, Ca, S, Mn, Cr, Ni, P, V — all determined. "Standards were Kakanui kaersutite for silicon, aluminum, titanium, potassium, sodium, iron, magnesium, and calcium, Canyon Diablo troilite for sulfur, rhodonite for manganese, chromium metal for chromium, nickel metal for nickel, apatite for phosphorus, and vanadium metal for vanadium" (pp.3–4) | F — 'Fluorine measurement was made using the LDE1 diffracting crystal'; no other element is named | P, F, Cl, Ca, Mn, Fe, Na, Mg, Si, S — all determined. "A total of 14 analyses were performed at 15 kV, 20 nA, using a 2 μm probe size, and included the elements P, F, Cl, Ca, Mn, Fe, Na, Mg, Si, and S" (p.3) | N — "Natural and synthetic standards were used" (p.7) without naming them; the analysed elements are given in Supplementary Table 4, which is not in the archived PDF | Fe, Mn, Mg, Ca — all determined. "Carbonate analyses were run at 15 kV and 10 nA, with an analytical spot size of 5 µm. Fe and Mn were analysed using a LIFL crystal, Mg using a TAPL crystal and Ca using a PETL crystal" (p.7) | F, P, Ca, Si, Mg, Fe, Al, S, K, Cl, Na, Mn — as for Target Species | N — "Well-characterized natural and synthetic materials were used as standards" (p.9); the paper gives beam conditions and count times for silicates, sulfides, oxides, phosphates and carbonates but names no element | Al, Ti, Ca, Cr, Mn, Ni, Mg, Fe, Si, Na, K — as for Target Species | N — the paper gives beam conditions and a detection limit for transition metals of about 250 ppm for the NHM London instrument (p.13) but names no element | Mg, Al, Fe, Ca, Ti, Na, Si, Mn, K, Cr — 'Two passes were used to collect X-ray intensities for Mg, Al, Fe, Ca, and Ti in pass 1, and Na, Si, Mn, K, and Cr in pass 2' (p.6)
Reported Variables and Units | The final variable(s) this procedure reports and their units — distinct from the fields recording what was *acquired* rather than what is reported. Record every reported variable, including intermediate quantities reported alongside final ones (e.g. both the 206Pb/238U ratio and the 206Pb/238U date). Where a reported variable is a nominal property with no magnitude (e.g. a mineral identification), record the variable and give the unit as 'N/A — nominal property'. | Basic | Read-Only | Text (free) | e.g., 'Element concentration (ppm); oxide (wt%)' ¦ 'd56Fe (permil vs IRMM-014); d57Fe (permil vs IRMM-014)' ¦ '206Pb/238U date (Ma); 207Pb/206Pb date (Ma)' ¦ 'Mineral identity (N/A — nominal property); match score (0-1)' | Source: Core module | 2026-08-25 | defines: reported property | A procedure may acquire many channels and report a small number of derived quantities; without this field a data consumer cannot tell which. | Y | Y | Y | Y |  | SiO2, TiO2, Al2O3, FeO, MgO, CaO, Na2O, K2O, Cr2O3, MnO; cations per formula unit; Ca/(Ca+Na+K); Ca-Eskola component; An content of the precursor plagioclase — oxide wt% with totals and 1 s.d. of the mean, cations on 6 or 8 oxygens (Table 1); Ca-Eskola (mol%) and An58–69 (p.1) | SiO2, TiO2, Al2O3, Cr2O3, FeO, MnO, MgO, CaO, Na2O, K2O — oxide wt% for maskelynite, melt inclusion glasses, silica glasses, coesite aggregates and mesostasis (p.2) | SiO2, TiO2, Al2O3, Cr2O3, FeO, MnO, MgO, CaO, Na2O, NiO, P2O5, K2O, V2O3, La2O3, Ce2O3; modal-area fraction per mineral (vol%); Mg#; En, Fs, Wo — oxide wt% 'of selected minerals' (Table 2), with V2O3, La2O3 and Ce2O3 for the oxides and merrillite; glass means with 1σ (Table 3); modal abundances from the X-ray maps (Table 1) | SiO2, TiO2, Al2O3, Cr2O3, FeO, MnO, MgO, CaO, Na2O, NiO, P2O5, K2O, V2O3, La2O3, Ce2O3; modal-area fraction per mineral (vol%); Mg#; En, Fs, Wo — oxide wt% 'of selected minerals' (Table 2), with V2O3, La2O3 and Ce2O3 for the oxides and merrillite; glass means with 1σ (Table 3); modal abundances from the X-ray maps (Table 1) | SiO2, TiO2, Al2O3, FeO, CaO, Na2O, K2O; cations per formula unit; density — oxide wt% with totals and 1 s.d. of the mean, cations on 8 oxygens (Table 1); 'Magnesium, Cr, and Mn were also analyzed but were below the detection limit in all cases'; calculated density 3.98 g cm-3 (p.3) | SO2, P2O5, Na2O, K2O, MgO, Al2O3, SiO2, CaO, TiO2, V2O3, FeO, Cr2O3, MnO, NiO; åkermanite content — oxide wt% with totals (Table 1); Åk14–31 (p.5) | Element concentrations in the carbonates (wt%: Fe, Mn — "Dolomite contains 2.0 ± 0.4 wt% Fe and 3.0 ± 0.9 wt% Mn (n = 37)", p.5); phase abundance as areal fraction of the section (areal%, e.g. sulfides "2.3 areal%", p.6); phase identifications from the X-ray maps (nominal) | F, Cl, Na2O, MgO, SiO2, SO3, P2O5, CaO, MnO, FeO — wt% with totals, per named grain (Table 1) | En, Fs, Wo; Ca-Eskola component; empirical formulae — mol% (p.2, p.4); the oxide analyses are in Supplementary Tables 1–4, not in the archived PDF | Carbonate end-member composition (mol%: MgCO3, FeCO3, MnCO3 — calcite is "near-end member composition (4 mol.% or less MgCO3 and FeCO3; 0.1 mol.% or less MnCO3)", p.2); phase identifications (nominal) | N — phosphate and carbonate compositions are reported by phase (p.2); the quantitative analyses are in the supplementary tables, not in the archived PDF | Fe + Co, S, Ni; modal abundance of carbonates, sulfides and magnetite — sulfide compositions in at% (Fig. 1); modal abundances '0.4–3.4%, ~3–8% and ~3–5%' from the EMPA phase maps (p.2) | Al, Ti, Ca, Cr, Mn, Ni, Mg, Fe, Si, Na, K — mineral compositions for the two element suites; values 'compiled in Supplementary Table 14', not in the archived PDF | major and minor element abundances; Mg# — 'Major and minor element abundances' of olivine and pyroxene; 'the Mg# of olivine grains is >83' (p.13) | element wt% maps; oxide wt% maps; cation stoichiometry; mineral endmember maps — 32-bit floating point .tiff
X-ray Detection Method per Monitored Element | The X-ray detection method used to measure the monitored element: wavelength-dispersive (WDS), in which a crystal spectrometer separates the X-rays by wavelength and a proportional counter counts them, or energy-dispersive (EDS), in which a solid-state detector sorts every photon by energy at once. Applies where a procedure uses both. | Basic | Read-Only | Controlled list | WDS ¦ EDS ¦ N/A ¦ None |  | 2026-09-28 | monitored property |  | Y | Y | Y | Y |  | all: WDS — 'WDS: 15 kV; 5 nA' | N | N | N | all: WDS — 'WDS: 15 kV, 5 nA' | N | all: WDS — 'wavelength-dispersive quantitative compositional mapping and analysis' | N | all: WDS — 'JEOL8100 ... with wavelength dispersive spectrometers (WDS)' | N | F, P, Ca, Si, Mg, Fe, Al, S, K, Cl: WDS; other: N — 'Wavelength-dispersive X-ray spectroscopy analyses of Mg,Na phosphate'; stated for the phosphate analyses only | N | N | N | all: WDS — p.6
X-ray Line | X-ray emission line measured. | Basic | Read-Only | Controlled list | Ka ¦ Kb ¦ La ¦ Lb ¦ Lg ¦ Ma ¦ Mb ¦ N/A ¦ None |  | 2026-09-10 | monitored property | Line choice affects sensitivity, matrix correction accuracy, and susceptibility to peak overlap and spectral interference. | Y | Y | Y | Y |  | Si, Al, Ca, Na, Fe, Mg, Mn, Ti, Cr, K: Ka — 'anorthite (SiKα, AlKα, CaKα); albite (NaKα); fayalite (FeKα); forsterite (MgKα); Mn2SiO4 (MnKα); TiO2 (TiKα); Cr2O3 (CrKα); and microcline (KKα)' | Mn: Ka; other: N — 'X-ray interference of the Ka line of Mn by the Kb line of Cr was corrected' | Ca, Al, Fe, Mg, P, Cr: Ka; other: N — maps: 'Ca Ka, Al Ka, Fe Ka, and Mg Ka' (four sections) and 'Fe Ka, P Ka, Al Ka, Ca Ka, and Cr Ka' (two olivine megacrysts); no line is stated for the point analyses | N | Si, Al, K, Ca, Na, Fe, Mg, Ti, Cr, Mn: Ka — 'Asbestos microcline (SiKa, AlKa, KKa), synthetic anorthite (CaKa), Amelia albite (NaKa), synthetic fayalite (FeKa), synthetic forsterite (MgKa), synthetic TiO2 (TiKa), synthetic Cr2O3 (CrKa), and synthetic Mn-olivine (MnKa)' | N | N | N | N | N | N | N | N | N | N
Diffracting Crystal | Analyzing crystal (monochromator). | Basic | Read-Only | Controlled list | LIF ¦ LLIF ¦ LIFL ¦ LIFH ¦ PET ¦ LPET ¦ PETL ¦ PETH ¦ PETJ ¦ TAP ¦ TAPL ¦ TAPH ¦ LTAP ¦ LDE1 ¦ LDE2 ¦ LDE3 ¦ RAP ¦ KAP ¦ ADP ¦ Unknown ¦ N/A ¦ None |  | 2026-09-10 | monitored property | Crystal choice determines the detectable wavelength range and dispersion. | N | N | Y | Y |  | N | N | N | N | N | N | F: LDE1 — p.3 | N | N | Fe, Mn: LIFL; Mg: TAPL; Ca: PETL — p.7 | N | N | N | N | N
WDS Spectrometer Channel | WDS spectrometer position(s) assigned to each target species, one entry per assignment. An target species may be assigned to more than one spectrometer with intensities aggregated (aggregate intensity counting), and one spectrometer serves several target species across a run, so the assignment — not the target species — is the unit carrying the spectrometer setup. | Advanced | Read-Only | Text (free) | e.g., 'Sp1' ¦ 'Sp2+Sp3 (aggregate intensity counting)' ¦ 'Sp4' |  | 2026-09-29 | monitored property | Different spectrometers may have different crystal configurations. | N | N | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N — spectrometer-to-element assignments are not stated
Sequence | Order in which spectrometer assignments are acquired, and — where the element suite exceeds the number of spectrometers — the passes the acquisition is divided into. Within a single pass all assigned spectrometers collect simultaneously, including at every pixel in X-ray mapping; a suite larger than the spectrometer count therefore requires the acquisition to be run more than once, each pass covering a different subset of elements. | Advanced | Read-Only | Integer | e.g., 1 ¦ 2 ¦ 3 ¦ 4 |  | 2026-09-10 | monitored property | Relevant for minimizing beam damage (volatile elements measured first) and for sequential multi-channel setups. | N | N | Y | N |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | Mg, Al, Fe, Ca, Ti: pass 1; Na, Si, Mn, K, Cr: pass 2 — 'Two passes were used to collect X-ray intensities for Mg, Al, Fe, Ca, and Ti in pass 1, and Na, Si, Mn, K, and Cr in pass 2'
Proportional Counter / Detector | Type of detector used. | Advanced | Read-Only | Controlled list | High Pressure P10 Flow ¦ Low Pressure P10 Flow ¦ Xe Sealed ¦ SDD ¦ Si(Li) ¦ Unknown ¦ N/A ¦ None |  | 2026-09-10 | monitored property | Affects sensitivity and count rate linearity. | N | N | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
WDS PHA Setting | Pulse height analyzer (PHA) setting for the WDS detector. Integral mode accepts all pulses above a threshold; Differential mode selects a narrow energy window to reject higher-order reflections and escape peaks. | Advanced | Read-Only | Controlled list | Integral ¦ Differential ¦ Unknown ¦ N/A ¦ None |  | 2026-09-10 | monitored property |  | N | N | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Peak Counting Time | Time spent counting X-ray intensity at the peak position, in seconds. Adjustments stay within procedure-defined bounds. | Basic | Editable | Numeric (s) | e.g., 10 ¦ 20 ¦ 30 ¦ 60 |  | 2026-10-01 | target material x monitored property |  | N | N | Y | N |  | N | N | N | N — Liu 2016 states no counting times; the earlier values matched Ma+2017 (same laboratory) and had no source in this paper | all [all: 20 s] — 'Counting times were 20 s on-peak and 10 s each on high and low background positions' (p.2) | all [all: 10–50 s] — 'Peak count times were 10–50 s' — a range; per-element times not given | N | N | N | N | N | silicates, sulfides, oxides [all: 20 s]; phosphates [all: 20 s]; carbonates [all: 10 s] — 'Quantitative analyses of silicates, sulfides and oxides were run using a focused beam at 15 kV, 20 nA, 20 s peak time and 10 s on each background ... Phosphate analyses were performed at 15 kV, 8 nA, 20 s peak and 10 s background; carbonates at 15 kV, 4 nA, 10 s peak and 5 s background' (Methods); no element is named | all [Al, Ti, Ca, Mn, Cr: 200 ms total (peak + background); Mg, Fe, Si: 20 ms total (peak + background); other: N] — 'The total peak + background counting time was 200 ms for Al, Ti, Ca, Mn and Cr, and 20 ms for Mg, Fe and Si' | N | N/A — mapping-only procedure; the 25 msec per-pixel dwell is under Dwell Time per Pixel
Background Counting Time | Total time spent counting at off-peak background position(s) in seconds, summed across all background positions. | Basic | Editable | Numeric (s) | e.g., 5 ¦ 10 ¦ 20 |  | 2026-10-01 | target material x monitored property |  | N | N | Y | N |  | N | N | N | N — as for Peak Counting Time | all [all: 10 s on each of the high and low background positions] — as above | N | N/A | N | N | N | N | silicates, sulfides, oxides [all: 10 s on each background]; phosphates [all: 10 s background]; carbonates [all: 5 s background] — 'Quantitative analyses of silicates, sulfides and oxides were run using a focused beam at 15 kV, 20 nA, 20 s peak time and 10 s on each background ... Phosphate analyses were performed at 15 kV, 8 nA, 20 s peak and 10 s background; carbonates at 15 kV, 4 nA, 10 s peak and 5 s background' (Methods); no element is named | N — only the total peak + background time is stated (see Peak Counting Time) | N | N/A — MAN background calibration; no off-peak counting
Background Position(s) | Location(s) of off-peak background measurement(s) relative to the peak, in mm or sin-theta, and whether on the high- or low-energy side. | Advanced | Editable | Text (free) | e.g., '+5 mm (High), -5 mm (Low)' ¦ 'High only (+4 mm)' ¦ 'MAN method (no off-peak positions)' |  | 2026-09-10 | monitored property |  | N | N | Y | N |  | N | N | N | N — as for Peak Counting Time | all: high and low — 'high and low background positions'; offsets not given | N | N/A | N | N | N | N | N — '10 s on each background' implies two positions, which are not given | N | N | N/A — MAN background calibration; no off-peak positions
EDS Live Time per Point or Pixel | EDS spectral acquisition live time per analysis point or per pixel in seconds. | Basic | Editable | Numeric (s) | e.g., 30 ¦ 60 ¦ 120 ¦ 0.005 (mapping) |  | 2026-10-29 | (none) |  | Y | Y | N | N |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
EDS Acquisition Mode | Spatial acquisition sub-strategy for EDS measurements: stationary-beam point acquisition, line scan (beam stepped along a transect at defined intervals), or area map / spectrum image (beam rastered over a pixel grid). Specifies how the beam is positioned during data collection within the declared analytical mode. Record 'N/A' where the procedure has no EDS detector. 'Point' covers what the literature also calls spot or point-spectrum analysis. 'Map' and 'Spectrum image' are distinct acquisitions, not synonyms: a map may retain element intensities alone, whereas a spectrum image retains a full spectrum at every pixel and can be requantified afterwards — record which was acquired. Where more than one mode was used, join them with '; ' rather than looking for a combined member. | Basic | Read-Only | Controlled list | Point ¦ Line scan ¦ Map ¦ Spectrum image ¦ N/A ¦ None |  | 2026-08-31 | (none) | Particularly important when a procedure includes linescans as a distinct acquisition approach not fully captured by the mode flag columns. | Y | Y | N | N |  | N/A — WDS procedure | N — WDS or EDS is not stated | N — WDS or EDS is not stated | N — WDS or EDS is not stated | N/A — WDS procedure | N — WDS or EDS is not stated | N/A — WDS procedure | N — WDS or EDS is not stated | N/A — WDS procedure | N — WDS or EDS is not stated | N/A — WDS procedure | N — WDS or EDS is not stated | N — WDS or EDS is not stated | N — WDS or EDS is not stated | N/A
Dwell Time per Pixel | Time the electron beam dwells on each pixel (step position) during X-ray mapping, in milliseconds. For WDS: one value per spectrometer assignment per pixel. For EDS: a single value, the real (clock) time per pixel; the detector live time, which excludes dead time, is recorded in EDS Live Time per Point or Pixel. | Basic | Editable | Numeric + unit | e.g., '10 ms' ¦ '100 ms' ¦ '0.5 s' |  | 2026-10-29 | monitored property |  | N | Y | N | Y |  | N/A | N/A | Fe, P, Al, Ca, Cr: ~0.5 s; other: N — 'a dwell time of ~0.5 s at each step' for the olivine megacryst maps; the section-map dwell time is not stated | N/A | N | N | N | N/A | N/A | N/A | N/A | N | N/A | N/A | all: 25 msec — the MAN background correction 'allows all map collection time to be dedicated to on-peak X-ray measurement'
Step Size / Pixel Size | Distance between adjacent measurement points in the X-ray map in micrometers, defining the spatial resolution. Report both X and Y step if they differ. | Basic | Editable | Numeric (µm) / Text | e.g., 0.5 ¦ 1 ¦ 2 ¦ 5 ¦ 10; or '1 um (X) x 2 um (Y)' |  | 2026-08-26 | (none) |  | N | Y | N | Y |  | N/A | N/A | 8-12 µm (BSE + Ca/Al/Fe/Mg Ka phase maps at UT); 2 µm (olivine megacryst Ka maps; instrument ambiguous) | N/A | N | N | N | N/A | N/A | N/A | N/A | N | N/A | N/A | 9.5 um (stage maps); ~1.5 um per pixel (BSE mosaic)
Map Dimensions | Number of pixels in the X-ray map in the X and Y directions. Based on the area of interest and selected step size. | N/A | Basic | Numeric pair (pixels x pixels) | e.g., 256x256 ¦ 512x512 ¦ 1024x768 |  | 2026-08-27 | (none) |  | N | Y | N | Y |  | N/A | N/A | N | N/A | N | N | N | N/A | N/A | N/A | N/A | N | N/A | N/A | 1,024 x 1,024 pixels per stage map; five stage maps per thin section
Map Area | Physical extent of the mapped region, given either as width × height in µm or as a total area in µm² or mm², and equal to (map width in pixels × step size) × (map height in pixels × step size). Complements the map's pixel-grid dimensions by recording the physical scale of the mapped region. | N/A | Basic | Numeric + unit | e.g., '100 × 100 µm' ¦ '500 × 300 µm' ¦ '1000 × 500 µm' ¦ '0.2 mm²' |  | 2026-08-27 | (none) | Based on the sample feature or region of interest, so it is chosen to cover the target. | N | Y | N | Y |  | N/A | N/A | N | N/A | N | N | N | N/A | N/A | N/A | N/A | N | N/A | N/A | N - not stated directly (1,024 pixels at 9.5 um step corresponds to ~9.7 mm per side)
Stage Scan vs. Beam Scan | For mapping modes, whether the map was acquired by moving the stage while the beam is held fixed (stage scan), or by deflecting the beam across the field while the stage is stationary (beam scan). | Advanced | Read-Only | Controlled list | Stage scan ¦ Beam scan ¦ Combined stage and beam scan ¦ Unknown ¦ N/A ¦ None |  | 2026-08-30 | (none) | Stage scan preserves beam geometry and spot size across the full map but is slower; beam scan is faster but may introduce geometric distortion at large deflections. | N | Y | N | Y |  | N/A | N/A | N | N/A | N | N | Stage scan | N/A | N/A | N/A | N/A | N | N/A | N/A | Stage scan (stated 'EPMA stage maps')
5. Data Processing | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Matrix Correction Method | X-ray matrix correction algorithm applied during quantitative EDS or WDS data reduction. For X-ray mapping, applies when raw count maps are converted to quantitative concentration maps. Where the k-factors or calibration constants themselves came from — measured standards, a vendor library, or theoretical cross-sections — is a separate question answered by this technique's calibration-standard field, not here; a procedure may be both absorption-corrected and standardless. | Basic | Read-Only | Controlled list / Text | PAP (Pouchou & Pichoir Full) ¦ XPP (Simplified PAP) ¦ PhiRhoZ Bastin (EPQ-91) ¦ Love-Scott I ¦ Love-Scott II ¦ Armstrong / Love-Scott ¦ ZAF ¦ CITZAF (Armstrong 1995) ¦ Bence-Albee ¦ Unknown ¦ N/A ¦ None |  | 2026-08-30 | (none) |  | Y | Y | Y | Y |  | CITZAF (Armstrong 1995) | Bence-Albee | N | N | CITZAF (Armstrong 1995) | N | CITZAF (Armstrong 1995) | N | ZAF | N | N | N | N | N | Full Phi(rho-z) correction applied at each pixel, of the form C = k x ZAF, where ZAF is the compositionally dependent correction for atomic number, X-ray absorption and characteristic fluorescence in both sample and standard
Mass Absorption Coefficients (MACs) | Database of mass absorption coefficients used in the matrix correction. | Basic | Read-Only | Controlled list | LINEMU ¦ CITZMU ¦ MCMASTER ¦ MAC30 ¦ MACJTA ¦ FFAST ¦ Unknown ¦ N/A ¦ None |  | 2026-08-30 | (none) | MAC database choice affects accuracy particularly for light elements (B, C, N, O, F, Na). | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
X-ray Background Correction Method | Method used to estimate and subtract background X-ray intensity beneath the peak. For WDS: typically 2-point off-peak linear interpolation or Mean Atomic Number (MAN) background model. For EDS: spectral background fitting or top-hat filter applied during spectral processing. | Basic | Editable | Controlled list | 2-point off-peak linear ¦ 2-point off-peak exponential ¦ 2-point off-peak polynomial ¦ 2-point off-peak (interpolation not specified) ¦ 1-point high with slope factor ¦ 1-point low with slope factor ¦ Mean Atomic Number (MAN) ¦ EDS background fit ¦ Top-hat filter ¦ N/A ¦ None |  | 2026-09-10 | monitored property |  | Y | Y | Y | Y |  | N | N | N | N — no background method is named | N | N | F: polynomial fit to the background determined on representative phases; other: Mean Atomic Number (MAN) — 'the mean atomic number (MAN) background calibration was made using these standards' | N | N | N | N | N — no background method is named | N | N | all: Mean Atomic Number (MAN) — MAN background calibration, using 'EPMA standards having a range of average atomic number Z'
Time-Dependent Intensity Correction | Type of time-dependent intensity (TDI) correction applied to compensate for beam-induced volatilization or migration of sensitive elements (e.g., Na, K, F in glasses, feldspars, carbonates). | Advanced | Editable | Controlled list | None ¦ Linear ¦ Exponential ¦ Double exponential ¦ Unknown ¦ N/A |  | 2026-09-01 | target species | Most commonly applied in point analysis; uncommon for X-ray mapping. | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Target Species Estimation Method | Whether elemental concentrations were calculated directly from measured X-ray intensities, or estimated by cation stoichiometry (e.g., oxygen calculated from cation proportions in silicates; carbon from stoichiometry in carbonates). | Advanced | Read-Only | Controlled list / Text | Direct (measured) ¦ Stoichiometry from cations ¦ By difference ¦ Unknown ¦ N/A ¦ None |  | 2026-09-01 | target species |  | Y | Y | Y | Y |  | N | N | N | N | N | N | O, CO2: Stoichiometry from cations; H2O: estimated by inspection of the analytical total; F: Direct (measured) — p.3 | OH: By difference; other: Direct (measured) — 'Hydroxyl was not measured directly in this study and therefore was calculated by difference based on 1–F–Cl=OH'; formulae on 13 anions (Ketcham 2015) | N | N | N | N | N | N | N
Halogen Correction on Oxygen | Whether oxygen content was adjusted to account for halogen substitution (F and/or Cl replacing OH) in halogen-bearing phases such as apatite, amphibole, and mica, where oxygen is calculated by stoichiometry. | Advanced | Editable | Controlled list / Text | e.g., 'Yes — F and Cl substitution in apatite; O calculated as 1-F-Cl=OH' ¦ 'No' ¦ N/A ¦ None |  | 2026-08-30 | (none) |  | Y | Y | Y | Y |  | N/A | N/A | N/A | N/A | N/A | N/A | N | N — OH is 'calculated by difference based on 1–F–Cl=OH' (under Target Species Estimation Method); no oxygen-equivalent correction for F and Cl is stated | N/A | N/A | N | N | N/A | N/A | N/A
WDS Dead Time Correction | Method used to correct for WDS proportional counter dead time at high count rates. Unlike EDS dead time — which is hardware-managed and reported as a session QC percentage (see EDS Dead Time) — WDS dead time correction is a user-selectable algorithm in the data reduction software. No separate measured WDS dead time value is reported; the correction is applied transparently during intensity-to-concentration conversion. Record the algorithm here and any instrument-specific constant alongside it — 'Default constant (manufacturer) — 3 µs, Cameca'. The instrument vendor itself is recorded by Instrument Manufacturer, not by this field's allowed values. | Basic | Read-Only | Controlled list / Text | Default constant (manufacturer) ¦ Adjusted constant ¦ Logarithmic ¦ High-precision (Probe for EPMA) ¦ Super-precision (Probe for EPMA) ¦ Unknown ¦ N/A ¦ None |  | 2026-08-31 | (none) | Dead time errors are most significant for major elements with high count rates (e.g., Si, Fe, Ca). | N | N | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
EDS Spectral Processing Type | Method used to process EDS spectra and extract net peak intensities from raw spectral data. Applied before quantification (see Matrix Correction Method). Common approaches include background fitting and subtraction followed by peak integration, and filter fit or Gaussian deconvolution for overlapping peaks. | Advanced | Read-Only | Controlled list | Background fitting ¦ Peak deconvolution ¦ Top-hat filter ¦ Gaussian fit ¦ Filter fit ¦ Unknown ¦ N/A ¦ None |  | 2026-08-30 | (none) |  | Y | Y | N | N |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Blank Correction | Method and reference material(s) used to determine and subtract blank signal contributions (e.g., carbon coat contribution to C signal, or background contamination for trace elements). | Advanced | Editable | Text (free) | e.g., 'Blank C signal subtracted using carbon-coated polished glass reference' ¦ 'None' |  | 2026-09-01 | target species |  | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Normalization / Standards-Based Correction | Post-acquisition normalization applied to the reported data beyond the primary calibration — for example correction to a reference value derived from secondary reference materials, or correction for a systematic bias those materials reveal. Record 'None' if no additional normalization is applied. | Advanced | Editable | Text (free) | e.g., 'Normalization to accepted values of Kakanui kaersutite run as secondary standard' ¦ 'None' | Source: Composition QC module | 2026-08-27 | reported property |  | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N — the k-ratio is the calibration, and is under Calibration Factor and Determination Method
Calibration Factor and Determination Method | An externally-calibrated factor that converts the measured quantity into the reported quantity, how it was determined, and its uncertainty. Applies where the conversion depends on a factor calibrated against a reference of independently known value, rather than on the instrument response alone. Distinct from the fields that name the calibration material and that state which approach applies to which target species, where the technique has them: this field records the resulting factor itself. | Basic | Editable | Text (free) |  | Source: Calibration Factor module | 2026-09-01 | reported property |  | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | all: k-ratio against the calibration standard — 'the k-ratio (k = (P-B)smp/(P-B)std) is the background corrected relative peak X-ray intensity (P-B) at each pixel in the map compared to the calibration standard', with 'C = k × ZAF'
Procedural Blank Level | The measured level of the analytical blank in the session, and — where the reported quantity is a ratio — its composition, since a blank subtracted from a ratio biases the result unless its own composition is known. Companion to the blank correction method. | N/A | Basic | Text (free) |  | Source: Blank module | 2026-09-01 | target species | Follows the criterion-versus-measurement split the library applies wherever a procedure sets a threshold and an analysis reports a value against it. | Y | Y | Y | Y |  | N/A — no chemical separation, so there is no procedural blank | N/A — no chemical separation, so there is no procedural blank | N/A — no chemical separation, so there is no procedural blank | N/A — no chemical separation, so there is no procedural blank | N/A — no chemical separation, so there is no procedural blank | N/A — no chemical separation, so there is no procedural blank | N/A — no chemical separation, so there is no procedural blank | N/A — no chemical separation, so there is no procedural blank | N/A — no chemical separation, so there is no procedural blank | N/A — no chemical separation, so there is no procedural blank | N/A — no chemical separation, so there is no procedural blank | N/A — no chemical separation, so there is no procedural blank | N/A — no chemical separation, so there is no procedural blank | N/A — no chemical separation, so there is no procedural blank | N/A
Analysis Inclusion and Rejection Criteria | The rules determining which individual results contribute to a combined result: the limits a result must meet (for example on analytical total or stoichiometry) and the grounds on which one is rejected (for example overlap onto a neighbouring phase, or an outlier test with its threshold), including any acceptance threshold on the dispersion statistic. An individual result is the value of the reported quantity obtained from one acquisition: a replicate measurement of the same solution or location, a spot or grain within a sample, or an independently prepared aliquot or digestion, whichever the procedure combines. Distinct from filtering the acquired signal during data reduction (removing spikes, cycles or scans, or discarding an acquisition whose signal is compromised): this field records the rules by which finished results enter a combined result. In the dataset: flag each individual result excluded from a combined value, with the rule that excluded it. | Basic | Editable | Text (free) | e.g., 'Totals 98.5–101.5 wt%; points overlapping a neighbouring phase rejected' ¦ 'All results included; no filter applied' | Source: Aggregation module | 2026-10-26 | (none) | The same measurements give different combined values under different acceptance rules, so the rules are needed to reproduce a combined value from its individual results. | Y | Y | Y | Y |  | N — no acceptance or rejection rule is stated. Re-read 2026-10-26 for the rules-only definition (Module_Aggregation v6); the earlier reading, kept for provenance: the contributing counts are stated per aggregate, each Table 1 column being the mean of n point analyses of one phase and textural setting (n = 6, 6, 6, 9, 17, 7 and 5; p.5), with one standard deviation of the mean. No acceptance or rejection rule, and no acquired-versus-included count, is stated | N — analyses are reported by phase with no contributing count and no acceptance or rejection rule stated | N — the modal fractions use every pixel of the mapped section ("the number of pixels attributed to each mineral was divided by the total number of pixels in the whole section", p.3); nothing is admitted or excluded | N — no acceptance or rejection rule is stated. Re-read 2026-10-26 for the rules-only definition (Module_Aggregation v6); the earlier reading, kept for provenance: contributing counts are stated for the glass averages ('EMP avg (n = 73)' and 'avg (n = 14)', Table 3, p.9). No acceptance or rejection rule is stated. The table's n = 7 and n = 13 are LA-ICP-MS averages, not microprobe ones | N — no acceptance or rejection rule is stated. Re-read 2026-10-26 for the rules-only definition (Module_Aggregation v6); the earlier reading, kept for provenance: each Table 1 column is the mean of n point analyses of one occurrence (n = 6, 2, 3, 3, 5 and 4; p.3), with one standard deviation of the mean. No acceptance or rejection rule is stated | N — the microprobe analyses are reported as "Representative electron-microprobe measurements" (p.5) with no contributing count and no selection rule. The counts on p.8 (n = 9, n = 7) are SIMS standard populations, not microprobe aggregates | N — no acceptance or rejection rule is stated. Re-read 2026-10-26 for the rules-only definition (Module_Aggregation v6); the earlier reading, kept for provenance: the carbonate compositions are means of stated counts ("Dolomite contains 2.0 ± 0.4 wt% Fe and 3.0 ± 0.9 wt% Mn (n = 37)"; "Magnesite contains 14.5 ± 2.7 wt% Fe and 4.6 ± 2.5 wt % Mn (n = 23)", p.5). No acceptance or rejection rule is stated | N — Table 1 reports one column per named apatite grain rather than an aggregate over results, and no acceptance or rejection rule is stated | N — no acceptance or rejection rule is stated. Re-read 2026-10-26 for the rules-only definition (Module_Aggregation v6); the earlier reading, kept for provenance: the reported averages state their contributing counts ("based on 12 analyses" for orthopyroxene, and "14 analyses" for augite, p.2; "based on 13 analyses" and "34 ± 7 mol% on average; 19 analyses" for the Ca-Eskola component, p.4). No acceptance or rejection rule is stated | N — compositions are reported by phase with no contributing count and no selection rule stated | N — compositions are reported by phase with no contributing count and no selection rule stated | N — no contributing count and no acceptance or rejection rule is stated for the microprobe analyses; the modal abundances come from classified phase-map pixels rather than from admitting or excluding results (p.9) | N — the analyses are "compiled in Supplementary Table 14" (p.11), not in the archived PDF, and no count or selection rule is stated in the text. The "Bennu (n = 58)" population (Fig. 5, p.6) is the SIMS oxygen-isotope dataset, not this procedure's | N — no contributing count and no acceptance or rejection rule is stated for the olivine and pyroxene analyses | N
Combination Method | Whether the reported variable is an individual result or combines several, and if combined, how: the statistical model used, including any criteria governing which model is applied, and what the individual results are combined over (for example per grain, per phase within a sample, or per sample). Record 'Not combined' where each individual result is reported on its own. Record the model only; the software implementing it belongs in Data Processing Software. In the dataset: report each combined value with its uncertainty and the number of individual results it combines, and identify, for each individual result, the combined value it contributes to. | Basic | Editable | Text (free) | e.g., Arithmetic mean per phase within each sample ¦ Weighted mean with MSWD-based uncertainty expansion ¦ York (1969) isochron regression ¦ Not combined | Source: Aggregation module | 2026-10-26 | reported property | A methodological choice that changes the result: an arithmetic and a weighted mean of the same data, or a Model-1 and a Model-3 regression, give different values and different uncertainties. | Y | Y | Y | Y |  | SiO2, TiO2, Al2O3, FeO, MgO, CaO, Na2O, K2O, Cr2O3, MnO: mean of n point analyses per phase and textural setting; other: N — Table 1 notes a and b | N | SiO2, TiO2, Al2O3, Cr2O3, MgO, CaO, MnO, FeO, NiO, Na2O, K2O, P2O5: average of n microprobe analyses per glass type; other: N — Table 3: 'Glass inclusion EMP avg (n = 73) 1σ' and the impact-melt 'avg (n = 14) 1σ', whose average 'only included the glassy regions of the impact melt pockets' | SiO2, TiO2, Al2O3, Cr2O3, MgO, CaO, MnO, FeO, NiO, Na2O, K2O, P2O5: average of n microprobe analyses per glass type; other: N — Table 3: 'Glass inclusion EMP avg (n = 73) 1σ' and the impact-melt 'avg (n = 14) 1σ', whose average 'only included the glassy regions of the impact melt pockets' | SiO2, TiO2, Al2O3, FeO, CaO, Na2O, K2O: mean of n point analyses per occurrence; other: N — Table 1 notes b and c | N | Element concentrations in the carbonates: mean per carbonate phase; other: N — dolomite n = 37, magnesite n = 23 (p.5) | N | En, Fs, Wo: average per pyroxene phase; Ca-Eskola component: average per clinopyroxene setting; other: N — 'The average compositions ... of orthopyroxene ... based on 12 analyses'; 'Ca-Esk components (41 ± 8 mol% on average; based on 13 analyses)' | N | N | N | N | N | N
Constants and Reference Values Used | Physical constants and reference values used in data reduction to calculate the final reported quantity (e.g., decay constants for age calculation, standard isotope ratios, or other citable reference values used in a correction or calculation), together with their source. Distinct from the Group 6 reference-material fields, which document accepted values for specific calibration/validation materials rather than universal physical constants. Record "None" if no citable, revisable physical constants feed into this procedure's data reduction. | Basic | Editable | Text (free) | e.g., "λ238U = 1.55125×10⁻¹⁰ yr⁻¹, λ235U = 9.8485×10⁻¹⁰ yr⁻¹ (Jaffey et al. 1971); 238U/235U = 137.818 (Hiess et al. 2012)" ¦ "None" | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
6. Quality Control & Uncertainty | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Primary Calibration Standard Name | Name and reference material identifier of the primary reference material(s) against which the instrument is calibrated — converting raw signal intensities to concentrations, or anchoring an isotope ratio as the bracketing standard or zero-delta reference. Give the material name, its source or supplier, and a citation for the accepted values used. Where calibration instead uses the vendor's stored library or theoretical response factors rather than measured reference materials — 'standardless' or 'semi-quantitative' quantification — record that here, naming the library or model used. 'None' means no calibration was performed at all, which is a different answer. | Basic | Editable | Text (free) | e.g., 'Smithsonian anorthite' ¦ 'Amelia albite' ¦ 'Synthetic fayalite' ¦ 'SrF2' | Source: Composition QC module | 2026-09-29 | target material x target species | Results calibrated against different published values for the same material are not directly comparable. | Y | Y | Y | Y |  | all [Si, Al, Ca: anorthite; Na: albite; Fe: fayalite; Mg: forsterite; Mn: Mn2SiO4; Ti: TiO2; Cr: Cr2O3; K: microcline] — p.3 | all [Si, Mg, Fe: natural kaersutite; Na, Al: jadeite; Ca, Mn: bustamite; K: K-feldspar; Ti: synthetic rutile; Cr: Cr2O3] — p.2 | N | N | all [Si, Al, K: Asbestos microcline; Ca: synthetic anorthite; Na: Amelia albite; Fe: synthetic fayalite; Mg: synthetic forsterite; Ti: synthetic TiO2; Cr: synthetic Cr2O3; Mn: synthetic Mn-olivine] — p.2 | all [Si, Al, Ti, K, Na, Fe, Mg, Ca: Kakanui kaersutite; S: Canyon Diablo troilite; Mn: rhodonite; Cr: chromium metal; Ni: nickel metal; P: apatite; V: vanadium metal] — pp.3–4 | all [F: synthetic F-phlogopite; other: N] — 'natural and synthetic minerals routinely used in the analytical facility' | all [F: SrF2; Na: albite; Mg: SW olivine; Si: quartz; P, Ca: Wilburforce apatite; S: barite; Cl: tugtupite; Mn: rhodonite; Fe: ilmenite] — p.3 | N — 'Natural and synthetic standards were used' | N — standards are named without their elements: magnetite (USNM 114887), calcite (USNM 13621), dolomite (USNM 10057), siderite (R-2460) and rhodonite (carbonates); chromite (USNM 117075), ilmenite (USNM 96189), magnetite (USNM 114887), manganite (USNM 157872), bytownite (R-2912), forsterite (P140), San Carlos olivine (USNM 111312/444) and Springwater olivine (USNM 2566) (magnetite and olivine) | "Mg,Na phosphate" [F, P, Ca: fluorapatite; Si: Fo92 olivine; Mg: Fo92 olivine, rhodonite; Fe: fayalite; Al: anorthite; S: baryte; K: potassium feldspar; Cl: scapolite]; carbonate [Na: albite; Si: Fo92 olivine; Mg: dolomite; Ca: calcite; Mn: Mn carbonate; P: apatite; S: baryte; Fe: fayalite] — 'The standards used for Mg,Na phosphate were fluorapatite (F, P, Ca), Fo92 olivine (Si, Mg), rhodonite (Mg), fayalite (Fe), anorthite (Al), baryte (S), potassium feldspar (K) and scapolite (Cl). For carbonates, the standards used were albite (Na), Fo92 olivine (Si), dolomite (Mg), calcite (Ca), Mn carbonate (Mn), apatite (P), baryte (S) and fayalite (Fe)' (p.7); rhodonite is printed against Mg | N — 'Well-characterized natural and synthetic materials were used as standards' | all [Mg, Si: Springwater olivine; Fe: fayalite; Ca: wollastonite; Na, Al: albite; K: orthoclase; Ti: rutile; Ni: Ni metal; Cr: chromite; Mn: rhodochrosite] — p.11 | N | N — 'EPMA standards having a range of average atomic number Z' are used for the MAN background calibration; individual standards not named
Secondary Reference Materials | Quality-control reference material(s) measured as unknowns alongside samples to assess accuracy independently and to monitor drift. Give the material name, its source, and a citation for the accepted or reference values used for comparison. | Basic | Editable | Text (free) | e.g., 'Kakanui kaersutite (Si, Al, Ti, Fe, Mg, Ca, Na, K; n=5; within 1% of GeoReM values)' ¦ 'Durango apatite (F, Cl, P, Ca; Wudarska et al. 2021)' ¦ 'USNM San Carlos olivine 111312 (Mg, Fe, Si)' | Source: Composition QC module | 2026-08-27 | defines: standard |  | Y | Y | Y | Y |  | N | N | N | N | feldspar standards — 'based on analysis of feldspar standards as unknowns'; not named individually | N — no EPMA secondary standard is named; San Carlos olivine standardised the SIMS oxygen-isotope measurements, and Kakanui kaersutite is a primary standard | Smithsonian Microbeam standards (specific materials and values N) | N | N | calcite, dolomite, rhodochrosite, magnetite, San Carlos olivine, Springwater olivine — carbonates: calcite, dolomite and rhodochrosite; magnetite and olivine: magnetite, San Carlos olivine and Springwater olivine (p.7) | N | N | N | N | N
X-ray Line Overlap Corrections Applied | Whether a spectral interference correction was applied. Common interferences include Ti Kb on V Ka, Cr Kb on Mn Ka, and Ba La on Ti Ka. | Basic | Read-Only | Controlled list / Text | Yes — specify the overlapping lines and the correction method ¦ No ¦ N/A ¦ None |  | 2026-09-01 | target species |  | Y | Y | Y | Y |  | N | Mn: Yes; other: N — 'X-ray interference of the Ka line of Mn by the Kb line of Cr was corrected' | N | N | N | N | N | N | N | N | N | N | N | N | N
Interfering Elements | Element(s) whose X-ray lines overlap with the measured peak, requiring a correction. | Advanced | Read-Only | Text (free) | e.g., 'Ti Kb' ¦ 'Cr Kb' ¦ 'Ba La' |  | 2026-09-29 | target species |  | Y | Y | Y | Y |  | N | Mn: Cr Kb; other: N — as above | N | N | N | N | N | N | N | N | N | N | N | N | N
Interference Correction Standard | Reference material used to quantify and calibrate the interference correction. | Advanced | Read-Only | Text (free) | e.g., 'Synthetic TiO2 (Ti Kb overlap)' ¦ 'Chromite USNM 117075 (Cr Kb overlap)' |  | 2026-09-29 | target species |  | Y | Y | Y | Y |  | N | N — the Cr Kb correction on Mn Ka is stated, but no correction standard is named | N | N | N | N | N | N | N | N | N | N | N | N | N
Detection Limit | Detection limit, one per reported concentration variable (one per target species, these being the same set). State the units and whether the values are procedure-typical estimates or session-specific measured values. The calculation method is recorded separately in Detection Limit Method. Record 'N/A' where the procedure reports no concentrations. In the dataset: report the detection limit of each analysis, where the procedure computes one. | Advanced | Basic | Numeric + unit / Text | e.g., '0.02 wt%' ¦ '0.03 wt%' ¦ '0.04 wt%' | Source: Composition QC module | 2026-10-27 | reported property |  | Y | Y | Y | Y |  | K2O: 0.02 wt% K; Cr2O3: 0.05 wt% Cr; MnO: 0.06 wt% Mn; other: N — Table 1 footnote c, attached to the b.d. entries of the K2O, Cr2O3 and MnO rows: 'b.d. = below detection limit: 0.02 wt% K, 0.05 wt% Cr, 0.06 wt% Mn' | K2O: 0.01 wt%; SiO2, Al2O3, MgO, CaO, Na2O: 0.02 wt%; TiO2, Cr2O3: 0.03 wt%; FeO: 0.05 wt%; MnO: 0.06 wt% — p.2 | SiO2, TiO2, Al2O3, MgO, CaO: <0.03 wt%; FeO, MnO, Cr2O3, NiO, Na2O, K2O, P2O5: <0.05–0.1 wt%; other: N — 'Detection limits are typically <0.03 wt% for SiO2, TiO2, Al2O3, MgO, and CaO; <0.05–0.1 wt% for FeO, MnO, Cr2O3, NiO, Na2O, K2O, and P2O5' (p.3) | SiO2, TiO2, Al2O3, MgO, CaO: <0.03 wt%; FeO, MnO, Cr2O3, NiO, Na2O, K2O, P2O5: <0.05–0.1 wt%; other: N — 'Detection limits are typically <0.03 wt% for SiO2, TiO2, Al2O3, MgO, and CaO; <0.05–0.1 wt% for FeO, MnO, Cr2O3, NiO, Na2O, K2O, and P2O5' (p.3) | Si: 0.05 wt%; Ti: 0.04 wt%; Al: 0.06 wt%; Fe: 0.06 wt%; Mg: 0.02 wt%; Ca: 0.02 wt%; Na: 0.03 wt%; K: 0.02 wt%; Cr: 0.05 wt%; Mn: 0.06 wt% — 'The detection limits (wt%) are 0.05 Si, 0.04 Ti, ...' (p.2) — stated per element, while Table 1 reports oxides | Al2O3, K2O, CaO: 0.03–0.04 wt%; Na2O, MgO, SiO2, FeO, MnO: 0.05 wt%; P2O5, SO2, TiO2, V2O3, Cr2O3, NiO: 0.06–0.09 wt% — 'typically' (p.4) | N | N | all: better than 0.02 wt% (typical) — 'Typical detection limits for oxides of most elements were better than 0.02 wt%' | N | N | N | Mg: 0.025 wt%; Fe: 0.025 wt%; Si, K, Na: 0.05 wt%; Ca: 0.005 wt%; Al: 0.02 wt%; Ti: 0.005 wt%; Cr: 0.015 wt%; Mn: 0.008 wt%; other: N — p.11 | N — 'Typical detection limits for transition metals were around 250 ppm'; the elements are not named | all: 0.1–0.2 element wt% — 'for all elements in this map set'
Detection Limit Method | Formula or approach used to calculate the detection limits, with a citation for the method where one exists. Must be consistent with the values reported in Detection Limit. | Basic | Read-Only | Controlled list / Text | 3σ blank ¦ 3σ background ¦ 3σ counting statistics ¦ 3× blank mean ¦ Poisson statistics ¦ N/A ¦ None | Source: Composition QC module | 2026-08-27 | reported property |  | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N — attributed to the MAN background correction dedicating all map collection time to on-peak measurement, 'which improves precision and detection limits'
Analytical Precision | Precision of repeated measurements on the same or equivalent reference material, expressed as 1-sigma relative standard deviation (%). State whether the repeats were made within one session (repeatability) or across sessions (intermediate precision, in ISO 5725 terms). Include reference material name, number of analyses (n), and the measured value. | Advanced | Basic | Text (free) | e.g., '0.5% 1-sigma RSD (n = 20)' ¦ '1.2% 1-sigma RSD (n = 15)' |  | 2026-10-29 | standard x reported property |  | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Analytical Accuracy | Offset between measured and accepted reference values for secondary standards, expressed as percent relative bias. Include reference material, reference value source, and the measured value. | Advanced | Basic | Text (free) | e.g., '+0.5% relative to GeoReM preferred value' ¦ '-1.2% relative to accepted value' |  | 2026-09-29 | standard x reported property |  | Y | Y | Y | Y |  | N | N | N | N | feldspar standards [Si, Al, Ca, Na, K: 1–2%; other: N] — 'The accuracy is 1–2% for Si, Al, Ca, Na, and K, based on analysis of feldspar standards as unknowns' — per element, while Table 1 reports oxides | N | N | N | N | N | N | N | N | N | N
Counting Statistics Error | How the uncertainty predicted from counting statistics — the theoretical limit set by the Poisson distribution of the counts accumulated — is computed for each reported quantity, and the sigma level it is quoted at. It is derived from the counts on the target species together with those on any background or blank subtracted from it. Distinct from the scatter actually observed within a measurement or between repeated measurements, which is recorded separately. In the dataset: report, for each analysis and reported quantity, the counting-statistics uncertainty at the stated sigma level. | Advanced | Editable | Text (free) | e.g., '1-sigma, from the counts on the peak and on the subtracted background' | Source: Composition QC module | 2026-10-27 | reported property | Where a procedure reports both, agreement indicates the measurement is shot-noise limited, and a larger observed scatter indicates a further source of variance. | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
EDS Dead Time | Percent dead time reported by the EDS detector during the session — the fraction of total acquisition time the detector spent processing rather than counting. This field documents the resulting percentage as a session QC metric. Unlike WDS dead time (see WDS Dead Time Correction), no user-selectable correction algorithm is required. | N/A | Basic | Numeric (%) | e.g., 10% ¦ 25% ¦ <40% |  | 2026-08-27 | (none) | Values above ~40% indicate excessive count rate and may degrade spectral quality and quantitative accuracy. EDS dead time correction is managed automatically by the detector electronics. | Y | Y | N | N |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Goodness-of-Fit or Dispersion Statistic | The statistic used to describe the scatter among the individual results contributing to a combined value, or to test whether that scatter exceeds what analytical uncertainty alone predicts: for example one standard deviation of the mean, or the MSWD with its probability of fit. A threshold for accepting it belongs with the inclusion criteria. In the dataset: report, for each combined value, the value of the statistic, and any further statistic (for example the standard error of the mean, or a 95% confidence interval) beside it, named. | Basic | Editable | Text (free) | e.g., One standard deviation of the mean ¦ MSWD with probability of fit | Source: Aggregation module | 2026-10-26 | reported property | Without the statistic named, the uncertainty given with a combined value cannot be interpreted, nor can whether the combined results form a single population be judged. | Y | Y | Y | Y |  | SiO2, TiO2, Al2O3, FeO, MgO, CaO, Na2O, K2O, Cr2O3, MnO: one standard deviation of the mean; other: N — Table 1 note b: 'Errors given inside parentheses are one standard deviation of the mean based on all of the analyses' | N | SiO2, TiO2, Al2O3, Cr2O3, MgO, CaO, MnO, FeO, NiO, Na2O, K2O, P2O5: 1σ, one standard deviation of the average; other: N — Table 3 note b: '1σ is 1 standard deviation of the average' | SiO2, TiO2, Al2O3, Cr2O3, MgO, CaO, MnO, FeO, NiO, Na2O, K2O, P2O5: 1σ, one standard deviation of the average; other: N — Table 3 note b | SiO2, TiO2, Al2O3, FeO, CaO, Na2O, K2O: one standard deviation of the mean; other: N — Table 1 note c, same wording | N | N — the carbonate means are given as '±' values ('2.0 ± 0.4 wt% Fe') without naming the statistic | N | N — the averages are given as '±' values ('41 ± 8 mol%') without naming the statistic | N | N | N | N | N | N
Additional Notes | Any procedure- or analysis-specific information not captured by a structured field anywhere in this TAPP — including anomalies, deviations from the registered procedure, instrument modifications, and supplementary context. Scope is the whole document, not Group 6: this is the last field of the TAPP and covers all six groups. Use sparingly; a structured field is preferred for anything that can be formally categorised. | Advanced | Advanced | Text (free) | Free text | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y |  | Ma et al. 2015, Earth Planet. Sci. Lett. — tissintite discovery paper (Tissint Mars meteorite). Instrument stated as "JEOL 8200 electron microprobe" (no JXA prefix). WDS explicitly stated ("WDS: 15 kV; 5 nA; beam in focused mode"). Point analysis only; no X-ray mapping reported. Probe for EPMA stated; CITZAF correction procedure (Armstrong 1995). Full standard suite with X-ray lines given. Detection limits: K=0.02, Cr=0.05, Mn=0.06 wt% from Table 1 footnote. Caltech GPS Division Analytical Facility. | Hu et al. 2020, Geochim. Cosmochim. Acta — coesite in NWA 8657 shergottite. JEOL JXA-8100 at IGGCAS; 15 kV, 10 nA; point analysis only; WDS or EDS not stated. Matrix correction: Bence-Albee (not PAP). Full primary standard suite stated (kaersutite, jadeite, bustamite, K-feldspar, rutile, Cr2O3). Mn Kα / Cr Kβ interference correction applied. Detection limits 0.01-0.06 wt% stated per oxide. Analytical software not stated. | Liu et al. 2016, Meteorit. Planet. Sci. — Tissint mineral chemistry. Protocol 1 of 2: University of Tennessee Cameca SX100. Same paper also uses Caltech GPS JXA-8200 (see Liu+2016_Cal column). Point analysis AND X-ray mapping performed at UT. Specific mapping: BSE + Ca/Al/Fe/Mg Ka maps (15 kV, 20 nA, step 8-12 µm). Olivine megacryst mapping (15 kV, 200 nA, step 2 µm, dwell ~0.5 s) described as "using the EMP" — instrument ambiguous (may be UT or Caltech instrument). Standards, matrix correction, and software not stated. | Liu et al. 2016, Meteorit. Planet. Sci. — Tissint mineral chemistry. Protocol 2 of 2: Caltech GPS Division JEOL JXA-8200. Point analysis only (no mapping attributed to Caltech instrument). Conditions stated jointly for UT and Caltech instruments. Standards, matrix correction, and software not stated for EPMA. | Ma et al. 2018, Meteorit. Planet. Sci. 53:50-61 (file dated 2017) — liebermannite (KAlSi3O8) discovery from Zagami. Instrument stated as "JEOL 8200 electron microprobe" (no JXA prefix in text). WDS explicitly stated ("WDS: 15 kV, 5 nA"). Probe for EPMA; CITZAF correction (Armstrong 1995) — NOT PAP. Full standard suite and X-ray lines stated. K-mapping by EPMA also performed (used for mineral identification) but mapping conditions (step size, dwell time, current) not stated. Na diffusion observed during analysis despite low 5 nA beam current. Detection limits stated (per-element wt% values). Analytical accuracy: 1-2% for Si, Al, Ca, Na, K (feldspar standards as unknowns). Caltech GPS Division Analytical Facility. | Frank et al. 2023, Meteorit. Planet. Sci. 58:1495-1511 — CAI in Ivuna CI chondrite. ARES NASA JSC. Instrument stated as "Cameca SX100 electron microprobe at ARES, Johnson Space Center" — NOT JEOL JXA-8530F as in v2 header. Accelerating voltage 20 kV (not 15 kV). Both point analysis (20 kV, 20 nA, 1 µm focused) and X-ray mapping performed. X-ray mapping described but conditions (step size, dwell time, mapping beam mode) N. WDS not explicitly stated. Matrix correction and background correction method N. Peak counting time 10-50 s. Primary standard suite fully documented. No EPMA secondary standard is named (San Carlos olivine standardised the SIMS work). Detection limits stated per element group. | Broussard et al. 2026, Meteorit. Planet. Sci. — OC002 CI chondrite links Bennu and Ryugu. Washington University in St. Louis. Instrument stated as "JEOL JXA-8200 electron microprobe" — NOT JXA-8230 as in v2 header. WDS explicitly stated ("wavelength-dispersive quantitative compositional mapping and analysis"). CITZAF matrix correction (Armstrong 1995) — NOT PAP or XPP. MAN background for most analytes; polynomial fit for F via LDE1 crystal. Both point analysis (15 kV, 25 nA) and quantitative stage mapping performed. O by stoichiometry from cations. F is the only explicitly named analyte in methods; full list N. An EDS spectrometer is on the instrument; its use is not stated. Smithsonian Microbeam standards as secondary QC. No peak counting time, beam diameter, detection limits, or interference corrections stated. | Seifert et al. 2026, Meteorit. Planet. Sci. — apatite in Bennu OSIRIS-REx samples. ARES NASA JSC. Instrument stated as "JEOL 8530 EMPA at NASA JSC" (no "JXA", no "F", no "+" suffix stated in paper). Analytical conditions: 15 kV, 20 nA, 2 µm probe size. Previous v2 values of 10/40-100 nA and 10 µm beam were WRONG — those were Durango apatite test conditions used to assess beam damage, not the actual protocol. Analytes: P, F, Cl, Ca, Mn, Fe, Na, Mg, Si, S. Apatite stoichiometry by Ketcham (2015) method (13-anion basis; OH by difference). Primary standards: SrF2, albite, olivine, quartz, apatite, barite, tugtupite, rhodonite, ilmenite. Sample preparation: fragments embedded in epoxy, dry-polished, ion-polished (one mount), carbon coated. 14 total analyses performed. | Pang et al. 2016, Sci. Rep. 6:26063 — NWA 8003 eucrite, Nanjing University. JEOL JXA-8100 (stated as "JEOL 8100"). WDS explicitly stated ("JEOL 8100 WDS"). ZAF matrix correction (NOT "ZAF or PAP" as in v2; paper states ZAF). Focused beam (20 nA) for most phases; defocused 2-5 µm for plagioclase and polymorphs. Natural and synthetic mineral standards (specific names N). Detection limit better than 0.02 wt% (as stated). Analytical software not stated. | McCoy et al. 2025, Nature 637:320-325 — Bennu evaporites. Protocol 1 of 2: Smithsonian Institution JEOL 8530 F+ Hyperprobe (Field Emission). Ir-coated specimens mounted on Ir-coated Parafilm. Carbonate analyses: 15 kV, 10 nA, 5 µm spot; LIFL (Fe,Mn), TAPL (Mg), PETL (Ca). Magnetite and olivine analyses: 15 kV, 10 nA, 1 µm spot; their own standard suite. Both primary and secondary standard suites fully documented with USNM catalog numbers. Acquisition software and matrix correction method N. WDS not explicitly stated in text (crystal designations LIFL/TAPL/PETL confirm WDS use). | McCoy et al. 2025, Nature 637:320-325 — Bennu evaporites. Protocol 2 of 2: U. Arizona K-ALFAA Cameca SX-100. 20 nm carbon coat. WDS explicitly stated for phosphate analyses. Mg,Na phosphate analyses: 15 kV, 8 nA, 1 µm. Carbonate analyses at K-ALFAA also mentioned; conditions N. Full primary standard suite documented for phosphates and carbonates. Acquisition software and matrix correction method N. | Zega et al. 2025, Nat. Geosci. — mineralogical evidence for hydrothermal alteration of Bennu. K-ALFAA, University of Arizona. Instrument stated as "SX-100 Ultra electron microprobe in the K-ALFAA". IMPORTANT: v2 had "no protocol details reported" — this was WRONG. The paper provides detailed EPMA conditions: X-ray maps and BSE images: 15 kV, 20 nA. Silicates/sulfides/oxides: 15 kV, 20 nA, focused, 20 s peak, 10 s on each background. Phosphates: 15 kV, 8 nA, 2 µm defocused, 20 s peak and 10 s background. Carbonates: 15 kV, 4 nA, 2 µm defocused, 10 s peak and 5 s background. Standards: "well-characterized natural and synthetic materials" (specific names N). Phase maps generated using XMapTools. WDS and matrix correction NOT explicitly stated in paper. | Barnes et al. 2025, Nat. Astron. — variety and origin of accreted materials in Bennu. Protocol 1 of 2: CRPG Nancy, JEOL JXA-8230. Instrument has 5 WDS spectrometers + 1 SDD EDS; per-analyte technique (WDS vs. EDS) not stated. Two analytical sessions: session 1 (no Na, K); session 2 (with Na, K). Counting times are stated as total peak + background combined: 200 ms for minor elements (Al, Ti, Ca, Mn, Cr) and 20 ms for major elements (Mg, Fe, Si). Full primary standard suite stated with element assignments. Full per-element detection limits stated. Matrix correction method not stated. Sample preparation done at Université Côte d'Azur (not at CRPG). Beam current not stated for NHM protocol; 3 nA mentioned in text is for SEM-EDS (different instrument). | Barnes et al. 2025, Nat. Astron. — variety and origin of accreted materials in Bennu. Protocol 2 of 2: NHM London, CAMECA SX100. Stated instrument: "CAMECA SX100 electron microprobe". Target minerals: olivine and pyroxene (anhydrous silicates). 20 kV, 1 µm focused beam. Beam current not stated for EPMA (3 nA in text refers to SEM-EDS on separate Zeiss EVO instrument). Detection limits ~250 ppm for transition metals. Standards, matrix correction, WDS spectrometer details not stated. Analyte list not given. SEM-EDS at NHM is a separate instrument (Zeiss EVO 15LS + Oxford X-Max80) calibrated at 20 kV, 3 nA. Carbon coat: initial coat for SEM/EPMA (thickness N); additional coat to ~30 nm total was for subsequent SIMS, not EPMA. | Multi-pass WDS mapping: two passes per stage map, five elements each; 18 hr per map; 20 x 10^6 fully quantitative analyses across all slides. Recorded in the acquisition-pass proposal (2026-09-08) as the evidence that EPMA partitions the target-species domain across passes.

## Legends
Procedure-Level Tier | Description | Analysis-Level Tier | Description
Basic | Mandatory for procedure registration. Must be provided to register a valid procedure. | Read-Only | Directly imported from the registered procedure; cannot be changed by the analyst. Changing this value means running a different procedure. Also used for fields relevant only at procedure level — the value is inherited from the procedure record and shown read-only in the analysis form.
Advanced | Optional for procedure registration. Strongly recommended but not required. | Editable | Imported from the registered procedure but may be adjusted within procedure-defined bounds (e.g., daily tuning, minor software updates). The procedure registers the target or typical value; the analyst confirms or adjusts it. Cannot be left void if the procedure-level tier is Basic.
N/A | Not applicable at procedure level. This field captures analysis-level information only. | Basic | Mandatory user input at analysis time. Value comes from the analysis itself and cannot be pre-specified in the procedure.
|  | Advanced | Optional user input at analysis time. Recommended for complete and reproducible documentation.
Mode Column | Definition
Purpose
EDS Point Analysis
EDS Mapping
WDS Point Analysis
WDS Mapping
Y | This field applies to this analytical mode and should be reported.
N | This field does not apply to this analytical mode.
Data Type | What to enter
Controlled list | Choose one of the values listed in "Example / Allowed Content". The list is CLOSED — it is meant to cover every case. If your value genuinely is not there, that is a gap in the list worth reporting, not a reason to write your own.
Controlled list / Text | Choose a listed value AND qualify it in the same cell — the detail is expected, not optional. Typically what was corrected, which target species or phases it applied to, the method or equation used, or the source it came from. An answer the list cannot express is also valid here.
Numeric (unit) | A number in the unit named in brackets. Enter the number only — the unit is fixed by the field.
Numeric + unit | A number AND its unit, because the unit varies between procedures. Write both, e.g. '50 us' or '0.5 s'.
Text (free) | Free text. No controlled vocabulary applies.
N/A ¦ None | Offered by every controlled list. "N/A" means the field does not apply to this procedure; "None" means it applies but nothing was used. Prefer either over leaving the cell blank.
Keyed By (Rule 7) | What the field's value repeats over
(none) | Scalar — one value per procedure or analysis. The default.
defines: monitored property per target species | ENUMERATES the monitored property domain — the header of the child table, not a column in it — and does so once per target species. One value per chemical species determined, at whatever granularity the procedure determines it.
defines: reported property | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
defines: sample | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it.
defines: sample > sampling unit | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per subdivision of the physical sample that carries its own row — grain, spot, aliquot, phase, sub-volume.
defines: standard | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per reference material or reference database entry.
defines: target material | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per material type the procedure is designed to analyse, as listed in Target Material.
defines: target species | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per chemical species determined, at whatever granularity the procedure determines it.
monitored property
reported property | One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
sample
sample > sampling unit | Nested — sampling unit exists only within sample. One value per subdivision of the physical sample that carries its own row — grain, spot, aliquot, phase, sub-volume.
standard x reported property | Cross-product, read as "for each standard, one value per reported property". One value per reference material or reference database entry. + One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
target material | One value per material type the procedure is designed to analyse, as listed in Target Material.
target material x monitored property | Cross-product, read as "for each target material, one value per monitored property". One value per material type the procedure is designed to analyse, as listed in Target Material.
target material x target species | Cross-product, read as "for each target material, one value per target species". One value per material type the procedure is designed to analyse, as listed in Target Material. + One value per chemical species determined, at whatever granularity the procedure determines it.
target species | One value per chemical species determined, at whatever granularity the procedure determines it.