## TAPP
Metadata Item | Description | Procedure-Level Tier | Analysis-Level Tier | Data Type | Example / Allowed Content | Comments | Last Update | Keyed By | Purpose | Literature Assessment | Budde+etal2016 ¦ Neptune Plus ¦ IfP Münster | Craddock+etal2008 ¦ Thermo NEPTUNE ¦ WHOI | Hopp+etal2021 ¦ Neptune (Plus spec) ¦ Univ Chicago | Hu+etal2022 ¦ Neptune Plus ¦ Univ Chicago | IbanezMejia+Tissot2020 ¦ Nu Plasma II ¦ MIT | Nie+Dauphas2019 ¦ Neptune ¦ Univ Chicago | Nowell+etal2008 ¦ Neptune ¦ Durham AHIGL | Nowell+etal2008 ¦ Nu Plasma ¦ NIGL | Pringle+Moynier2017 ¦ Neptune Plus ¦ IPGP | Schönbächler+etal2025 ¦ Neptune Plus ¦ ETH Zurich | vanKooten+etal2026 ¦ Thermo Neoma ¦ Univ Copenhagen | Broussard+etal2026 ¦ Neptune Plus ¦ WUSTL | Barnes+etal2025 ¦ Neptune Plus ¦ WUSTL | Barnes+etal2025 ¦ Neptune Plus ¦ ETH Zurich
1. Procedure Identification | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Procedure Name | A short descriptive name for this analytical procedure, including a version number. Should identify the instrument, the technique, and the scope of what is measured — target species, material, or feature. | Basic | Read-Only | Text (free) |  | Source: Core module | 2026-09-01 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Technique | Top-level analytical technique identifier. | Basic | Read-Only | Controlled list | Solution MC-ICP-MS | Source: Core module | 2026-06-18 | (none) |  |  | Solution MC-ICP-MS | Solution MC-ICP-MS | Solution MC-ICP-MS | Solution MC-ICP-MS | Solution MC-ICP-MS | Solution MC-ICP-MS | Solution MC-ICP-MS | Solution MC-ICP-MS | Solution MC-ICP-MS | Solution MC-ICP-MS | Solution MC-ICP-MS | Solution MC-ICP-MS | Solution MC-ICP-MS | Solution MC-ICP-MS
Procedure Author | Person(s) or laboratory responsible for developing and registering this procedure. ORCID recommended for individuals. | Basic | Read-Only | Text (free) |  | Source: Core module | 2026-07-24 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Laboratory | Name of the laboratory or institution hosting the instrument. | Basic | Editable | Text (free) |  | Source: Core module | 2026-06-18 | (none) |  |  | Institut für Planetologie, University of Münster | Woods Hole Oceanographic Institution | University of Chicago | University of Chicago | Massachusetts Institute of Technology | University of Chicago | Arthur Holmes Isotope Geology Laboratory, Durham | NERC Isotope Geosciences Laboratory (NIGL) | Institut de Physique du Globe de Paris | ETH Zurich; sample digestion and separation at Tokyo Institute of Technology | Centre for Star and Planet Formation, Globe Institute, University of Copenhagen | Washington University in St. Louis | Washington University in St. Louis | Institute of Geochemistry and Petrology, ETH Zurich
Laboratory ID | Persistent identifier for the laboratory (e.g., ROR ID). | Advanced | Editable | URI / DOI |  | Source: Core module | 2026-06-18 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Procedure Start Date | First date this procedure configuration was used in production. | Basic | Read-Only | Date |  | Source: Core module | 2026-07-24 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Funding Source for Procedure Development | Grants and other funding sources that supported instrument acquisition, major upgrades, procedure development, and associated personnel time. Include grant numbers and funding agencies where applicable. | Advanced | Read-Only | Text (free) |  | Source: Core module | 2026-07-24 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Procedure Reference(s) | DOI or URL for peer-reviewed publications or technical reports describing, validating, or benchmarking this procedure. | Advanced | Read-Only | URI / DOI |  | Source: Core module | 2026-07-24 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Procedure DOI | DOI or unique persistent identifier for the registered procedure used in this analysis, generated upon procedure submission to a data repository. If a DOI has been applied for but not yet minted, enter "pending". | N/A | Basic | URI / DOI |  | Source: Core module | 2026-08-25 | (none) | Encourages formal procedure registration and ensures traceability. |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Session Identifier | Identifier for the analytical session this record describes — the laboratory's own run, sequence or batch identifier as generated by the instrument or acquisition software. Distinct from any persistent identifier a repository mints on submission. | N/A | Basic | Text (free) |  | Source: Core module | 2026-08-25 | (none) | The analysis record corresponds to one session, which may cover several samples, and this is the link back to the raw instrument files. |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Analyst | Name(s) of the analyst(s) who performed the analysis session. ORCID is recommended for persistent identification. | N/A | Basic | Text (free) |  | Source: Core module | 2026-08-25 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Analysis Start Date | Date on which the analytical session began. For sessions spanning multiple days, use the date of the first session. Format: YYYY-MM-DD. | N/A | Basic | Date |  | Source: Core module | 2026-08-25 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Analysis End Date | Date on which the analytical session ended. May equal Analysis Start Date for single-day sessions. Format: YYYY-MM-DD. | N/A | Basic | Date |  | Source: Core module | 2026-08-25 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Funding Source for Analysis | Grants and other funding sources that supported the specific sample analyses and data collection reported here. Include grant numbers and funding agencies. Distinct from procedure development funding, which may have been awarded years earlier. | N/A | Basic | Text (free) |  | Source: Core module | 2026-08-25 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Coupled Technique(s) | Other analytical techniques applied to the same sample(s) whose results are intended to be interpreted together with data from this procedure. Document coupling with any technique whose results are functionally linked to this dataset — providing calibration inputs, complementary spatial context, or required companion measurements. Name the instrument, laboratory or purpose alongside each technique where known, and join multiple techniques with '; '. Enter "None" if no coupling is intended. | Advanced | Editable | Text (free) | e.g., 'EPMA (Brown University CAMECA SX-100)' ¦ 'FIB-SEM (Crossbeam 540) for 3D tomography; EDS for mineral analysis' ¦ 'NanoSIMS; STXM; Raman' ¦ 'None' ¦ 'N/A' | Source: Core module | 2026-08-31 | (none) |  |  | TIMS — Ba isotopes on a Thermo Scientific Triton Plus at the same institute; Hf-W on the same sample digestions | Laser-ablation MC-ICP-MS — the same NEPTUNE, with a NewWave UP213 laser, "such that laser ablation and solution aspiration can be operated simultaneously" | Prior Pt, Mo, Ni and/or W isotope analyses on the same digestions | N | ID-TIMS U-Pb on an Isotopx X-62 at MIT, and solution Q-ICP-MS (Agilent 7700) for Zr and Hf concentrations, on aliquots of the same dissolutions | N | N | N | N | Mg, K, Ca, Ti, Cr, Fe, Cu, Zn, Mo and Nd isotope data "all obtained from the same sample digestions and are therefore directly comparable" | ICP-MS for Sr and Rb weathering assessment; Si isotopes on a separate NaOH-fusion aliquot | Laser-fluorination oxygen isotopes on a Thermo Finnigan MAT 253 Plus; K loss monitored by Thermo Fisher iCAP Q ICP-MS | High-resolution ICP-MS (Thermo Element XR) at LLNL for bulk elemental abundances, on splits of the same digest | Coordinated dissolution shared with the WUSTL K/Cu/Zn procedure; SIMS oxygen isotopes
Coupling Description | Description of how this procedure is coupled with the technique(s) listed above. Include: (1) the functional relationship — what data or context flows between techniques, or how results are combined (e.g. which output from the coupled technique serves as input to data reduction for this technique); and (2) the analytical sequence — which technique is performed first and why (e.g. non-destructive before destructive). Required when Coupled Technique(s) is not "None". | Advanced | Editable | Text (free) |  | Source: Core module | 2026-08-25 | (none) |  |  | N | Functional: the laser is connected directly to the spray chamber so ablated particles mix with 2% HNO3 and are "effectively analyzed as a wet plasma ensuring that ablated aerosols are closely matrix-matched to solution standards". Sequence: interchangeable — "Our setup allows for interchangeable bulk and in situ S isotope measurement" | N | N | Functional: 3 M HCl washes from the U-Pb anion chemistry were collected and became the Zr aliquots, so the same crystal yields a U-Pb date and a Zr isotopic composition. Sequence: U-Pb purification first, Zr purification from its washes | N | N | N | N | N | N | Functional: pre-cut and post-cut fractions either side of the K collection were measured by Q-ICP-MS "to monitor for K loss during column chemistry". Sequence: Q-ICP-MS check before MC-ICP-MS measurement | N | N
Coupled Procedure DOI | Registered procedure DOI for the coupled technique named above. If the coupled procedure has not yet been registered, enter the DOI of a publication describing the coupled method, or "pending". Enter "None" if no coupling is planned. | N/A | Advanced | URI / DOI |  | Source: Core module | 2026-08-25 | (none) | Provides a stable, citable link to the companion method independent of whether a dataset has been deposited. |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Coupled Dataset or Publication Reference | DOI or other persistent identifier for the co-registered dataset or publication where both datasets are reported together. Accepts: a dedicated dataset DOI (if separately deposited), a shared dataset DOI (if co-submitted in the same package), or a publication DOI. Use "same submission" if the coupled dataset is included in this data package, or "pending" if not yet assigned. If coupling is documented through a shared sample identifier only, that information is already captured in Sample Persistent Identifier (Group 2). | N/A | Advanced | URI / DOI / Text |  | Source: Core module | 2026-08-25 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
2. Samples | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Target Material | General description of the material type(s) this procedure is designed to analyse. | Basic | Read-Only | Controlled list / Text | Basalt ¦ Chondrite ¦ Seawater ¦ Mineral separate ¦ Synthetic solution ¦ N/A ¦ None | Source: Core module | 2026-08-30 | (none) | Used for discoverability and procedure matching, and because the material type constrains sample preparation, calibration and matrix-matching requirements. |  | Chondrules, matrix separates and bulk rock of the Allende CV3 chondrite | Sulfate minerals (anhydrite, barite, gypsum) and sulfide minerals (pyrite, chalcopyrite) | Iron meteorites and terrestrial basalt geostandards | Calcium-aluminium-rich inclusions (CAIs) | Single zircon and baddeleyite crystals, and bulk rock | Silicate rocks — geostandards including basalts, granites and peridotites, and the Allende chondrite | Osmium isotope reference material solutions | Osmium isotope reference material solutions | Whole-rock terrestrial igneous rocks, chondrites, achondrites and Apollo lunar samples | Ryugu returned samples, carbonaceous chondrites, eucrites and terrestrial rock reference materials | Bulk chondrite powders | CI chondrite Oued Chebeika 002 and geostandard | Bennu returned sample aggregate | Bennu returned sample aggregate
Sample Name | Name or identifier of each sample analysed in this session, as used in the laboratory — a sample mount, section or aliquot counts as one entry where that is what the laboratory tracks. Should match the identifier used in associated publications or data tables. | N/A | Basic | Text (free) | e.g., 'Allende MS-1 digest A' ¦ 'NWA 8657 aliquot 2' ¦ 'BHVO-2 replicate 3' | Source: Core module | 2026-08-25 | defines: sample | The analysis record corresponds to one session and may cover several samples. |  | Three matrix separates, six chondrule fractions (C2, C3, C4; C3m, C3i, C3n) and two bulk rock samples of Allende; BHVO-2 | IAEA-S-1, S-2, S-4, NBS-123; in-house standards S_Alfa and S_Spex; anhydrite mineral standard Sch-M-2; pyrite FVG-1 | Toluca, Gibeon, Duchesne, Skookum, Tlacotepec and 18 further iron meteorites; BHVO-2, BCR-2; IRMM-524a | Group II CAIs including FG-FT-4, FG-FT-8 and FG-FT-9 | FC-1 zircon and baddeleyite crystals; ZrNIST reference solution | BHVO-2, BCR-2, BE-N, W-2, AGV-2, GSR-1, GS-N, G-A, G-3; DTS-2b and PCC-1 synthetic mixes; Allende; NIST SRM984 | UMd, DTM, LOsST and DROsS Os isotope reference materials | DTM and LOsST Os isotope reference materials | GS-N, AGV-2, BCR-2, BHVO-2, EW9309 10D, AHANEMO2 D20B; Allende (duplicate splits); NIST SRM984 | Ryugu A0106, A0106-A0107 and C0108; Tagish Lake, Tarda, Ivuna (PB and high PT), Orgueil, Murchison, Colony; eucrites Bouvante and Bereba; BHVO-2, BCR-2, AGV-1, SCo-1; NIST SRM 3169 | BHVO2 and DTS-2b processed alongside the samples | Oued Chebeika 002; geostandard BHVO-2; NIST SRM 3141a | OREX-803015-101 (LLNL split) and OREX-803015-100 (ETH split) of Bennu aggregate; BHVO-2 | OREX-803015-100, a 5.2 mg aliquot of Bennu aggregate
Sampling Unit Type | The kind of physical subdivision of the sample to which one row of reported values corresponds — the unit that is analysed and reported, as distinct from the sample as a whole. Where units nest (e.g. confined tracks within grains), state both levels. The units themselves are listed in Sampling Unit Name. | Basic | Read-Only | Controlled list / Text | Whole sample ¦ Aliquot ¦ Grain ¦ Spot ¦ Analysis point ¦ Phase ¦ Sub-volume ¦ Region of interest ¦ N/A ¦ None | Source: Core module | 2026-09-15 | (none) |  |  | Digestion aliquot — "All samples (0.3–0.5 g) were digested in closed Savillex beakers"; chondrule fractions "comprise between 155 and ~3000 chondrules each" | Purified solution aliquot — "Less than 50 mg of sample was accurately weighed"; "A precise solution volume, corresponding to 500 µg of S" taken for column purification | Solution aliquot of a digestion — "the Fe isotopic compositions were analyzed on solution aliquots (~1-2 mg Fe) of digestions"; five meteorites cut as "~50 mg pieces" | Fraction of a CAI digestion — "Approximately 30% of the matrix cut", "equivalent to 24% fraction of the whole CAI" | Single crystal — "Single zircon and baddeleyite crystals selected for analysis were individually handpicked"; each "individually loaded into clean PFA microcapsules" | Digestion aliquot — "Samples of about 100 mg or less were digested" | Reference material solution aliquot — 200 ng/ml to 2.5 µg/ml Os, ~300 µl consumed per analysis | Reference material solution aliquot — ~6400 µl consumed per analysis | Weighed powder aliquot — "An aliquot of <=125 mg of powdered sample was weighed depending on the Rb concentration of the sample; masses were calculated to yield >20 ng Rb" | Digestion aliquot — Ryugu "aliquots of <25 mg were analyzed with ~40 to 70 ng Zr"; Tagish Lake 30 mg, Tarda 90 mg, Ivuna 40 and 44 mg "from a larger homogenized powder (550 mg)" | Fraction of a bulk digestion — "Another 5% fraction was used to determine Al/Mg ratios by multi-collector (MC)-ICPMS" | N | Split of a single digest — "The solution was then split two ways: about half stayed at WUSTL and half was sent to Lawrence Livermore National Laboratory ... the aliquot was further split into two aliquots" | A 5.2 mg aliquot of Bennu aggregate
Sampling Unit Name | The name or label of each sampling unit analysed in this session, as the laboratory records it, together with the sample it belongs to — e.g. a spot number, a grain label, a map or region-of-interest name, or an aliquot identifier. Where units are too numerous to name individually, such as map pixels or reconstructed voxels, name the acquisition area they belong to instead. | N/A | Basic | Text (free) | e.g., NWA 8657-1: spots 1–24 ¦ Zircon mount ZM-3: grains Z1–Z30 ¦ Core 73001: map areas A1–A5 ¦ Bennu aggregate OREX-803015-100: aliquot for Ti | Source: Core module | 2026-09-15 | defines: sample > sampling unit | Gives every field keyed by sampling unit a row to attach its values to, and ties each unit to the sample it belongs to; without it those fields collapse to one value per session. |  | Labelled within Allende: bulk "MS-A", "MS-B"; matrix "M1", "M2", "M3"; chondrules "C1", "C2", "C3m", "C3n", "C3i", "C4" (Table 1, p.3). "C3b" and "C2–C4c" in the same table are weighted means, not units. BHVO-2 by name only | Sample name only — Table 3 lists "IAEA-S-1", "IAEA-S-2", "IAEA-S-4", "NBS-123" and the in-house standards Alfa and Spex (p.4); replicates are counted ("# of replicates"), not labelled | Sample name only — meteorites by name, e.g. "Toluca, Gibeon, Duchesne, Skookum, Tlacotepec" (p.5); the digestion aliquots and ~50 mg pieces carry no labels of their own | Labelled: each CAI by name under its specimen number — Table 1's "Sample" and "CAI name" columns give e.g. ME-3364-25.2 "FG-FT-3", ME-2639-16.2 "FG-FT-4", AL3S5 "FG-FT-8", AL4S6 "FG-FT-9", AL8S2 "FG-FT-10" (p.3); † marks a second analysis of FG-FT-4, -8 and -9. BCR-2 by name only | Labelled: single crystals of FC-1 — zircons "z1"–"z16" with "3_z…" and "4_z…" series (e.g. "4_z15R", "4_z17"), baddeleyites "b1"–"b8" with "3_b…" and "4_b…" series, and bulk rock "WR1" (Table 1, pp.4–6); "CA"/"Untr." beside each label marks chemical abrasion, not identity | Labelled for the lunar rocks by Apollo sample and split number: "12002.613", "12018.301", "12052.353", "10017.413", "74275.361", "77215.276" (Table 1, p.2), the last a "white-colored fragment"; terrestrial rocks by name only (BCR-2, BHVO-2 …) | Labelled for LOsST only: "LOsST 17-03-06 (Aliq 1)" (Table 8b, p.22) — the Neptune and Nu Plasma runs "were made on two different aliquots of the LOsST RM" (p.24); UMd, DTM and DROsS by name and session date | Labelled for LOsST only: "23-05-06NIGL (Aliq 2)" (Table 8b, p.22), the second of the "two different aliquots of the LOsST RM" (p.24); DTM by name only | Labelled for the one duplicated sample: "Allende I" and "Allende II" (Table 1, p.4), "duplicate splits from the same powder aliquot" (p.3); every other sample by name only | Labelled by digestion where a sample was digested more than one way: "Ivuna PB" and "Ivuna high PT", "30 mg Tagish Lake (labeled high PT)" and "90 mg Tarda (high PT)" (p.4); Ryugu A0106, A0106-A0107 and C0108 and the other samples by name only | Labelled where a meteorite was analysed twice: "NWA 16569 (1)", "NWA 16569 (2)", "NWA 16554 (1)", "NWA 16554 (2)" (Table 1, p.2); the other chondrites by name only | Labelled: two solutions of the OC002 fragment — "Approximately 7 mg of sample from the LAB24-2 OC002A and OC002B solutions ... were used for potassium stable isotope analysis" (p.3) | Labelled: split "OREX-803015-0" — "An ~20.66 mg split of Bennu aggregate (OREX-803015-0) was dissolved at WUSTL" (p.7); the half of the solution kept at WUSTL is given no identifier of its own. The paper states its scheme: splits take "suffixes of -100, -101, -102" on the parent's number (p.7) | Labelled: aliquot "OREX-803015-100" — "a 5.2 mg aliquot of Bennu aggregate (OREX-803015-100)" (p.7)
Sample Persistent Identifier | Globally unique, persistent identifier for each sample listed in Sample Name. IGSN (International Geo Sample Number) is the recommended standard for geological and cosmochemical samples, as used by Astromat, EarthChem and SESAR. Where a sample and its sub-samples are separately registered, record the identifier at the level actually analysed. | N/A | Advanced | URI / IGSN |  | Source: Core module | 2026-09-15 | sample |  |  | N | N | N | N | N | N | N | N | N | N | N | N | Curatorial sample identifiers stated: OREX-803015-101 (LLNL) and OREX-803015-100 (ETH Zurich) | OREX-803015-100
Sample Description | Brief description of sample provenance, form, or preparation state relevant to this analysis. | N/A | Basic | Text (free) |  |  | 2026-06-18 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Sample Preparation Method | The form in which the sample is presented to the instrument, and the preparation that brought it to that form — for example mounting, sectioning, polishing, coating, crushing, fusion, or extraction of an electron-transparent section. Record 'None' where the material is analysed as received. | Basic | Editable | Controlled list / Text | Whole-rock powder ¦ Mineral separate ¦ Biogenic material (cleaned) ¦ Natural water ¦ Leachate or partial-dissolution fraction ¦ Dissolved residue ¦ Pre-digested liquid ¦ N/A ¦ None | Source: Core module | 2026-08-27 | (none) |  |  | Chondrule, matrix and bulk rock separates; preparation detailed in the supplementary material | Mineral standard cut as a 2 mm thick section, polished and mounted on a 45x25 mm petrographic slide for the laser half; solution half dissolved from weighed mineral | Iron meteorite pieces "cut using a diamond saw, polished with SiC abrasive paper, and cleaned in ethanol" | N | Crushing in a stainless steel mortar, sieving through 375 µm plastic mesh, washing in a plastic gold pan, hand magnet, Frantz LB-1 magnetic separation, methylene iodide heavy liquid, hand picking under high-purity ethanol | Whole-rock powder | N/A — reference material solutions, no solid preparation | N/A — reference material solutions, no solid preparation | "Whole rock samples were crushed by hand using an agate mortar until a fine powder was obtained. A minimum of 0.5 g of terrestrial rock or meteorite and 100 mg of lunar samples was crushed in order to avoid non-representational sample analysis" | Homogenised powder — Ivuna aliquots taken "from a larger homogenized powder (550 mg)" | Bulk powder; for the Si aliquot, NaOH fusion in silver crucibles | N | N | N
Digestion Acid(s) | Acid mixture used for sample dissolution. Record acid type(s) and concentration(s) where known (e.g., concentrated HF–HNO3 in Parr bombs is standard for silicate dissolution). | Basic | Read-Only | Text (free) | e.g., HF–HNO3 ¦ HF–HNO3–HClO4 ¦ HNO3–HCl ¦ HNO3 only ¦ Aqua regia ¦ N/A ¦ None | Source: Solution Introduction module | 2026-08-31 | preparation step | For isotope-ratio work, complete matrix dissolution is required to avoid isotopic fractionation during partial dissolution. |  | "HF–HNO3(–HClO4), followed by inverse aqua regia" | 5 ml HNO3 (50%), then 3 ml concentrated HNO3 + 2 mL HCl (50%); residue dissolved in 4 mL 2% HNO3 | Iron meteorites: aqua regia (3:1 HCl-HNO3). Basalts: HF-HNO3 (2:1) followed by several steps of aqua regia. All converted to chloride and redissolved in 0.25 ml 10 M HCl | HF/HNO3 in 3:1 proportion with a few drops of HClO4, then -- after evaporation to dryness -- a 2:1 mixture of HCl:HNO3; dried down and dissolved in concentrated HNO3, diluted in 3 M HNO3 and centrifuged. The first HF/HNO3-HClO4 attack was missing from this cell before 2026-09-08. | 29 M HF; after conversion to a chloride matrix for U-Pb. Zr aliquots taken up in 3 M HNO3 + 0.5 M HF | Three steps of concentrated HF–HNO3–HCl–HClO4; step (i) "4 ml 28 M HF + 2 ml 15 M HNO3 + 1 ml 10 M HClO4" | N/A — reference material solutions in 3 or 5 mol/l Teflon-distilled HCl | N/A — reference material solutions in Teflon-distilled 3 mol/l HCl | "a mixture of concentrated HF/HNO3"; after evaporation "6N HCl was added" to dissolve fluoride complexes | Concentrated HF-HNO3, then a HNO3-HCl mixture, then a HNO3-H2O2 mixture; Ivuna high PT: concentrated HF-HNO3 for 3 days then concentrated HCl for 2 days | Cr/Mg route: 6 M HCl loading, 10 M HCl pretreatment, 0.5 M HCl, 0.5 M HNO3, 1 M HF, 6 M HCl elutions; Si route: NaOH fusion then Milli-Q water and HNO3 | N | "concentrated HF and HNO3 in a 3:1 ratio", followed by fluxing in concentrated HNO3 and HCl with 1 ml H2O2 added to remove organics; brought up in 5 ml 0.5 M HNO3 | Coordinated dissolution shared with the WUSTL split — see the WUSTL column
Digestion Vessel Type | Type of vessel used for acid digestion. | Basic | Read-Only | Text (free) | e.g., Savillex beaker ¦ Parr bomb ¦ TFE/TFM bomb ¦ Carius tube ¦ HPA-S ¦ Open beaker ¦ N/A ¦ None | Source: Solution Introduction module | 2026-08-31 | (none) | Vessel type constrains maximum temperature and pressure and determines suitability for refractory mineral dissolution (e.g., Parr bombs required for complete silicate digestion at high temperatures). |  | "closed Savillex beakers" | 15 ml PTFE digestion vessel | Hot plate, closed vessel not specified beyond "on a hot plate" | N | "clean Teflon microcapsules" inside "a large-volume Parr digestion vessel" | 30 ml fluoropolymer vessel | N/A | N/A | "closed Teflon bombs" | Octagonal-body Savillex vials; Parr bomb (Ivuna PB, BCR-2, AGV-1); hotplate (BHVO-2) | N | N | "in a closed beaker" | N
Digestion Step | Each distinct acid digestion step applied to dissolve the sample, listed in the order performed and named by its attack. A step is distinct when its acid mixture, vessel, temperature or duration differs from the one before; an evaporation or dry-down that carries no attack of its own is part of the step it follows, and an identical attack repeated on the residue is a repeat of that step rather than a new one. Enumerating the steps is what allows Digestion Acid(s), Digestion Temperature and Digestion Duration to be recorded per step. Record N/A where the sample is introduced without acid digestion. | Basic | Read-Only | Text (free) | e.g., '1: HF-HNO3, Parr bomb ¦ 2: aqua regia reflux of the residue' ¦ '1: HF-HNO3-HClO4 ¦ 2: HNO3-HCl ¦ 3: HNO3-H2O2' ¦ 'Single step' ¦ 'N/A (no acid digestion)' | Source: Solution Introduction module | 2026-09-08 | defines: preparation step | Naming each step is what lets acid, temperature and duration be read per step; a bare count cannot say which conditions belong to which attack, and cannot settle whether a repeated attack or an intervening evaporation counts. |  | 1: HF-HNO3 (-HClO4), closed Savillex beakers on a hotplate ¦ 2: inverse aqua regia. Both steps stated; conditions not given beyond 'on a hotplate'. | 1: 5 ml HNO3 (50%), hot plate below 70 deg C, taken to dryness ¦ 2: 3 ml concentrated HNO3 + 2 ml HCl (50%), sealed PTFE vessel, 70 deg C, taken to dryness. The subsequent 4 ml 2% HNO3 is the final uptake, not a step. | Iron meteorites, 1: aqua regia (3:1 HCl-HNO3), 120 deg C, 24 h on a hot plate. Basalts, 1: HF-HNO3 (2:1), 150 deg C, 48 h on a hot plate ¦ 2: 'several steps of aqua regia', number not stated. Both routes then converted to chloride and taken up in 0.25 ml 10 M HCl. | 1: HF/HNO3 in 3:1 proportion with a few drops of HClO4, hot plate 160 deg C, 2 weeks ¦ 2: evaporated to dryness and redissolved in a 2:1 mixture of HCl:HNO3, 1 week on a hot plate. "These steps were performed twice to ensure complete digestion" -- the PAIR is repeated, so under the 2026-09-08 grain rule the members are 2, not 4. The subsequent concentrated HNO3 and 3 M HNO3 are dry-down and uptake, not steps. Resolves the cell left open on 2026-09-08. | Untreated crystals, 1: 29 M HF in a PFA microcapsule inside a Parr vessel, 215 deg C, 48 h. Chemically abraded crystals (19 zircons), 1: 12-h partial dissolution in 29 M HF at 215 deg C under pressure, after annealing at 900 deg C for 60 h ¦ 2: the same complete digestion. The annealing is a pre-treatment, not a digestion step. | 1: 4 ml 28 M HF + 2 ml 15 M HNO3 + 1 ml 10 M HClO4 ¦ 2: not stated ¦ 3: not stated. The paper numbers three steps of concentrated HF-HNO3-HCl-HClO4 but gives the composition only of step (i). | N/A - reference material solutions, no digestion. | N/A - reference material solutions, no digestion. | 1: concentrated HF/HNO3, closed Teflon bombs, 130 deg C, >48 h ¦ 2: after evaporation of the HF/HNO3, 6N HCl at 130 deg C to dissolve fluoride complexes. Samples were then evaporated to dryness and were ready for chemistry. | Main route 1: concentrated HF-HNO3 ¦ 2: HNO3-HCl ¦ 3: HNO3-H2O2. Ivuna high-PT route 1: concentrated HF-HNO3, 3 days ¦ 2: concentrated HCl, 2 days. | Bulk and chondrule route, 1: 3:1 7 M HNO3 : 28 M HF in Parr bombs, 3 days (1 day at 150 deg C, 2 days at 210 deg C) ¦ 2: dried down and taken up in aqua regia, 2 further days on a hotplate. Si route, 1: NaOH fusion in silver crucibles, 720 deg C, 13 min, the fusion cake dissolved in Milli-Q water and acidified with HNO3 -- a fusion rather than an acid digestion. | 1: 3:2 concentrated HF : double-distilled HNO3 in PFA vials, hotplate 150 deg C, about 1 week ¦ 2: dried down, dissolved in 1 ml concentrated HNO3, with 1 ml H2O2 added slowly in 0.1 ml increments ¦ 3: dried down, dissolved in double-distilled HCl. The final 5 ml 2% HNO3 is the uptake, not a step. | 1: concentrated HF and HNO3 in a 3:1 ratio, closed beaker, 170 deg C, 48 h ¦ 2: fluxing in concentrated HNO3 and HCl, with 1 ml H2O2 added slowly during the HNO3 flux to remove organics. The 5 ml 0.5 M HNO3 is the uptake, not a step. | Coordinated dissolution shared with the WUSTL split - see the WUSTL column.
Digestion Temperature | Temperature at which acid digestion was carried out (°C). | Basic | Editable | Numeric (°C) | e.g., '90–120 °C (hotplate, Savillex beaker)' ¦ '150–190 °C (Parr bomb, refractory phases)' ¦ '220 °C (HPA-S)' | Source: Solution Introduction module | 2026-08-25 | preparation step |  |  | N | "less than 70 °C" for the first evaporation; 70 °C for the total digestion | Iron meteorites 120 °C; basalts 150 °C | 160 deg C (hot plate) | 215 °C | N | N | N | 130 °C for both the HF/HNO3 and the 6N HCl steps | 180 °C (hotplate), 120 °C (HNO3-HCl), 220 °C (Savillex, Tagish Lake and Tarda), 170 °C (Parr bomb, Ivuna PB), 160 °C (Ivuna high PT) | 150 deg C for 1 day then 210 deg C for 2 days (Parr bomb); hotplate for the aqua regia step; 720 deg C for the NaOH fusion (Si route). The 130 deg C 10 M HCl step previously recorded here is Cr(VI) speciation during column chemistry, not a digestion. | 150 deg C (hotplate). The 70 and 140 deg C previously recorded here belong to the carbonate-removal pre-treatment and to the cosmogenic-radionuclide dissolution, which are different preparations in the same paper. | 170 °C | N
Digestion Duration | Duration of the primary acid digestion step (hours or days). | Basic | Editable | Text (free) | e.g., '12 h' ¦ '24 h' ¦ '48 h (refractory accessory phases)' ¦ '3 days (multi-step with intermediate dry-down)' | Source: Solution Introduction module | 2026-08-25 | preparation step |  |  | N | Not stated for the individual steps beyond "taken to dryness" | Iron meteorites 24 hours; basalts 48 hours | 2 weeks for the HF-HNO3-HClO4 step and 1 week for the HCl-HNO3 step, the pair performed twice | 48 h at 215 deg C in a Parr vessel. The 60 h previously recorded here is the chemical-abrasion ANNEALING at 900 deg C, not a digestion; the abrasion leach itself is 12 h. | N | N | N | ">48 h" | 3–7 days (hotplate), 12 h (HNO3-HCl), "about a week" (Tagish Lake and Tarda), 3 days + 2 days (Ivuna high PT) | 3 days in the Parr bomb (1 + 2 days) then 2 days in aqua regia; 13 min for the NaOH fusion. The 3 h and >1 week previously recorded here are Cr speciation during column chemistry. | About 1 week for the HF-HNO3 step. The 20 h previously recorded here is the cosmogenic-radionuclide dissolution, a different preparation. | 48 h | N
Final Solution Matrix | Acid type and concentration of the final dissolved sample solution introduced to the instrument. For isotope-ratio procedures, match acid molarity precisely between samples and bracketing standards. Procedures targeting HFSE may add trace HF to the final solution to keep those elements in solution. | Basic | Read-Only | Text (free) | e.g., 0.3 M HNO3 ¦ 2% HNO3 ¦ 0.5 M HCl ¦ 3 M HCl ¦ N/A ¦ None | Source: Solution Introduction module | 2026-09-08 | acquisition pass | Directly affects instrument response and ion transmission. Mismatched molarity introduces matrix-induced mass bias offsets. |  | N | 2% (w/w) HNO3, 50 ppm S stock | 0.3 M HNO3 (measured at 10 µg/g Fe in 0.45 M HNO3); all sample and standard solutions "prepared with the same 0.3 M HNO3 solution" | 15–25 ppb for the most abundant isotope | 0.59 M HNO3 + 0.28 M HF, samples and bracketing standards matched in matrix and at 60 ng/g total Zr | 0.3 M HNO3, ~15–25 ppb Rb | 3 or 5 mol/l Teflon-distilled HCl | 3 mol/l Teflon-distilled HCl | 0.1N HNO3 | 0.5 M HNO3 - 0.005 M HF at 30 ppb Zr (also 17 and 60 ppb) | 0.5 M HNO3 (Cr); 6 M HCl elution of the final Cr cut | 300 ppb K solution | 200 ppb for K and Zn; 100 ppb for Cu | N
Sample Aliquot Mass or Volume | Mass (mg) of solid material digested or volume (mL) of liquid taken for dissolution. | Advanced | Editable | Numeric (mg or mL) | e.g., '50 mg (whole-rock silicate)' ¦ '25 mg (mass-limited meteorite)' ¦ '100 mg (low-abundance trace elements)' ¦ '1–5 mL (aqueous sample)' | Source: Solution Introduction module | 2026-08-25 | sample | Used for yield calculations and concentration back-calculation. |  | 0.3–0.5 g digested; ~100 ng Mo consumed per measurement | <50 mg weighed; 500 µg S taken for column purification | ~1-2 mg Fe per analysis; ~50 mg meteorite pieces | N | Single crystals; ~50 µl (5% of sample) taken for concentration measurement | ~100 mg or less; ~40 ng Rb typical | Up to 600 ng Os consumed per analysis; ~300 µl of solution | ~6400 µl of solution per analysis | <=125 mg powder, calculated to yield >20 ng Rb | Ryugu <25 mg with ~40–70 ng Zr; 15 ng Zr consumed per 30 ppb analysis | N | N | N | 5.2 mg
Chromatographic Separation Applied | Whether chromatographic separation of element groups (e.g., using ion exchange resin) was performed prior to analysis. Record resin type and elution matrix conditions if applied. Record 'None' if the direct digestion solution is analyzed. | Basic | Read-Only | Controlled list / Text | Yes ¦ N/A ¦ None | Source: Solution Introduction module | 2026-08-31 | (none) | Separation may be required to remove matrix elements that cause variable mass bias, to eliminate isobaric interferences, or to isolate an element group before measurement (e.g., Fe and Cu–Zn for isotope dilution, Hf separation from REE, removal of Cr and Ni from Fe fractions). It is usually mandatory for isotope-ratio procedures and optional for concentration measurement. |  | Yes — two-stage anion exchange for W, with Mo collected in 3 M HNO3 and further purified on Eichrom TRU Resin; Ba separated on AG50-X8 | Yes — cation exchange AG50-X8 (H+ form), 2.5 ml resin, conditioned with 1.4 N HNO3; S passes through while matrix elements are retained. Yield 98±4% | Yes — AG1-X8 (200-400 mesh) anion resin, 3 ml, 10.5 cm PFA columns; repeated with new resin. Overall Fe yield >99% | Yes — U/TEVA, TODGA, then two-step FPLC on Ln-Spec resin (70 cm x 1.6 mm, 1.4 ml of 25–50 µm resin, 94 steps, 188 ml, 16 h at 70 °C, 0.17 ml/min). Overall yields >95% | Yes — AG-1X for U-Pb; Ln-Spec (~300 µl, 25–50 µm) for Zr, giving >95% Zr, undetectable REEs and <3% of initial Hf; TODGA first stage for bulk rocks | Yes — five steps: AG50W-X8 cation, a second cation column, AG1-X8 anion in 2 M HF for Ti, a 40 cm Eichrom Sr resin column for Rb-K, and an AG50W-X8 clean-up. Yields >95% | N/A — reference material solutions | N/A — reference material solutions | Yes — DGA resin Ca removal (1.8 mL), then AG50 X12 (20 mL and 10 mL) in 3N HCl, then AG50 X8 (1 mL) in 0.5N HCl. Reduces K/Rb by a factor of 200 to K/Rb<2 and gives 88Sr/85Rb<0.005 | Yes — four-step separation on anion exchange (AG 1-X8), DGA and LN resin; two-stage anion exchange for Ivuna; three-stage AG 1-X8 + LN for terrestrial samples | Yes — AG1-X8 anion (1 ml) for Fe, then AG50-X12 cation (1 ml) twice for Cr and Mg | Yes — twice through 1.5 mL Bio-Rad AG50W-X8 100–200 mesh cation resin, loading, matrix elution and K elution all in 0.5 M HNO3 | Yes — AG1-X8 200–400 mesh anion resin, 5 ml 1.5 M HBr to elute the matrix and 3 ml 0.5 M HNO3 to elute Zn | Yes — three-step anion exchange chromatography; yields 75–100%
Isotope Dilution Spike | Isotopic spike material added to the sample before or during digestion for isotope dilution (ID) analysis or double-spike mass bias correction. For double-spike procedures, add and equilibrate the spike before digestion. For concentration ID, the spike must be added before digestion for complete equilibration. Record spike name or certificate reference, spike isotopes, and whether added pre- or post-digestion. Record 'None' if ID or double-spike technique is not used. | Basic | Read-Only | Text (free) | e.g., IRMM-3636 (233U–236U double spike) ¦ 117Sn–122Sn double spike (in-house) ¦ 57Fe–58Fe double spike ¦ None ¦ N/A | Source: Solution Introduction module | 2026-08-31 | (none) | Spiking before digestion corrects for isotopic fractionation during dissolution and column chemistry. |  | N/A — no isotope dilution spike; mass bias handled by standard-sample bracketing or internal normalization | N/A — no isotope dilution spike; mass bias handled by standard-sample bracketing or internal normalization | N/A — no isotope dilution spike; mass bias handled by standard-sample bracketing or internal normalization | N/A — no isotope dilution spike; mass bias handled by standard-sample bracketing or internal normalization | In-house 91Zr-96Zr double spike, added at a 0.43:0.57 spike-to-sample Zr mass ratio | N/A — no isotope dilution spike; mass bias handled by standard-sample bracketing or internal normalization | N/A — no isotope dilution spike; mass bias handled by standard-sample bracketing or internal normalization | N/A — no isotope dilution spike; mass bias handled by standard-sample bracketing or internal normalization | N/A — no isotope dilution spike; mass bias handled by standard-sample bracketing or internal normalization | N/A — no isotope dilution spike; mass bias handled by standard-sample bracketing or internal normalization | N/A — no isotope dilution spike; mass bias handled by standard-sample bracketing or internal normalization | N/A — no isotope dilution spike; mass bias handled by standard-sample bracketing or internal normalization | N/A — no isotope dilution spike; mass bias handled by standard-sample bracketing or internal normalization | N/A — no isotope dilution spike; mass bias handled by standard-sample bracketing or internal normalization
3. Instrument & Software | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Instrument Manufacturer | Manufacturer of the instrument that performs the measurement, recorded as a controlled value. Where a procedure couples a sample-introduction system to an analysing instrument, this records the analysing instrument. Instrument Model gives the specific designation. | Basic | Read-Only | Controlled list | Thermo Fisher Scientific ¦ Agilent ¦ PerkinElmer ¦ Nu Instruments ¦ Analytik Jena ¦ Shimadzu ¦ Unknown ¦ N/A ¦ None | Source: Core module | 2026-08-30 | (none) | Recording it as a controlled value lets procedures be found by vendor. |  | Thermo Fisher Scientific | Thermo Fisher Scientific | Thermo Fisher Scientific | Thermo Fisher Scientific | Nu Instruments — "a Nu Plasma II MC-ICP-MS fitted with an enhanced sensitivity interface" | Thermo Fisher Scientific | Thermo Fisher Scientific | Thermo Fisher Scientific | Thermo Fisher Scientific | Thermo Fisher Scientific | Thermo Fisher Scientific | Thermo Fisher Scientific | Thermo Fisher Scientific | Thermo Fisher Scientific
Instrument Model | Model designation of the instrument that performs the measurement, including any generation or configuration suffix. Conventionally written with the manufacturer name included; Instrument Manufacturer records the vendor separately, as a controlled value, so that procedures remain findable by vendor. | Basic | Read-Only | Text (free) | e.g., Thermo Scientific Neptune ¦ Thermo Scientific Neptune Plus ¦ Nu Instruments Nu Plasma II ¦ Nu Instruments Nu Plasma 1700 ¦ Nu Instruments Nu Sapphire ¦ IsotopX Phoenix ¦ N/A ¦ None | Source: Core module | 2026-08-31 | (none) |  |  | Neptune Plus | NEPTUNE ("Thermo Electron NEPTUNE") | Neptune "upgraded to Neptune Plus specifications" | Neptune Plus "with the addition of an OnTool booster" | Nu Plasma II | Neptune | Neptune | Nu Plasma | Neptune Plus | Neptune Plus | Neoma | Neptune Plus | Neptune Plus | Neptune Plus
ICP-MS Type | Type of ICP-MS system - the mass analyser and, where the instrument has one, the collector configuration. | Basic | Read-Only | Controlled list | Multi-collector sector-field (MC-ICP-MS) ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-30 | (none) | Multi-collector sector-field instruments combine a magnetic sector and electrostatic analyser with a fixed array of Faraday cups (and optionally ion counters) that simultaneously collect multiple isotope beams, enabling high-precision static multi-collection. |  | Multi-collector sector-field ICP-MS [all columns state "multi-collector inductively coupled plasma mass spectrometer"] | Multi-collector sector-field ICP-MS [all columns state "multi-collector inductively coupled plasma mass spectrometer"] | Multi-collector sector-field ICP-MS [all columns state "multi-collector inductively coupled plasma mass spectrometer"] | Multi-collector sector-field ICP-MS [all columns state "multi-collector inductively coupled plasma mass spectrometer"] | Multi-collector sector-field ICP-MS [all columns state "multi-collector inductively coupled plasma mass spectrometer"] | Multi-collector sector-field ICP-MS [all columns state "multi-collector inductively coupled plasma mass spectrometer"] | Multi-collector sector-field ICP-MS [all columns state "multi-collector inductively coupled plasma mass spectrometer"] | Multi-collector sector-field ICP-MS [all columns state "multi-collector inductively coupled plasma mass spectrometer"] | Multi-collector sector-field ICP-MS [all columns state "multi-collector inductively coupled plasma mass spectrometer"] | Multi-collector sector-field ICP-MS [all columns state "multi-collector inductively coupled plasma mass spectrometer"] | Multi-collector sector-field ICP-MS [all columns state "multi-collector inductively coupled plasma mass spectrometer"] | Multi-collector sector-field ICP-MS [all columns state "multi-collector inductively coupled plasma mass spectrometer"] | Multi-collector sector-field ICP-MS [all columns state "multi-collector inductively coupled plasma mass spectrometer"] | Multi-collector sector-field ICP-MS [all columns state "multi-collector inductively coupled plasma mass spectrometer"]
Instrument Serial Number or Lab Identifier | Serial number or laboratory-internal identifier for the specific instrument unit. | Advanced | Editable | Text (free) |  | Source: ICP-MS module | 2026-08-25 | (none) | Supports traceability to instrument service records. |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Torch Type | Type of plasma torch installed (e.g. standard quartz, high-matrix, low-flow). | Advanced | Read-Only | Text (free) | Standard quartz ¦ High-matrix ¦ Low-flow ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-30 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Interface Cone Configuration | Geometry and designation of the sampler and skimmer cones installed during analysis. State both cones. | Basic | Read-Only | Text (free) | e.g., Ni H-sampler + Ni H-skimmer (standard) ¦ Ni H-sampler + Ni X-skimmer ¦ Ni Jet-sampler + Ni X-skimmer (with interface pump) ¦ N/A ¦ None ¦ For Neptune-type instruments, H-cones (standard) and X-cones (high-sensitivity, ~3× higher transmission) are common options. | Source: ICP-MS module | 2026-08-31 | (none) | Jet cones combined with an interface pump upgrade are used on some instruments for enhanced sensitivity. The choice affects ion beam intensity. |  | Standard sample and (H) skimmer cones | X-cones | H skimmer cones | N | N | Normal sampler and skimmer cones | N | N | Sample cone Jet; skimmer cone H | "Normal skimmer and sampler cones were utilized" | A Jet and X cone | N | N | N
Sampler and Skimmer Cone Material | Material composition of the sampler and skimmer cones. State both sampler and skimmer materials; a single statement is acceptable where they are identical. | Advanced | Read-Only | Controlled list / Text | Ni sampler and Ni skimmer ¦ Pt sampler and Pt skimmer ¦ Ni sampler and Pt skimmer ¦ Al sampler and Al skimmer ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-31 | (none) | Nickel is standard for most aqueous and nitric matrices; platinum is used for HCl-rich, high-TDS or organic matrices for its greater corrosion resistance; aluminium is used in some laboratories for high-purity work or enhanced sensitivity. |  | Ni | Ni | "We used Ni or Pt sampler and H skimmer cones ... The main motivation for using Pt cones was an increase in sensitivity and a decrease in the frequency of cone cleaning" | N | N | Ni | N | N | N | N | N | N | N | N
Torch Depth | Distance between the load coil and the sampling cone tip (mm), also called injector depth or torch position depending on the instrument manufacturer. | Advanced | Editable | Text (free) | 5 mm ¦ Standard position (manufacturer default) ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-25 | (none) | Affects ion transmission efficiency, oxide formation, and doubly-charged species production. |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Guard Electrode | Whether a guard electrode - also called a grounded shield electrode, extraction lens, or Fassel torch shield - is installed and active on the torch assembly. | Basic | Read-Only | Controlled list | On ¦ Off ¦ Not installed ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-30 | (none) | Capacitively decouples the plasma from the load coil, reducing secondary discharge and improving ion extraction efficiency. |  | N | "Pt-guard electrode: On, grounded" | N | N | N | N | N | N | N | N | N | N | N | N
Faraday Cup Array Configuration | Description of the multi-collector Faraday cup array. Specify the total number of Faraday cup detectors (fixed and moveable), their labeled positions (e.g., L4, L3, L2, L1, Ax, H1, H2, H3, H4 for a 9-cup Neptune), and the presence and position of any ion counter (SEM) or Daly detector in the collector block. | Basic | Read-Only | Text (free) | Neptune: 9 Faraday cups (L4–H4) + 1 SEM (axial or off-axis); Nu Plasma II: 16 Faraday cups + 3 ion counters ¦ N/A ¦ None | Source: MC-ICP-MS module | 2026-08-30 | (none) | The number and span of cups constrains which isotope combinations can be simultaneously collected. |  | N | Nine Faraday cups — "equipped with nine Faraday Cups" (p.3); Table 1 gives detection system "Faraday cups" and acquisition mode "Static, analogue detectors" (p.3). No ion counter stated | N | N | N | Nine Faraday collectors — "The MC-ICPMS at the University of Chicago is equipped with nine Faraday collectors" (p.8). No ion counter stated | "The Durham Neptune has a 9 Faraday collector array equipped with 10^11 Ω resistor amplifiers which allow a maximum beam of 50 V per channel" (p.3). An SEM ion counter is also present — abundance sensitivity "was determined by scanning the low mass tail of a 30 V 192Os beam using the SEM" (p.3); its position is not stated | Seven Faraday cups — "fitted with a 7 Faraday ‘U–Pb’ collector block and 10^11 Ω resistor amplifiers which permitted maximum beam sizes of 10 V per channel" (p.4). No ion counter stated | N | N | N | N | N | N
Faraday Cup Amplifier Resistor Values | Resistance values (Ω) of the feedback resistors in the Faraday cup amplifiers. Standard amplifiers use 10¹¹ Ω resistors, yielding 1 V per ~6.24 × 10⁶ ion counts per second. Report the resistor value per cup position, or note 'all 10¹¹ Ω' if uniform. | Basic | Read-Only | Text (free) | e.g., All cups: 10¹¹ Ω ¦ L1: 10¹³ Ω (234U); all others: 10¹¹ Ω ¦ N/A ¦ None | Source: MC-ICP-MS module | 2026-09-10 | monitored property | High-gain amplifiers (10¹² or 10¹³ Ω) are fitted to selected cups to improve signal-to-noise for very low-intensity beams (e.g., 234U, low-abundance spike isotopes). |  | N | N | 10^10 Ω for 56Fe+; 10^11 Ω for 54Fe, 57Fe, 58Fe; 10^12 Ω for the 53Cr and 60Ni interference monitors | N | N | "All three collectors were equipped with the 10^11 Ω amplifiers" | 10^11 Ω | N | N | 10^11 Ω for 90Zr–96Zr and 95Mo; 10^12 Ω for 99Ru and 101Ru | 10^11 Ω for 24Mg, 25Mg, 26Mg | N | N | N
Faraday Cup Gain Calibration Method | Method used to calibrate the relative gain (amplification factor) of each Faraday cup amplifier at the start of or during each analytical session. Common approaches: instrument internal gain calibration routine (measures signal in each cup sequentially using a common beam), amplifier rotation procedure, or measurement of a known isotope ratio solution in each cup. For instruments with SEM detectors, also describes the SEM-to-Faraday cross-calibration sequence. Specify frequency (per session, per block, weekly). | Basic | Read-Only | Text (free) | Instrument internal gain calibration routine; baseline and gain calibrated at session start and daily ¦ SEM-to-Faraday cross-calibration measured per session using known ratio solution ¦ N/A ¦ None | Source: MC-ICP-MS module | 2026-09-10 | monitored property | Relative gain differences between amplifiers cause systematic errors in isotope ratios if uncorrected. |  | N | N | N | N | N | N | "Instrument electronic baselines and amplifier gains were then measured, on peak with the line of sight valve closed"; "Although amplifier gains were measured at the start of each session the Virtual Amplifier was used in rotation mode to cancel out amplifier gains" | N | N | N | N | N | N | N
Mass Resolution Setting | Operating mass resolution of the mass analyser. Where the analyser permits selection, state the mode; where individual target species are assigned to different modes, state each. | Basic | Editable | Controlled list / Text | Unit resolution (quadrupole, fixed) ¦ Low resolution (LR) ¦ Medium resolution (MR) ¦ High resolution (HR) ¦ Multiple modes assigned per target species ¦ N/A ¦ None | Source: ICP-MS module | 2026-09-01 | (none) | MC-ICP-MS instruments allow selection of low resolution (LR; m/Δm ≈ 300–400), medium resolution (MR; m/Δm ≈ 3000–8000), or high resolution (HR; m/Δm ≈ 10,000–12,000). Most isotope ratio measurements use low resolution after chemical separation has removed interfering elements. Medium or high resolution is used when polyatomic interferences (e.g., ArN+, ArO+, ArOH+ on Fe isotopes) cannot be fully resolved by chemistry alone. |  | N | "High (entrance slit); Low (detector slit)" | Medium or high resolution — "the measurements were made on the flat-topped peak shoulder in either medium-resolution (MR) or high-resolution (HR) mode" | N | N | Low resolution | N | N | N | N | Medium resolution, M/ΔM > 6,000 | Measured "on the left 'shoulder' of the peak to resolve the difference between 40Ar1H+ and 41K+" | High-mass-resolution slit for K; low-mass-resolution slit for Cu and Zn | Medium mass resolution, R ≈ 6,600–7,000 (R = m/m0.95 − m0.05)
Collision/Reaction Cell (CRC) Configuration | Whether a collision or reaction cell is installed and its operating mode. In standard mode (STD), no cell gas is introduced and the cell acts only as an ion guide. In kinetic energy discrimination (KED) mode, helium thermalizes ions so that polyatomic interferences, which have a larger collision cross-section, are selectively retarded by a cell exit barrier voltage. In dynamic reaction cell (DRC) mode, a reactive gas selectively neutralizes specific interferences through ion-molecule reactions. Specific gas types, flow rates and cell voltages are documented in Group 4. Record 'Not installed' where the instrument has no collision/reaction cell. | Basic | Read-Only | Controlled list / Text | Not installed ¦ STD (standard mode, no gas) ¦ KED (kinetic energy discrimination, He gas) ¦ DRC (dynamic reaction cell, reactive gas) ¦ KED+DRC ¦ MS/MS (pre-cell mass filter) ¦ N/A ¦ None | Source: collision/reaction cell module | 2026-09-10 | monitored property | On MS/MS instruments a second mass filter preceding the cell enables precursor-ion selection. |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Nebulizer Type | Type and material of the pneumatic nebulizer used for sample introduction. | Basic | Read-Only | Text (free) | e.g., PFA self-aspirating MicroFlow (50–100 µL/min) ¦ PFA-ST ¦ Glass concentric ¦ ESI Apex-HF nebulizer ¦ N/A ¦ None | Source: Solution Introduction module | 2026-08-31 | (none) | Affects droplet size distribution and sample introduction efficiency. |  | Savillex C-Flow PFA nebulizer | PFA-50, Elemental Scientific, Inc. | Cyclonic glass spray chamber (wet) or ESI Apex Ω desolvating nebulizer (dry) | N | Cetac Aridus II desolvator nebulizer | N | ESI PFA-50 micro-flow nebuliser | ESI PFA-50 low uptake nebuliser | N | PFA nebulizer with an Aridus II desolvating nebulizer system | N | Elemental Scientific APEX Omega desolvating nebulizer | N | N
Spray Chamber Type and Cooling Temperature | Type of spray chamber and cooling temperature if thermostatted (°C). | Basic | Read-Only | Text (free) | e.g., Stable Introduction System (SIS): cyclonic + Scott double-pass ¦ Scott double-pass (glass) at 2°C ¦ Cyclonic at 5°C ¦ Spray chamber downstream of Aridus II ¦ N/A ¦ None | Source: Solution Introduction module | 2026-08-31 | (none) | Chamber type and temperature control aerosol droplet size and the solvent vapor load reaching the plasma. In some procedures a spray chamber is placed downstream of a desolvating nebulizer to improve signal stability. |  | N | SSI cyclonic spray dual chamber, Elemental Scientific, Inc.; cooling not stated | Cyclonic glass spray chamber for wet-plasma MR-mode work; cooling not stated | N | N | "a dual cyclonic-Scott-type quartz spray chamber"; cooling not stated | Glass Expansion micro-cyclonic "Cinnabar" spray chamber; cooling not stated | GE Cinnabar micro-cyclonic spray chamber | N | N | N | N | Quartz glass dual cyclonic spray chamber for Cu and Zn; cooling not stated | N
Desolvation System | Desolvating nebulizer or membrane desolvator used upstream of the plasma to reduce solvent load. Record 'None' if not used. | Advanced | Read-Only | Controlled list / Text | Aridus I ¦ Aridus II ¦ Apex Q ¦ Apex IR ¦ Apex HF ¦ None ¦ N/A | Source: Solution Introduction module | 2026-09-08 | acquisition pass | Lowers oxide production rates and increases sensitivity, and is commonly used for low-abundance target species or small sample sizes. Note the trade-off: desolvation can introduce additional instrumental mass bias instability relative to wet plasma introduction. |  | Cetac Aridus II | None — and deliberately: "passing solutions through a desolvating nebulizer to obtain dry plasma conditions is not viable for bulk analysis" | ESI Apex Ω for HR-mode dry plasma work, "with no auxiliary N2 flow"; none for MR-mode wet plasma | N | Cetac Aridus II — "Analyses were conducted in dry plasma mode" | None — spray chamber introduction | None — "Although greater sensitivity could be attained using a desolvating nebuliser such systems have been shown to suffer severe memory problems for Os" | None | APEX, used alongside the spray chamber as an alternative introduction system in different sessions | Aridus II | ESI Apex HF with an actively cooled membrane unit (Cr); ESI Apex Omega (Mg) | Elemental Scientific APEX Omega | Elemental Scientific APEX Ω for K ("dry plasma technique"); none for Cu and Zn | N
Sample Uptake Rate | Peristaltic pump speed or self-aspiration flow rate for sample introduction (µL/min or mL/min). | Basic | Editable | Numeric (µL/min or mL/min) | e.g., '50–100 µL/min (self-aspirating PFA nebulizer)' ¦ '200–400 µL/min (micro-flow, peristaltic)' ¦ '1 mL/min (conventional concentric)' | Source: Solution Introduction module | 2026-08-26 | (none) |  |  | ~50 µl/min | 50 µL/min | ~100 µl/min | N | N | 100 µl/min | ~80 µl/min, free aspiration | Not stated; ~6400 µl consumed over a ~16 min analysis | Peristaltic pump at 5 rpm | ~0.05 mL/min | 30 µl/min for Cr and for Mg | N | N | N
Acquisition Software | Instrument control, automation and data acquisition software used to collect the raw data, including version number. Distinct from Data Processing Software(s), which covers everything applied to the data after acquisition. | Basic | Editable | Text (free) |  | Source: Core module | 2026-08-26 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Data Processing Software(s) | All software applied to the data after acquisition in order to produce the reported quantities, including version numbers. List every package used. Distinct from Acquisition Software, which controls the instrument and collects the raw data. | Basic | Editable | Text (free) |  | Source: Core module | 2026-08-26 | (none) |  |  | N | N | N | N | Mathematica — "Data were reduced using a minimization approach implemented in Mathematica" | N | Microsoft Excel — "Following analysis all intensity data was exported and re-processed offline using Excel" | Online processing on the instrument — "Samples were processed on-line for W and Re interferences and instrumental mass bias" | N | N | N | N | N | N
4. Measurement Information | [6 empty] | 2026-08-26 |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Analytical Mode | Primary analytical mode(s) executed under this procedure. Records one value for a single-mode procedure and all applicable modes for a multi-mode one. Applies even where a procedure defines no mode flag columns. | Basic | Read-Only | Controlled list | Solution nebulisation (continuous) ¦ Flow injection | Source: Core module | 2026-08-26 | (none) | Serves as the procedure-level declaration of measurement type, distinct from the mode flag columns, which state per-field applicability. Ensures every procedure record is self-describing and comparable across the library. |  | Solution nebulisation (continuous) | Solution nebulisation (continuous) | Solution nebulisation (continuous) | Solution nebulisation (continuous) | Solution nebulisation (continuous) | Solution nebulisation (continuous) | Solution nebulisation (continuous) | Solution nebulisation (continuous) | Solution nebulisation (continuous) | Solution nebulisation (continuous) | Solution nebulisation (continuous) | Solution nebulisation (continuous) | Solution nebulisation (continuous) | Solution nebulisation (continuous)
RF Power | Radiofrequency forward power applied to the ICP, in watts. | Basic | Editable | Numeric (W) | e.g., 1200 ¦ 1250 ¦ 1300 ¦ 1350 | Source: ICP-MS module | 2026-08-31 | (none) | Controls ionization efficiency and oxide production rates. |  | N | ~1150 W | N | N | N | N | N | N | 1200 W | N | Stated qualitatively — measured "at low radiofrequency power and sample gas inflow" to reduce gas-based interferences | N | N | N
Coolant (Plasma) Gas Flow Rate | Flow rate of the outer (coolant/plasma) argon gas stream that sustains the ICP plasma, in L/min. | Basic | Editable | Numeric (L/min) | 13–16 L/min | Source: ICP-MS module | 2026-08-26 | (none) | Influences plasma temperature and oxide ion formation. |  | N | ~15 L/min Ar | N | N | N | N | N | N | 16 L/min | N | N | N | N | N
Auxiliary Gas Flow Rate | Flow rate of the intermediate (auxiliary) argon gas stream, between the torch body and the injector tube, which positions the plasma relative to the load coil, in L/min. Distinct from the outer coolant (plasma) gas and from the gas stream that carries sample into the plasma. | Basic | Editable | Numeric (L/min) | 0.8–1.1 L/min | Source: ICP-MS module | 2026-08-26 | (none) |  |  | N | ~0.8 L/min Ar | N | N | N | N | N | N | 1.01 L/min | N | N | N | N | N
Nebulizer Gas Flow Rate | Flow rate of the carrier argon gas delivered through the nebulizer (L/min). | Basic | Editable | Numeric (L/min) | 0.8–1.1 L/min | Source: Solution Introduction module | 2026-08-26 | (none) | Controls aerosol transport and strongly influences signal sensitivity and stability. |  | N | ~0.8–0.9 L/min Ar (sample gas) | N | N | N | N | N | N | 1.03 L/min (sample gas) | N | N | N | N | N
Make-up Gas and Flow Rate | Supplementary gas added to the sample-carrying stream between the sample introduction system and the plasma, with its identity and the procedure-registered target flow rate. Record any small nitrogen or hydrogen addition with its own flow, whose unit commonly differs from the make-up flow. Record 'None' explicitly where no supplementary gas is added, to distinguish it from not reported. | Advanced | Editable | Numeric (L/min) / Text | e.g., '0.25 (Ar, no desolvator fitted)' ¦ '0.85 Ar + 4 ml/min N2 (sensitivity enhancement)' ¦ 'None' | Source: ICP-MS module | 2026-08-26 | (none) | Argon make-up is standard and maintains total gas delivery where the carrier flow alone is insufficient — downstream of an ablation cell, or of a desolvation system that has removed solvent load. Small nitrogen or hydrogen additions enhance sensitivity for some elements. |  | N | N | None — the Apex Ω was run "with no auxiliary N2 flow" | N | N | N | N | N | N | N | None — "The samples were measured without the use of an auxiliary gas to the introduction system to reduce gas-based interferences" | N | N | N
Plasma Thermal Mode | Whether the ICP plasma is operated under cool plasma (<=900 W RF) or normal (hot) plasma (>1000 W RF) conditions, with the RF power that defines it. This field records the RF-POWER axis only. Solvent loading also affects plasma thermal state, but dry versus wet aerosol is recorded by Desolvation System — do not answer this field with 'dry plasma' or 'wet plasma'. Document independently of the RF Power field. | Basic | Read-Only | Controlled list / Text | Normal plasma (>1000 W RF) ¦ Cool plasma (≤900 W RF) ¦ Mixed: specify mode per analytical sub-run ¦ N/A ¦ None | Source: ICP-MS module | 2026-09-08 | acquisition pass | Normal plasma (>1000 W RF) is standard for solution MC-ICP-MS. Cool plasma is rarely used in MC-ICP-MS due to stability requirements for high-precision isotope ratio work. |  | N | Wet plasma — solutions "introduced as a 'wet' aerosol (in 2% HNO3) into the ICP torch via a cyclonic spray dual chamber"; dry plasma deliberately rejected as "not viable for bulk analysis" | Both, by mode — "either a cyclonic glass spray chamber (wet plasma, MR-mode, Pt cones) or an ESI Apex Ω desolvating nebulizer system (dry plasma, HR-mode, Ni cones)" | N | Dry plasma — "Analyses were conducted in dry plasma mode using a Cetac Aridus II desolvator nebulizer" | N | N | N | N | N | N | N | Dry plasma for K — "all K isotope analyses were undertaken using a 'dry plasma' technique with the Elemental Scientific APEX Ω high-sensitivity desolvation system"; wet plasma for Cu and Zn via a quartz glass dual cyclonic spray chamber | N
ICP Tuning | Description of the approach used to optimise ICP plasma conditions prior to analysis, including the reference material used for tuning and the acceptance criteria (e.g., oxide production threshold, sensitivity targets, mass calibration). | Advanced | Editable | Text (free) | e.g., 'Daily tune on the bracketing standard at run concentration; optimised for maximum intensity and peak flatness with UO+/U+ <0.5%' ¦ 'Tuned on 100 ng/ml Nd; oxide production monitored on CeO+/Ce+' | Source: ICP-MS module | 2026-08-26 | (none) |  |  | N | N | N | N | N | N | "At the start of each analytical session the Neptune was tuned for maximum sensitivity and optimal peak shape using an Os solution, either the UMd or DTM RMs, and the mass calibration was updated by peak-centering on the centre-cup mass 187Os" | N | N | "Tuning was performed to minimize interferences of 40Ar2 16O+ and 40Ar2 14N+ on 96Zr+ and 94Zr+ and with it the on-peak background corrections" | Measured "at low radiofrequency power and sample gas inflow" deliberately, to reduce gas-based interferences | N | N | N
Instrument Warm-up / Session Duration Limit | Minimum warm-up time required after plasma ignition before analyses begin, and any maximum session duration enforced to maintain stable operating conditions. | Advanced | Read-Only | Text (free) | e.g., '>=1 h warm-up after plasma ignition; sessions <=12 h' ¦ '30 min warm-up; no session duration limit' | Source: ICP-MS module | 2026-08-26 | (none) |  |  | N | N | N | N | N | N | "Instrument electronic baselines and amplifier gains were then measured ... while the Neptune was allowed to warm up for half an hour" | N | N | N | N | N | N | N
Instrument Sensitivity | Instrument sensitivity achieved in the session, with the isotope or monitored property it was measured on and the conditions it applies to. May be expressed either as detected signal per unit concentration or per unit mass of target species delivered — counts per second per ppb, volts per ppm, counts per picogram — or as useful yield, the percentage of sampled atoms ultimately detected as ions, with the method used to derive it cited. A sensitivity the procedure requires before analyses may begin belongs with the tuning acceptance criteria. | N/A | Advanced | Numeric + unit / Text | e.g., '572 V/ppm total Zr' ¦ '~2.5 x 10^6 cps/ppb on 115In' ¦ '0.04 count pg-1 ml (111Cd)' ¦ '0.1% useful yield' ¦ 'N/A' | Source: ICP-MS module | 2026-09-10 | monitored property | Useful yield is the more comparable of the two wherever the amount of material consumed varies between procedures, as it does with spot size, fluence and repetition rate. |  | ~1.3 x 10^-10 A total ion beam for a ~100 ppb Mo solution at ~50 µl/min uptake | N | Typical 56Fe+ intensities 1.4 nA (wet plasma, MR) to 2 nA (dry plasma, HR) | 3.5–10 V for the most abundant isotope at 15–25 ppb (10 V for 140Ce, 4 V for 142Nd, 3.5 V for 152Sm) | "The average instrumental sensitivity using this setup was around 572 V/ppm of total Zr" | ~1–1.5 V for 85Rb at 15–25 ppb in low resolution | "The sensitivity for Os with this set-up was approximately 50 V for a 1 µg ml-1 Os solution at a free-aspiration rate of around 80 µl min-1" | N | N | Total Zr ion beams 3.5–13 V at 30 ppb | Typical Mg signal intensities 80–90 V | N | N | N
Target Species | The chemical species this procedure is designed to determine, recorded at whatever resolution the chemistry is resolved — element(s) for this technique; valence species where a procedure resolves oxidation state; compounds where it resolves molecules. Isotopes are not target species: isotopes of an element are the same chemical species. The subset actually measured may be narrower than the registered suite where species were excluded due to interferences or scope reduction. | Basic | Editable | Text (free) |  | Source: Target Species module | 2026-09-01 | defines: target species |  |  | Mo (and Ba by TIMS) | S | Fe | The rare earth elements — La, Ce, Nd, Sm, Eu, Gd, Dy, Er, Yb and Y | Zr | Rb | Os | Os | Rb | Zr | Fe, Cr and Mg | K | K, Cu and Zn | Ti
Monitored Masses | Specific masses monitored in this procedure, grouped by the target species element they serve where they serve one. Covers atomic isotopes and, where a reaction cell shifts an target species onto a different mass, the product mass actually measured. Includes interference-monitor and internal-standard masses, which serve no target species and so have no parent element. The target species list is given by the Target Species field and is never inferred from the element symbols appearing here. | Basic | Read-Only | Text (free) | e.g., ⁷Li ¦ ²⁵Mg ¦ ⁵⁶Fe ¦ ²⁰⁸Pb ¦ N/A ¦ None |  | 2026-09-10 | defines: monitored property per target species |  |  | ⁹²Mo, ⁹⁴Mo, ⁹⁵Mo, ⁹⁶Mo, ⁹⁷Mo, ⁹⁸Mo, ¹⁰⁰Mo (Mo); ⁹¹Zr, ⁹⁹Ru (interference monitors, no target species) — "Isobaric interferences of Zr and Ru on Mo masses were corrected by monitoring 91Zr and 99Ru" (p.2); mass bias normalised to 98Mo/96Mo and ε⁹²/⁹⁴/⁹⁵/⁹⁷/¹⁰⁰Mo reported (pp.2–3) | ³²S (L3), ³³S (C), ³⁴S (H3) — Table 1 "Cup configuration", p.3. All three serve the single target species S | ⁵⁴Fe, ⁵⁶Fe, ⁵⁷Fe, ⁵⁸Fe (Fe); ⁵³Cr, ⁶⁰Ni (interference monitors, no target species) — "Ion beams of 54Fe+, 56Fe+, 57Fe+, and 58Fe+ were analyzed in static mode on Faraday collectors... Possible isobaric interferences from 54Cr+ and 58Ni+ were measured simultaneously by monitoring 53Cr+ and 60Ni+" (p.6) | N — the paper states "The cup configurations used for isotopic analyses of the REEs are provided in table S2" (p.9); that supplementary table is not in the archived PDF | Masses 90, 91, 92, 93, 94, 95, 96 and 98 — "Masses 90, 91, 92, 93, 94, 95, 96, and 98 were measured in static mode at 0.5 amu spacing in the Nu Plasma II collector block, allowing direct monitoring of all Zr isotopes and Mo interferences (masses 95 and 98)" (p.11). 95 and 98 carry the Mo monitors and serve no target species | ⁸⁵Rb, ⁸⁷Rb (Rb); ⁸⁸Sr (interference monitor, no target species) — "Rubidium-85 and -87 were measured on L2 and axial (A) Faraday collectors, respectively", and the ⁸⁷Sr contribution "was corrected for by monitoring 88Sr" (p.8) | ¹⁸⁴Os, ¹⁸⁶Os, ¹⁸⁷Os, ¹⁸⁸Os, ¹⁸⁹Os, ¹⁹⁰Os, ¹⁹²Os (Os); ¹⁸⁵Re and ¹⁸²W/¹⁸⁴W/¹⁸⁶W (interference monitors, no target species) — the Os masses are those whose ratios to ¹⁸⁸Os the paper measures and reports; the Re and W monitors are named throughout the interference-correction discussion (§3.6, pp.12–18), and the L3 detector is named as carrying ¹⁸⁴Os (p.26). The paper gives no single cup-configuration table | N — the paper describes the Nu Plasma acquisition only as two-sequence static multi-collection and gives no mass list or cup configuration for that instrument | ⁸⁴Sr (L2), ⁸⁵Rb (L1), ⁸⁶Sr (C), ⁸⁷Rb + ⁸⁷Sr (H1), ⁸⁸Sr (H2) — Table 2, p.3. ⁸⁵Rb and ⁸⁷Rb serve the target species Rb; the Sr masses are interference monitors with no target species, ⁸⁸Sr being the one used to correct ⁸⁷Sr on ⁸⁷Rb | ⁹⁰Zr, ⁹¹Zr, ⁹²Zr, ⁹⁴Zr, ⁹⁶Zr (Zr); ⁹⁵Mo, ⁹⁹Ru, ¹⁰¹Ru (interference monitors, no target species) — "Faraday cups with 10¹¹ Ω amplifiers were used to collect Zr masses 90Zr to 96Zr and 95Mo, whereas 10¹² Ω amplifiers were applied for the collection of 99Ru and 101Ru" (p.6) | ²⁴Mg, ²⁵Mg, ²⁶Mg (Mg) — "The isotopes 24Mg, 25Mg and 26Mg were analysed using 10¹¹ Ω resistors" (p.8) | ³⁹K, ⁴¹K (K) — δ⁴¹K is defined from the ⁴¹K/³⁹K ratio (p.4). ⁴⁰Ar¹H⁺ is named as the interference on ⁴¹K⁺ (p.4) but is not itself a monitored mass | ³⁹K, ⁴¹K (K); ⁶³Cu, ⁶⁵Cu (Cu); ⁶⁴Zn, ⁶⁶Zn (Zn) — the three delta values are defined from the ⁴¹K/³⁹K, ⁶⁵Cu/⁶³Cu and ⁶⁶Zn/⁶⁴Zn ratios (p.7) | ⁴⁶Ti, ⁴⁷Ti, ⁴⁸Ti, ⁴⁹Ti, ⁵⁰Ti (Ti); ⁴⁴Ca, ⁵¹V, ⁵²Cr, ⁵³Cr (interference monitors, no target species) — "Titanium isotopes were collected in two cup configurations. First, all five Ti isotopes and 44Ca were measured enabling correction of the Ca interference on 46Ti and 48Ti. The second configuration included 49Ti, 50Ti, 51V, 52Cr and 53Cr to correct for isobaric interferences from V and Cr on 50Ti" (p.8). ⁴⁵Sc belongs to the LLNL procedure in the same paper, not this one
Reported Variables and Units | The final variable(s) this procedure reports and their units — distinct from the fields recording what was *acquired* rather than what is reported. Record every reported variable, including intermediate quantities reported alongside final ones (e.g. both the 206Pb/238U ratio and the 206Pb/238U date). Where a reported variable is a nominal property with no magnitude (e.g. a mineral identification), record the variable and give the unit as 'N/A — nominal property'. | Basic | Read-Only | Text (free) | e.g., 'Element concentration (ppm); oxide (wt%)' ¦ 'd56Fe (permil vs IRMM-014); d57Fe (permil vs IRMM-014)' ¦ '206Pb/238U date (Ma); 207Pb/206Pb date (Ma)' ¦ 'Mineral identity (N/A — nominal property); match score (0-1)' | Source: Core module | 2026-08-26 | defines: reported property | A procedure may acquire many channels and report a small number of derived quantities; without this field a data consumer cannot tell which. |  | εiMo relative to the Alfa Aesar solution standard, εiMo = [(iMo/96Mo)sample/(iMo/96Mo)standard − 1] x 10^4 | δ34S and δ33S in permil vs V-CDT | µ-notation Fe isotope ratios relative to IRMM-524a | Mass-dependent REE isotopic fractionation relative to the OL-REE standards, in delta notation | δ9x/90ZrNIST in permil — δ91/90Zr, δ92/90Zr, δ94/90Zr and δ96/90Zr | δ87Rb in permil relative to NIST SRM984 | 187Os/188Os, 186Os/188Os and 184Os/188Os ratios | 187Os/188Os, 186Os/188Os and 184Os/188Os ratios | δ87Rb in permil = [(87Rb/85Rb)sample/(87Rb/85Rb)standard − 1] x 1000 | ε91Zr, ε92Zr and ε96Zr relative to NIST SRM 3169 | µ-notation Fe relative to IRMM-014, Cr relative to SRM979, Mg relative to DTS-2b | δ41K in permil relative to NIST SRM 3141a | δ41K, δ65Cu and δ66Zn in permil, each defined explicitly against its bracketing standard | ε46Ti, ε48Ti and ε50Ti (parts per 10^4) relative to an in-house Alfa Aesar Ti wire standard, εiTi = [(iTi/47Ti)sample/(iTi/47Ti)standard − 1] × 10^4, "where i refers to the isotope masses 46Ti, 48Ti and 50Ti" (p.8)
Mass Resolution Assignment | Mass resolution mode used for acquisition. One target species may be acquired at more than one resolution, so the assignment is per acquired mass rather than per element. The overall mode(s) used in the procedure are recorded in Mass Resolution Setting (Group 3). | Basic | Read-Only | Text (free) | e.g., HR: K; LR: Cu, Zn ¦ All LR ¦ All MR ¦ N/A ¦ None |  | 2026-09-08 | acquisition pass | The selected resolution determines which polyatomic interferences are physically resolved by the magnetic sector.
Collector Configuration | Assignment of isotope masses to Faraday cup or ion counter detector positions for this procedure, and the target species each position serves. Lists the mass number monitored in each collector position with its target species element where it has one — e.g. 'L3=116Sn (Sn); L2=117Sn (Sn); Ax=120Sn (Sn); H1=121Sb (interference monitor, no target species)'. Interference-monitor and internal-standard positions serve no target species and carry no parent element; the target species list itself is given by the Target Species field and is never inferred from the mass labels here. For static multi-collection procedures, one configuration applies throughout the measurement. For multi-dynamic procedures, list all configurations and the cycling sequence. | Basic | Read-Only | Text (free) |  | Source: MC-ICP-MS module | 2026-09-10 | monitored property | Different element systems require different mass assignments within the collector array span. |  | N | 32S(L3), 33S(C), 34S(H3) | 54Fe, 56Fe, 57Fe, 58Fe in static mode, with 53Cr and 60Ni monitored simultaneously | Static mode for most REEs; a subconfiguration for Dy and Yb to monitor isobaric interferences | Masses 90, 91, 92, 93, 94, 95, 96 and 98 "measured in static mode at 0.5 amu spacing in the Nu Plasma II collector block" | 85Rb, 87Rb+87Sr and 88Sr on three collectors, 88Sr on H1 | L4=182W, L3=184Os, L2=185Re, L1=186Os, Ax=187Os, H1=188Os, H2=189Os, H3=190Os, H4=192Os, with 184W, 186W and 187Re as interference monitors | Two-sequence static multi-collection | L2=84Sr, L1=85Rb, C=86Sr, H1=87Rb+87Sr, H2=88Sr | 90Zr–96Zr and 95Mo on 10^11 Ω cups; 99Ru and 101Ru on 10^12 Ω cups | 49Ti, 51V, 56Fe alongside 50Cr, 52Cr, 53Cr, 54Cr; 24Mg, 25Mg, 26Mg | N | N | Two cup configurations: (1) ⁴⁶Ti–⁵⁰Ti and ⁴⁴Ca; (2) ⁴⁹Ti, ⁵⁰Ti, ⁵¹V, ⁵²Cr, ⁵³Cr — "Titanium isotopes were collected in two cup configurations" (p.8). Cup positions are not stated
Number of Blocks per Measurement | Number of measurement blocks acquired per sample or standard solution introduction. In MC-ICP-MS, data are structured as blocks of cycles; each block typically begins with a baseline (on-peak zero) measurement before the analytical cycles. Adjust for samples with low signal or where higher precision is required. | Basic | Editable | Integer | e.g., 1 ¦ 2 ¦ 3 ¦ 5 | Source: MC-ICP-MS module | 2026-08-31 | (none) | Multiple blocks per sample allow inspection of signal stability and within-measurement drift. |  | N | N | N | N | N | A single block | 9 blocks | 1 block | Blocks of 20 cycles | N | N | N | N | N
Number of Cycles per Block | Number of measurement cycles acquired per block. One cycle corresponds to a single set of simultaneous Faraday cup readings integrated for the duration specified in Integration Time per Cycle. Total integration time per sample = (Number of Cycles per Block) × (Integration Time per Cycle) × (Number of Blocks per Measurement). | Basic | Editable | Integer | e.g., 25 ¦ 40 ¦ 50 ¦ 80 | Source: MC-ICP-MS module | 2026-09-08 | acquisition pass |  |  | 100 isotope ratio measurements, preceded by 40 baseline integrations | 20 cycles | 25 (HR) or 50 (MR) cycles | 40 cycles in the main configuration; the subconfiguration measured twice | 50 cycles | 25 cycles | 5 cycles per block | 50 cycles | 20 cycles | 60 ratios | Fe 200 cycles; Cr 100 cycles; Mg 100 cycles | "Each sample was measured approximately 20 times" | N | 40 cycles
Integration Time per Cycle | Duration of signal integration per measurement cycle (seconds). | Basic | Editable | Numeric (s) | e.g., 4.194 s ¦ 8 s ¦ 10 s ¦ 16.8 s | Source: MC-ICP-MS module | 2026-09-10 | monitored property | Determines counting statistics per cycle. Longer integration times improve shot-noise precision but increase the impact of signal drift within the integration window. For high-gain (10¹² or 10¹³ Ω) amplifier channels, longer integration times are often required to accumulate sufficient charge. |  | 8.4 s | 8.5 s | 8.369 s | 4.142 s in the subconfiguration | 5 s | 4.194 s | 4 s | 8 s for sequence 1 and 4 s for sequence 2 | 8.389 s | 4.2 s | Fe 8.3 s; Cr 8.3 s; Mg 16.7 s | N | N | 8.39 s for the first cup configuration and 4.19 s for the second — "A sample measurement consisted of 40 cycles with 8.39 s integration time for the first configuration and 4.19 s for the second" (p.8). The 4 s in the same paper is the LLNL procedure's (p.8), not this one
Baseline Measurement Approach | Method and timing for measuring the ion beam baseline (on-peak zero) during acquisition, and the portion of the acquisition it is taken from. Three approaches are in use: deflecting the ion beam (beam-off or electrostatic deflector); aspirating an acid blank; or collecting a defined number of laser-off cycles at the start of the same block. State which, and how many cycles or how long. | Basic | Read-Only | Text (free) | e.g., Beam deflector on-peak zero; 1 baseline block of 1 cycle (4.194 s) per measurement block ¦ On-peak zero measured before each sample by aspirating 0.3 M HNO3 for 10 cycles ¦ N/A ¦ None | Source: MC-ICP-MS module | 2026-08-31 | (none) | The choice depends on how sample reaches the plasma — an acid blank applies only to solution introduction, and for laser ablation no blank can be aspirated mid-ablation. |  | 40 on-peak-zero baseline integrations of 8.4 s preceding each measurement | N | "On peak zero intensities from a blank solution measured at the beginning of each sequence were subtracted from all individual measurements" | N | "On-peak-zero correction was done using the mean acid blank intensities before data processing to account for blank contribution as well as any 'memory' effects from the Aridus II sample introduction system during the run" | N | Electronic baselines "measured, on peak with the line of sight valve closed"; peak centering and baselines "were not carried out at the start of each analysis to reduce measurement time and conserve sample but were repeated several times during an analytical session" | N | N | Both: "Electronic baselines were measured for 30 s prior to each analysis. An on-peak background correction was performed." | On-peak baseline — Fe 25 x 16.7 s; Cr 75 s; Mg 25 x 16.7 s | N | N | N
Wash Time Between Samples | Duration of instrument rinse between successive sample solutions (seconds). | Basic | Read-Only | Numeric (s) | e.g., 120 ¦ 180 ¦ 210 ¦ 300 | Source: Solution Introduction module | 2026-08-31 | (none) | Controls carryover and memory effects. For isotope-ratio work, complete washout is critical to prevent cross-contamination between samples with differing isotopic compositions. |  | N | 2 min for solution work (4 min for laser) | 210 s | 300 s rinsing between bracketed measurements | N | 60 s wash in 0.45 M HNO3, with a 90 s take-up time | "Teflon-distilled (TD) 3 or 5 mol/l HCl acid was aspirated between analyses until the 192Os beam decreased to acceptable background levels"; not required in single-RM sessions | TD 3 mol/l HCl aspirated between analyses until the Os beam decreased to acceptable background levels | N | N | N | N | N | N
Analysis Sequence | The repeating order in which calibration or bracketing standards, quality-control and secondary reference materials, blanks and unknowns are interleaved within a measurement session. Adjustments must maintain the bracketing strategy defined in the procedure. | Basic | Editable | Text (free) | e.g., SSB: sample bracketed by two standards (Std1–Sample–Std2) ¦ Double-spike: spiked standards bracketing spiked samples ¦ SSB + secondary RM every 5 samples ¦ N/A ¦ None ¦ For SSB procedures, each sample is bracketed by two standards (standard–sample–standard); the bracketing standard is the same isotopic material as the zero-delta reference. | Source: ICP-MS module | 2026-08-31 | (none) |  |  | Bracketing runs of the Alfa Aesar solution standard; BHVO-2 digestions "analyzed together with each set of samples" | N | "Sample analyses were bracketed by measurements of the reference material IRMM-524a" | Standard-sample bracketing — "On average, LREEs were measured nine times bracketed by OL-REE isotope standard spaced apart by 300-s rinsing time" | "Each sample measurement was individually bracketed by measurements of the ZrNIST solution spiked at the same level as our samples and matched in concentration (60 ng/g) as well as acid matrix"; each measurement preceded by an acid blank | Standard-sample bracketing | N | N | Standard-sample bracketing; an external pure Rb ICP-MS solution "analyzed as an external standard during each analytical session to monitor the reproducibility" | Standard sample bracketing against NIST SRM 3169; "The Zr standard material NIST SRM 3169 was analyzed in each session" | Standard-sample bracketing, "ten individual standard-bracketed sample analyses" per reported value | Standard-sample bracketing against NIST SRM 3141a; BHVO-2 measured alongside the samples | Standard-sample bracketing for all analyses; BHVO-2 "analysed alongside all sample analyses" | N
Collision Gas Type | Type of collision gas introduced into the collision/reaction cell for PHYSICAL energy discrimination (KED), where an inert gas thermalizes ions and polyatomic interferences are retarded by their larger collision cross-section. Typically He. Record a gas here when it is used without intended chemistry; where the same gas is also used to drive a mass shift, record that use under Reaction Gas Type — O2 is attested in both roles. Record 'None' if the cell is in STD mode, and 'N/A' where Collision/Reaction Cell (CRC) Configuration does not include KED or the instrument has no cell. | Basic | Read-Only | Controlled list / Text | He ¦ H2 ¦ He+H2 ¦ O2 ¦ N/A ¦ None | Source: collision/reaction cell module | 2026-09-10 | monitored property | Helium (He) is the standard collision gas for kinetic energy discrimination, being low in mass and chemically inert. |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Collision Gas Flow Rate | Flow rate of the collision gas, typically He, introduced into the collision/reaction cell in KED mode, in mL/min. Record 'None' if the cell is in STD mode, and 'N/A' where Collision/Reaction Cell (CRC) Configuration does not include KED or the instrument has no cell. | Basic | Editable | Numeric (mL/min) | e.g., 0.5 ¦ 4.5 ¦ 5.2 ¦ N/A | Source: collision/reaction cell module | 2026-09-01 | (none) | Controls the degree of ion thermalization and KED efficiency; higher flow rates give greater interference suppression at the cost of target species sensitivity. |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Cell Exit Discrimination Voltage | Bias voltage applied at the collision/reaction cell exit to discriminate between target species ions and low-energy polyatomic interferences in KED mode, in volts (V). Record 'None' if the cell is in STD mode, and 'N/A' where Collision/Reaction Cell (CRC) Configuration does not include KED or the instrument has no cell. | Basic | Editable | Numeric (V) | e.g., −2 ¦ −5 ¦ N/A | Source: collision/reaction cell module | 2026-09-01 | (none) | A negative bias preferentially retards slow polyatomic ions while transmitting faster target species ions, controlling the degree of polyatomic suppression. |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Reaction Gas Type | Type of reactive gas introduced into the cell for interference removal through ION-MOLECULE CHEMISTRY, either on-mass or by mass shift. Common reaction gases include NH₃ (e.g., for Fe, Ca, K isotopes), O₂ (e.g., for As, Ge, Te mass shift) and CH₄. The distinction from Collision Gas Type is the mechanism, not the gas: record a gas here when a reaction is intended. Record 'None' if no reactive gas is used, and 'N/A' where Collision/Reaction Cell (CRC) Configuration does not include a reaction mode or the instrument has no cell. | Advanced | Read-Only | Controlled list / Text | NH3 ¦ O2 ¦ CH4 ¦ N/A ¦ None | Source: collision/reaction cell module | 2026-09-10 | monitored property |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Reaction Gas Flow Rate | Flow rate of the reactive gas introduced into the dynamic reaction cell (DRC), in mL/min. Record 'None' if DRC mode is not used, and 'N/A' where Collision/Reaction Cell (CRC) Configuration does not include DRC or the instrument has no cell. | Advanced | Editable | Numeric (mL/min) | N/A | Source: collision/reaction cell module | 2026-08-27 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Internal Standard Element | Element(s) used as the internal standard, the monitored isotope, and how the internal-standard concentration was established for unknowns. Where the element serves isotope-ratio work it must have a known isotopic composition and a mass close to the target species's. Record 'None' together with the alternative used where no internal standard is applied. | Basic | Read-Only | Text (free) | e.g., Cu (65Cu/63Cu for Fe procedures) ¦ Ni (62Ni/58Ni for Fe procedures) ¦ Tl (205Tl/203Tl for Pb procedures) ¦ Cd (113Cd/111Cd for Sb procedures) ¦ None ¦ N/A | Source: ICP-MS module | 2026-09-08 | acquisition pass |  |  | N/A — mass bias corrected by standard-sample bracketing, internal normalization or a double spike rather than by an added internal standard element | N/A — mass bias corrected by standard-sample bracketing, internal normalization or a double spike rather than by an added internal standard element | N/A — mass bias corrected by standard-sample bracketing, internal normalization or a double spike rather than by an added internal standard element | N/A — mass bias corrected by standard-sample bracketing, internal normalization or a double spike rather than by an added internal standard element | N/A — mass bias corrected by standard-sample bracketing, internal normalization or a double spike rather than by an added internal standard element | N/A — mass bias corrected by standard-sample bracketing, internal normalization or a double spike rather than by an added internal standard element | N/A — mass bias corrected by standard-sample bracketing, internal normalization or a double spike rather than by an added internal standard element | N/A — mass bias corrected by standard-sample bracketing, internal normalization or a double spike rather than by an added internal standard element | N/A — mass bias corrected by standard-sample bracketing, internal normalization or a double spike rather than by an added internal standard element | N/A — mass bias corrected by standard-sample bracketing, internal normalization or a double spike rather than by an added internal standard element | N/A — mass bias corrected by standard-sample bracketing, internal normalization or a double spike rather than by an added internal standard element | N/A — mass bias corrected by standard-sample bracketing, internal normalization or a double spike rather than by an added internal standard element | N/A — mass bias corrected by standard-sample bracketing, internal normalization or a double spike rather than by an added internal standard element | N/A — mass bias corrected by standard-sample bracketing, internal normalization or a double spike rather than by an added internal standard element
Internal Standard Concentration | Target concentration of internal standard element(s) in all measured solutions (µg/L or ppb). Record 'N/A' where Internal Standard Element is 'None'. | Advanced | Read-Only | Numeric (µg/L) | e.g., '5 µg/L In' ¦ '10 µg/L Rh' ¦ '1–10 µg/L (In, Rh, Re multi-element)' ¦ '20 µg/L (high-dilution runs)' | Source: Solution Introduction module | 2026-08-26 | (none) |  |  | N/A — no added internal standard element | N/A — no added internal standard element | N/A — no added internal standard element | N/A — no added internal standard element | N/A — no added internal standard element | N/A — no added internal standard element | N/A — no added internal standard element | N/A — no added internal standard element | N/A — no added internal standard element | N/A — no added internal standard element | N/A — no added internal standard element | N/A — no added internal standard element | N/A — no added internal standard element | N/A — no added internal standard element
Oxide Production Method and Threshold | Method used to quantify plasma oxide production and the acceptance threshold applied before commencing analysis. Record both the monitored mass ratio(s) and the maximum allowed threshold(s). Measured values are recorded in Oxide Production. The monitored ratio is procedure-dependent: ThO+/Th+ (m/z 248/232), UO+/U+ (254/238), CeO+/Ce+ (156/140) and BaO+/Ba+ are all in use. | Basic | Read-Only | Text (free) | e.g., CeO+/Ce+ (m/z 156/140) < 0.5% ¦ CeO+/Ce+ < 0.3%; BaO+/Ba+ (m/z 154/138) < 0.1% ¦ N/A ¦ None | Source: ICP-MS module | 2026-09-01 | (none) | Stricter oxide thresholds may be required for target species sensitive to oxide interferences (e.g., BaO+ on REE isotopes). |  | N | N | N | N | N | N | N | N | N | Argide and Ar-Ar-oxide interferences on 94Zr and 96Zr minimised by tuning; no numeric threshold stated | N | N | N | N
Oxide Production | Measured oxide production ratio obtained during session-start tuning, for the proxy specified in Oxide Production Method and Threshold. Record the measured value and confirm whether the procedure threshold was met. | N/A | Basic | Text (free) | CeO+/Ce+ = 0.3% (threshold < 0.5% met) ¦ BaO+/Ba+ = 0.07% (threshold < 0.1% met) | Source: ICP-MS module | 2026-08-26 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Peak Flatness Method and Threshold | Method used to verify that every collector sits on a flat region of its peak top simultaneously, and the acceptance threshold applied before analysis begins. Typically assessed by scanning across the peak plateau, or by comparing ratios measured at slightly offset mass positions. | Basic | Read-Only | Text (free) |  | Source: MC-ICP-MS module | 2026-08-26 | (none) | Specific to simultaneous collection: a sequential instrument centres one mass at a time, whereas a multi-collector array must have all cups on flat peak tops at once or the measured ratio carries a peak-shape bias that no downstream correction removes. |  | N | N | Measurements made "on the flat-topped peak shoulder" in MR or HR mode; no numeric threshold stated | N | N | N | Tuned for "optimal peak shape"; mass calibration updated by peak-centering on the centre-cup mass 187Os. No numeric threshold stated | N | N | N | N | Measurements taken "on the left 'shoulder' of the peak"; no numeric threshold stated | N | N
Peak Flatness | Peak flatness actually achieved in the session, assessed against the criterion registered by the procedure. | N/A | Basic | Text (free) |  | Source: MC-ICP-MS module | 2026-08-26 | (none) | Companion measured value to Peak Flatness Method and Threshold, following the criterion-versus-measurement split used for Oxide Production. |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Ion Counter Dead Time | Dead time of the ion-counting detector(s), used in the dead-time correction applied to high count rates. Distinct from pulse/analog cross-calibration, which relates the two detector modes rather than correcting counting losses within the pulse-counting mode. | Basic | Editable | Numeric (ns) | e.g., '23 ns (SEM behind L4)' ¦ '20 ns each on the four ion counters' ¦ 'N/A (Faraday cups only)' |  | 2026-09-10 | monitored property |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Collision/Reaction Gas Mixture Ratio | Where the collision or reaction cell is supplied with a mixture of gases rather than a single gas, the identities and proportions of that mixture. Recorded separately from the gas identity. Record 'N/A' where a single gas is used. | Advanced | Editable | Text (free) | e.g., 'He:H2 92% : 8%' ¦ 'N/A (single gas)' ¦ 'None' | Source: collision/reaction cell module | 2026-09-10 | monitored property | The proportions change reaction efficiency and interference suppression independently of which gases are used.
Reaction Product Ion / Mass-Shift Transition | Where a monitored mass is produced by a reaction in the collision/reaction cell, the precursor ion, the reagent gas and the product ion measured. Records the mass-shift chemistry relating the mass measured to the target species it reports, which the monitored mass alone does not state. Record 'N/A' where the target species is measured on its own mass. | Advanced | Read-Only | Text (free) | e.g., '125Te + 16O -> 141TeO+' ¦ '77Se + 16O -> 93SeO+' ¦ 'N/A (on-mass measurement)' | Source: collision/reaction cell module | 2026-09-10 | monitored property
Acquisition Pass | The acquisition passes the procedure is divided into, named and described so that fields keyed by acquisition pass can point at them. A pass is a sub-procedure: a distinct traversal of the measurement with its own configuration, run in sequence on the same material. State what distinguishes each pass — resolution mode, cup or cell configuration, plasma or introduction path, spot size — since that differs by technique. Identical repeats of one configuration are replicates, not passes, and belong in Number of Replicates. | Basic | Read-Only | Text (free) | e.g., 'Pass 1: LR (M/dm 300); Pass 2: MR (M/dm 4000)' ¦ 'Pass 1: KED (He), on-mass; Pass 2: MS/MS O2 mass-shift' ¦ 'Pre-ablation; Run 1 (major, MR, 30 um); Run 2 (trace, LR, 130 um)' ¦ 'Single pass' | Source: ICP-MS module | 2026-09-08 | defines: acquisition pass
Number of Acquisition Passes | Number of acquisition passes the procedure runs. A count of the passes enumerated in Acquisition Pass, recorded separately so multi-pass procedures are findable without parsing that field. | Basic | Read-Only | Integer | e.g., 1 ¦ 2 ¦ 3 | Source: ICP-MS module | 2026-09-08 | (none)
Doubly-Charged Species Monitor | The mass ratio monitored to estimate doubly-charged ion (M²⁺) formation during instrument tuning. The monitor species and the mass positions monitored should be stated explicitly. Analogous to Oxide Production Method and Threshold for oxide monitoring. | Advanced | Editable | Text (free) | e.g., Ba²⁺/Ba⁺ (m/z 69/138) ¦ Ce²⁺/Ce⁺ (m/z 70/140) ¦ ²³⁸U²⁺/²³⁸U⁺ (m/z 119/238) ¦ N/A ¦ None | Source: ICP-MS module | 2026-09-11 | (none) | Doubly-charged ions appear at half the mass of the parent ion and can cause isobaric interferences on target species in that mass region. |  | N | N (⁶⁴Zn²⁺, ⁶⁶Zn²⁺ and ⁶⁸Zn²⁺ are listed as interferences on the S masses, Table 2, p.4, and resolved by mass resolution, p.5; no M²⁺ tuning monitor or production ratio is stated) | N | N | N | N | N | N | N ("double-charged Er or Yb isotopes" are named as interferences that the chemistry separates from Rb, p.2; no M²⁺ tuning monitor or production ratio is stated) | N | N | N | N | N
Doubly-Charged Species Production | Measured percentage of doubly-charged ion production for the monitored species at the time of instrument tuning. The acceptable threshold is typically <1% or <3%. Record both the threshold and the measured value. | Advanced | Editable | Text (free) | e.g., Ce²⁺/Ce⁺ < 3% (threshold); 1.8% (measured) ¦ N/A ¦ None | Source: ICP-MS module | 2026-09-11 | (none) | Elevated doubly-charged production indicates incomplete ionization and potential interference on elements at approximately half the mass of abundant matrix components. |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
5. Data Processing | [6 empty] | 2026-08-26 |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Mass Bias Correction Strategy | Primary strategy used to correct for instrumental isotopic mass fractionation. Four main strategies: (1) Sample-standard bracketing (SSB): alternating sample and isotopic standard measurements; bias interpolated linearly between bracketing measurements. (2) Double-spike: a mixture of two enriched isotopes of the target species element added before digestion provides an internal monitor of both instrumental and chemical mass fractionation. (3) Internal normalization: an element of known isotopic composition added to samples and standards; measured ratio of normalizing element used to calculate the mass bias factor applied to target species ratios. (4) SSB + internal normalization: both strategies applied simultaneously for redundant bias correction. | Basic | Read-Only | Text (free) | e.g., SSB ¦ Double-spike ¦ Internal normalization ¦ SSB + internal normalization ¦ N/A ¦ None |  | 2026-09-01 | (none) | In solution MC-ICP-MS, instrumental mass bias is typically 3–5%/amu and must be corrected before isotope ratios can be interpreted geochemically. |  | Internal normalization to 98Mo/96Mo = 1.453173 using the exponential law, plus bracketing against the Alfa Aesar standard | Standard-sample bracketing against matrix-matched purified S solutions | Internal normalization to 57Fe/56Fe = 0.023095 or 57Fe/54Fe = 0.362549 using the exponential law, with IRMM-524a bracketing | Standard-sample bracketing against OL-REE standards — "SSB is advantageous over the double-spike approach because one can distinguish mass-dependent fractionation from isotopic anomalies" | 91Zr-96Zr double spike inversion, with ZrNIST bracketing after inversion | Standard-sample bracketing against NIST SRM984 | Instrumental mass bias correction applied offline in Excel alongside abundance sensitivity and W/Re interference corrections | "Samples were processed on-line for W and Re interferences and instrumental mass bias" | "Measurements were made using standard-sample bracketing to correct for instrumental mass bias" | Internal normalization to 94Zr/90Zr = 0.3381 using the exponential law; an initial Mo correction uses a mass bias relative to 91Zr/90Zr = 0.21798; results reported by standard sample bracketing to NIST SRM 3169 | Standard-sample bracketing | Standard-sample bracketing against NIST SRM 3141a | "To correct for instrument mass bias, the sample–standard bracketing technique was used for all analyses" | Internal normalization to 49Ti/47Ti = 0.749766 using the exponential law, plus bracketing against an in-house Alfa Aesar Ti wire standard — "the isotope data were normalized to a 49Ti/47Ti ratio of 0.749766 (ref. 72), using the exponential law"; results reported "applying the sample–standard bracketing method" (p.8)
Mass Fractionation Law | Mathematical law used to parameterize instrumental isotope mass fractionation as a function of mass difference. The linear law assumes fractionation proportional to mass difference. The power law is an alternative formulation. For double-spike procedures, the law is embedded in the inversion algorithm and must be consistent between spike calibration and sample data reduction. | Basic | Read-Only | Controlled list | Exponential ¦ Linear ¦ Power ¦ N/A ¦ None | Source: MC-ICP-MS module | 2026-08-30 | (none) | The exponential law (Maréchal et al. 1999) is most widely used and theoretically best justified for MC-ICP-MS. |  | Exponential law | N | Exponential law | N | Power-law form used in the double-spike inversion equation | N | N | N | N | Exponential law | N | N | N | N
Double-Spike Isotope Pair | The pair of enriched isotopes used to construct the double spike. Both isotopes must be naturally rare and free from isobaric interferences. The double spike composition must be accurately calibrated before use. Specify both isotope masses and the spike certificate or calibration reference. Record 'N/A' where the procedure does not use a double spike. | Basic | Read-Only | Text (free) | 233U–236U (IRMM-3636) ¦ 117Sn–122Sn (in-house calibrated) ¦ 57Fe–58Fe ¦ 61Ni–62Ni ¦ 204Pb–207Pb ¦ N/A ¦ None | Source: MC-ICP-MS module | 2026-08-30 | (none) | The pair is chosen to minimise error magnification in the inversion. |  | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | 91Zr-96Zr | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used
Double-Spike Mixing Ratio | Target proportion of double-spike signal relative to total target species signal in the spiked mixture, expressed as spike fraction (0–1) or spike:sample ratio. The optimum is target species-system specific and is typically determined using the Double Spike Toolbox or equivalent. The achieved mixing ratio may deviate from the target within acceptable bounds (typically ±20% of optimal); the double-spike inversion corrects for actual mixing ratios. Record 'N/A' where the procedure does not use a double spike. | Advanced | Editable | Text (free) | ~0.40 (spike fraction; optimal for 117Sn–122Sn) ¦ ~0.5 spike:sample ¦ N/A ¦ None | Source: MC-ICP-MS module | 2026-09-01 | (none) | An optimal mixing ratio minimises error propagation through the double-spike inversion. |  | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | 0.43:0.57 spike-to-sample Zr mass ratio, described as optimal | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used
Double-Spike Inversion Algorithm | Mathematical algorithm used to simultaneously solve for the true sample isotope ratio and the mass fractionation factor from the measured spike-sample mixture. Common implementations: iterative Newton-Raphson solution, matrix algebraic approach, or published software (e.g., Rudge et al. 2009 Double Spike Toolbox). Specify the algorithm or software reference and version. Record 'N/A' where the procedure does not use a double spike. | Basic | Read-Only | Text (free) | Rudge et al. (2009) Double Spike Toolbox ¦ Iterative Newton-Raphson (Compston & Oversby 1969) ¦ In-house spreadsheet implementing exponential-law inversion ¦ N/A ¦ None | Source: MC-ICP-MS module | 2026-08-30 | (none) |  |  | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | "a minimization approach implemented in Mathematica, taking into account all ratios (i.e., 91/90Zr, 92/90Zr, 94/90Zr, and 96/90Zr) with different weighs being assigned to each based on their associated uncertainty", cross-checked against two exact three-ratio solutions | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used | N/A — no double spike used
Internal Normalization Element and Isotope Ratio | Element and isotope ratio used as the internal normalization standard for mass bias correction. Specify the element, the monitored isotope ratio (e.g., 113Cd/111Cd for Sb measurements), and the certified or consensus value used for correction. Record 'N/A' where the procedure does not use internal normalization. | Basic | Read-Only | Text (free) | e.g., 60Ni/58Ni = 0.38579 (from IUPAC 2021) for Fe-Ni doping ¦ 65Cu/63Cu = 0.44561 for Fe-Cu doping ¦ 113Cd/111Cd = 0.95542 (Abouchami et al. 2012) for Sb ¦ 205Tl/203Tl = 2.38714 for Pb ¦ N/A ¦ None | Source: MC-ICP-MS module | 2026-09-01 | (none) | The measured ratio of the normalizing element is used to calculate the mass bias factor β using the chosen mass fractionation law, which is then applied to correct the target species isotope ratios assuming β_norm ≈ β_species. |  | 98Mo/96Mo = 1.453173 | N | 57Fe/56Fe = 0.023095 or 57Fe/54Fe = 0.362549, the certified ratios of IRMM-014 | N | N | N/A — Rb has two stable isotopes, so internal normalization is not possible; bracketing used instead | N | N | N/A — Rb has two stable isotopes; bracketing used instead | 94Zr/90Zr = 0.3381 (Minster & Ricard 1981) | N | N | N | N
Isotope Ratio Reported | Specific isotope ratio pair(s) reported as primary data output. The denominator isotope is typically the reference mass used in the delta notation (e.g., 54Fe for Fe isotopes, 235U for U isotopes). Report all ratio pairs routinely calculated and reported. | Basic | Read-Only | Text (free) | e.g., 56Fe/54Fe and 57Fe/54Fe ¦ 238U/235U ¦ 122Sn/118Sn ¦ 123Sb/121Sb ¦ 206Pb/204Pb; 207Pb/204Pb; 208Pb/204Pb ¦ N/A ¦ None |  | 2026-08-31 | reported property |  |  | iMo/96Mo | 34S/32S and 33S/32S | Fe isotope ratios relative to IRMM-524a | N | 91/90Zr, 92/90Zr, 94/90Zr, 96/90Zr | 87Rb/85Rb | 187Os/188Os, 186Os/188Os, 184Os/188Os | 187Os/188Os, 186Os/188Os, 184Os/188Os | 87Rb/85Rb | 91Zr/90Zr, 92Zr/90Zr, 96Zr/90Zr | N | N | 41K/39K, 65Cu/63Cu, 66Zn/64Zn | 46Ti/47Ti, 48Ti/47Ti, 50Ti/47Ti — εiTi is defined on iTi/47Ti "where i refers to the isotope masses 46Ti, 48Ti and 50Ti" (p.8)
delta or epsilon Value Reference Standard | International or community-accepted isotopic reference standard used as the zero-delta anchor for expressing isotopic compositions in delta (δ, per mil) or epsilon (ε, per ten thousand) notation: δⁱX = [(ⁱX/ʲX)sample / (ⁱX/ʲX)reference − 1] × 1000‰. The reporting reference is a per-study or per-publication decision. Specify the standard name, lot or batch number where applicable, and the certified or consensus isotope ratio used for normalization. | N/A | Basic | Text (free) | e.g., IRMM-014 (Fe; Taylor et al. 1992) ¦ CRM-112a / SRM-960 (U) ¦ NIST SRM 981 (Pb) ¦ NIST SRM 3102a (Sb) ¦ Sn_IPGP (in-house Sn standard) ¦ JMC Lyon (Zn) ¦ N/A ¦ None |  | 2026-09-01 | target species |  |  | Alfa Aesar solution standard | V-CDT scale via IAEA-S-1, S-2, S-4 and NBS-123 | IRMM-524a, "that has an identical isotopic composition to IRMM-014" | OL-REE series, prepared in-house from high-purity ESPI oxide powder | ZrNIST | NIST SRM984 | N | N | NIST SRM984 RbCl; the basalt geostandard BCR-2 used as an alternative bracketing standard in some sessions | NIST SRM 3169 | IRMM-014 (Fe), SRM979 (Cr), DTS-2b (Mg) | NIST SRM 3141a | NIST-SRM 3141a (K), NIST-SRM 976 (Cu), JMC-Lyon (Zn) | N
Calibration Strategy per Target Species | Approach used to convert measured ion signals to reported concentrations, and specifically any case where different target species or target species groups within one procedure are calibrated differently - different primary standards for different mass ranges or phases, or one element serving as internal standard while others are externally calibrated. Where a single strategy applies to all target species, record that strategy. Where the procedure reports isotope ratios only and no concentrations, record 'Not applicable (isotope ratios only)'. | Basic | Read-Only | Text (free) | External calibration (all target species) ¦ Isotope dilution (all target species) ¦ ID for Zn (67Zn spike); external calibration for all other elements ¦ External calibration + standard addition ¦ N/A ¦ None | Source: ICP-MS module | 2026-09-01 | target species |  |  | N | N | N | One OL-REE series standard per REE, with separate cup configurations for Dy and Yb | N | N | N | N | N | N | One bracketing standard per analyte — IRMM-014 for Fe, SRM979 for Cr, DTS-2b for Mg | N | One bracketing standard per analyte — NIST-SRM 3141a for K, NIST-SRM 976 for Cu, JMC-Lyon for Zn; run concentrations also differ by analyte (200 ppb for K and Zn, 100 ppb for Cu) | N
Signal Integration Interval Method | Method used to define which portion of the acquired signal is integrated for each reported measurement, and the rule for excluding unstable or anomalous portions. State whether the interval is the full acquisition, a stable sub-interval selected by inspection, or selected by an automated criterion, and name the criterion. | Basic | Read-Only | Controlled list / Text | Full acquisition ¦ Stable interval selected by inspection ¦ Automated threshold ¦ Automated outlier rejection ¦ Fixed time window ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-26 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Signal Integration Time | The integration time actually used for the reported signal, in seconds - the time retained after the signal-integration interval has been applied, and therefore equal to or shorter than the total acquisition. Where the procedure computes it from an acquisition structure, record that basis. | N/A | Basic | Numeric (s) | Calculated from blocks × cycles per block × integration time per cycle. | Source: ICP-MS module | 2026-08-26 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Blank / Background Correction Method | Method used to subtract background and, where relevant, procedural blank from the measured signal. State what is subtracted - instrument background acquired with no sample reaching the plasma, an on-peak zero acquired on the same matrix, or a separately measured procedural blank - over what interval it is measured, and at what granularity it is applied. | Basic | Read-Only | Controlled list / Text | Gas blank ¦ On-peak zero ¦ Solution blank ¦ Procedural blank ¦ On-peak zero + procedural blank ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-26 | (none) | For isotope ratio procedures where procedural blank is minor relative to the sample signal, on-peak zero is often sufficient. |  | On-peak-zero baseline integrations subtracted | N | On-peak zero from a blank solution subtracted from all measurements | N | On-peak-zero correction using mean acid blank intensities | N | Corrections applied offline for abundance sensitivity, W and Re atomic interferences and instrumental mass bias | N | N | "An on-peak background correction was performed"; background corrections averaged 0.3, 2 and 98 ppm for 91Zr/90Zr, 92Zr/90Zr and 96Zr/90Zr | On-peak baseline measurement preceding each analysis | N | N | N
Spike / Outlier Filtering Approach | Criteria used to identify and exclude anomalous data - signal spikes, individual cycles, or whole replicate measurements - before the reported value is calculated. State where in the reduction sequence the filter is applied. | Advanced | Editable | Text (free) | 2σ exclusion on per-cycle ratios ¦ Grubbs test ¦ Visual inspection ¦ None ¦ N/A | Source: ICP-MS module | 2026-08-30 | (none) |  |  | N | N | N | N | N | N | N | N | "any ratio outside 2σ was discarded" | N | N | N | N | N
Spectral Interference Corrections Applied | Whether mathematical corrections for isobaric, polyatomic or residual interferences are applied in data reduction, supplementary to any suppression already achieved by chemical separation, mass resolution, or a collision/reaction cell. Detail for each affected mass is carried by Interfering Species and Interference Correction Method. | Basic | Read-Only | Controlled list / Text | e.g., 'Yes — ⁸⁷Rb on ⁸⁷Sr, corrected from the ⁸⁵Rb monitor' ¦ 'Yes — doubly-charged Er and Yb on the Sr and Rb masses' ¦ 'No explicit corrections applied; medium resolution resolves the polyatomic interferences' ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-27 | (none) | In solution MC-ICP-MS, chemical separation is the primary interference mitigation; mathematical corrections address residual contributions. |  | Yes — Zr and Ru on Mo masses | Yes — the paper is built around correcting isobaric interferences and mass bias for S | Yes — 54Cr on 54Fe and 58Ni on 58Fe, plus argide molecular interferences | Yes — Er on Dy, and neighbouring-element isotopes generally | Yes — Mo on Zr | Yes — 87Sr on 87Rb | Yes — W and Re on Os | Yes — W and Re on Os | Yes — 87Sr on 87Rb, Sr isotopes monitored | Yes — Mo and Ru on Zr, plus Fe- and Cr-argides and Ar-Ar-oxides | Yes — 49Ti, 51V and 56Fe monitored to correct interferences on Cr | Yes — 40Ar1H+ on 41K+, resolved by peak-shoulder measurement | Yes — ArH+ on K, lowered by dry plasma and a high-resolution slit | N
Interfering Species | The isobaric, polyatomic and doubly charged species that overlap the measured masses and are corrected in data reduction - direct isobars, oxides and argides, hydrides, and abundance-sensitivity tailing from an adjacent large beam. Name each species and the mass it affects. | Basic | Read-Only | Text (free) | e.g., 54Cr+ on 54Fe+; 58Ni+ on 58Fe+ ¦ 204Hg+ on 204Pb (corrected via 202Hg monitor) ¦ 238U abundance sensitivity tail on 235U and 236U ¦ None ¦ N/A ¦ Includes direct isobars (e.g., 54Cr+ on 54Fe+, 58Ni+ on 58Fe+), abundance sensitivity tailing (e.g., 238U tail on 235U and 236U in U isotope measurements), and hydride interferences (e.g., 238UH+ at mass 239). | Source: ICP-MS module | 2026-09-10 | monitored property |  |  | Zr and Ru, monitored on 91Zr and 99Ru | N | 54Cr on 54Fe, 58Ni on 58Fe, and argide ions 40Ar13C+, 40Ar14N+, 40Ar16O+, 40Ar16O1H+ and 40Ar18O+ | 162Er and 164Er on 162Dy and 164Dy | Mo, monitored on masses 95 and 98 | 87Sr on 87Rb, monitored at 88Sr | 182W/184W/186W and 185Re/187Re | W and Re | 87Sr on 87Rb; also double-charged Er and Yb isotopes | 95Mo, 99Ru, 101Ru; 40Ar2 16O+ and 40Ar2 14N+ on 96Zr+ and 94Zr+ | 49Ti, 51V, 56Fe on the Cr masses | 40Ar1H+ on 41K+ | ArH+ on K | N
Interference Correction Method | Equation or procedure used to calculate and remove each interference contribution, together with how its magnitude was established - a monitor mass measured simultaneously and scaled by natural abundance ratios, a production-rate factor measured on a reference material or interference standard solution, or a tailing factor measured on a pure standard. Name the reference material used. | Basic | Read-Only | Text (free) | e.g., Monitor mass correction: 202Hg signal × (204Hg/202Hg)natural to correct 204Pb ¦ Abundance sensitivity tailing factor measured weekly from pure U standard ¦ None ¦ N/A | Source: ICP-MS module | 2026-09-10 | monitored property |  |  | Monitoring 91Zr and 99Ru | N | Simultaneous monitoring of 53Cr and 60Ni on 10^12 Ω amplifiers; argides resolved by measuring on the flat-topped peak shoulder in MR or HR mode | Subconfiguration scaling — 166Er measured only in the subconfiguration, 163Dy in both; the 166Er/163Dy ratio is scaled onto the main configuration and multiplied by 162Er/166Er and 164Er/166Er to subtract the interferences | A doping calibration: "Offset δ94/90ZrNIST = −63.4 · x" where x is the total Mo/Zr atomic ratio; the reported data are uncorrected for residual Mo | Correction assuming a constant terrestrial 87Sr/88Sr = 0.085, with a sensitivity test over 0.0835–0.0885 showing a maximum shift of 0.008 permil | Monitor isotopes 184W, 186W and 187Re; abundance sensitivity determined by scanning the low mass tail of a 30 V 192Os beam using the SEM and fitting a curve to the tail — 0.5 to 1 ppm in the Os mass range | Online W and Re correction; abundance sensitivity greater than 1 ppm "since the analyser pressure was higher at ~2 x 10^-8 mbar" | N | Mo and Ru corrections calculated from the signals on mass 95, 99 and 101, with an initial Mo correction using a mass bias relative to 91Zr/90Zr = 0.21798 | Simultaneous monitoring of 49Ti, 51V and 56Fe | N | N | N
Isotope Dilution Data Reduction Method | Mass balance approach used to calculate sample mass fractions, and where applicable isotope ratios, from spike-sample isotope ratio measurements. For double-spike procedures the ratios and concentrations are co-determined through the spike inversion; for conventional isotope dilution the IUPAC ID equations are used. Record 'None' where neither isotope dilution nor double-spike methods are applied. | Basic | Read-Only | Controlled list / Text | IUPAC isotope dilution equations ¦ Double-spike inversion ¦ ID-IS (isotope dilution internal standardisation) ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-31 | (none) |  |  | N/A — no isotope dilution applied | N/A — no isotope dilution applied | N/A — no isotope dilution applied | N/A — no isotope dilution applied | Double-spike inversion, Rmeas = [p·RSpike + (1−p)·RStd·(Mx/Mn)^α]·(Mx/Mi)^β, solved by weighted minimisation over four ratios | N/A — no isotope dilution applied | N/A — no isotope dilution applied | N/A — no isotope dilution applied | N/A — no isotope dilution applied | N/A — no isotope dilution applied | N/A — no isotope dilution applied | N/A — no isotope dilution applied | N/A — no isotope dilution applied | N/A — no isotope dilution applied
Memory Effect Mitigation | Procedure applied to identify and minimise carry-over of high-concentration or isotopically distinct material from a preceding measurement into the current one. Mitigation is applied primarily at measurement time, by allowing sufficient washout or rinse between successive introductions. At data processing level, record any flagging or exclusion of measurements where the required washout may not have been achieved. | Advanced | Editable | Text (free) | e.g., Extended acid rinse between samples; first measurement excluded if preceded by high-signal sample ¦ Blank measurement inspected before each sample; session halted if blank > 0.5% of sample signal ¦ None required for this target species suite ¦ N/A ¦ None | Source: ICP-MS module | 2026-09-01 | (none) |  |  | N | Wash-out 2 min for solution | 210 s washout between all measurements | N | On-peak-zero acid blank before each sample "to account for blank contribution as well as any 'memory' effects from the Aridus II sample introduction system during the run" | 60 s wash in 0.45 M HNO3 | Desolvating nebulisers deliberately avoided because of "severe memory problems for Os"; ESI PFA-50 low-uptake nebuliser and GE Cinnabar micro-cyclonic spray chamber chosen "in the hope these would reduce the long Os washout times and poor memory usually associated with solution introduction of Os"; wash acid aspirated until the 192Os beam fell to background — a 99.99% decrease reached after 220 s for DTM | TD 3 mol/l HCl aspirated between analyses until the Os beam decreased to acceptable background levels | N | N | N | N | N | N
Uncertainty Level | The convention at which reported uncertainties are quoted — 1-sigma, 2-sigma, or a 95% confidence interval — and whether a measured spread is reported as a standard error or a standard deviation. State the convention applying to all values reported under this procedure, and state each separately where different reported quantities use different conventions. Distinct from Uncertainty Propagation Method, which addresses how component uncertainties are combined rather than how the result is quoted. | Basic | Read-Only | Controlled list / Text | 2σ (95% confidence) ¦ 1σ (68% confidence) ¦ 95% confidence interval ¦ 2 standard errors (2SE) ¦ 1 standard deviation (1SD) ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-26 | (none) | A reported uncertainty is not interpretable without it: the same numeric value means different things at 1-sigma and at 95% confidence. |  | 2 s.d. for external reproducibility (n = 24 for Mo, n = 14 for Ba) | "external reproducibility is reported at the 2σ error level"; long-term reproducibility "typically 0.20‰ and 0.45‰ (2σ) for solution and laser" | N | Not stated in the section read | "the external reproducibility (at 2σ) of the spiked ZrNIST measurements from each run, which in all cases was similar in magnitude or slightly larger than the internal uncertainty determined from counting statistics" | N | 2SD for short- and long-term reproducibility; within-run errors as "2 standard errors of the mean (2SE = 2SD/n^0.5; where n = 45 for the Neptune analyses" | 2SD and 2SE, with n = 50 for the Nu Plasma analyses | "the 2 standard error (2se) is reported unless stated otherwise"; for samples analysed fewer than 3 times, "the largest 2 se reported for a sample analyzed multiple times has been used" | Both quoted: "external precision expressed as 2 standard deviations (2SD)" and 2SE per analysis | "the mean and 2 x standard error (SE) of ten individual standard-bracketed sample analyses" | Stated as ± values on δ41K without an explicit convention in the section read | 2 s.d. | 2 s.d.
Uncertainty Propagation Method | The approach used to propagate analytical uncertainty through the data reduction chain to the final reported value. State which sources are included in the propagation: counting statistics, calibration standard uncertainty, internal standard uncertainty, drift correction, and any systematic contributions. Distinct from Uncertainty Level, which states the convention at which the resulting uncertainty is quoted. | Advanced | Editable | Text (free) | 2σ of bracketing standard reproducibility propagated to sample ¦ Monte Carlo simulation through double-spike inversion ¦ ISO GUM quadrature propagation ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-30 | reported property |  |  | N | N | N | N | External reproducibility of the spiked ZrNIST measurements from each run adopted per determination, compared against internal counting-statistics uncertainty | N | N | N | "Errors are determined from repeated measurements" | N | N | N | N | N
Normalization / Standards-Based Correction | Post-acquisition normalization applied to the reported data beyond the primary calibration — for example correction to a reference value derived from secondary reference materials, or correction for a systematic bias those materials reveal. Record 'None' if no additional normalization is applied. | Advanced | Editable | Text (free) | Spike isotope decay correction applied ¦ Age correction for U-series disequilibrium ¦ None ¦ N/A | Source: Composition QC module | 2026-08-30 | reported property |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Calibration Factor and Determination Method | An externally-calibrated factor that converts the measured quantity into the reported quantity, how it was determined, and its uncertainty. Applies where the conversion depends on a factor calibrated against a reference of independently known value, rather than on the instrument response alone. Distinct from the fields that name the calibration material and that state which approach applies to which target species, where the technique has them: this field records the resulting factor itself. | Basic | Editable | Text (free) |  | Source: Calibration Factor module | 2026-09-01 | reported property |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Procedural Blank Level | The measured level of the analytical blank in the session, and — where the reported quantity is a ratio — its composition, since a blank subtracted from a ratio biases the result unless its own composition is known. Companion to the blank correction method. | N/A | Basic | Text (free) |  | Source: Blank module | 2026-09-01 | target species | Follows the criterion-versus-measurement split the library applies wherever a procedure sets a threshold and an analysis reports a value against it. |  | "Total procedural blanks were between 0.7 and 1.2 ng and thus negligible, given that several hundred ng of Mo were analyzed for each sample" | "The procedural blank, resulting from chemical processing and purification is ~0.05% (~0.25 µg per 500 µg S used for column chemistry)" | "the total procedural blank is ~70 ng and thus negligible considering that 1-2 mg Fe was purified for each sample" | N | "The total mass of non-radiogenic Pb measured in our FC-1 zircon and baddeleyite fractions is indistinguishable from the range of Pb determined in total procedural blanks" | "The Rb blank of the procedure (digestion and column chemistry) is ~0.14 ng, which accounts for less than 0.5% of total Rb from a typical sample (40 ng)" | N | N | N | "Total procedural blanks prepared together with Tarda and Tagish Lake contained 0.08 and 0.24 ng Zr, while total blanks treated alongside Ivuna were 0.09 and 0.13 ng Zr" | N | N | N | "The total procedural blank for Ti was 3.7 ng, resulting in a maximum blank contribution of 0.18% for Ti"
Analysis Inclusion and Rejection Criteria | The rules determining which individual results contribute to a reported aggregate value, together with the outcome of applying them: how many results were obtained, how many were included, and on what grounds any were excluded. An individual result is the value of the reported quantity obtained from one acquisition: a replicate measurement of the same solution or location, a spot or grain within a sample, or an independently prepared aliquot or digestion, whichever the procedure aggregates. Distinct from filtering the acquired signal during data reduction (removing spikes, cycles or scans, or discarding an acquisition whose signal is compromised): this field records which finished results enter the reported value, and on what grounds. | Basic | Basic | Text (free) |  | Source: Aggregation module | 2026-09-15 | (none) | Criteria and outcome are combined in one field because neither is interpretable without the other, following the precision/accuracy precedent. |  | Partially — "For samples analyzed several times, reported values represent the mean of pooled solution replicates". No acceptance or rejection rule stated | N | N | Partially — "On average, LREEs were measured nine times"; replicate matrix cuts were measured but "are not used, however, for data interpretation to avoid unnecessary influence of stable isotopic fractionation potentially induced by Mo chemistry" — an explicit exclusion, on chemical rather than statistical grounds | Partially — Table 1 records "Number of times the same purified Zr solution was measured independently in the MC-ICP-MS" and "Reported values are weighted means of all replicate" analyses. No rejection rule stated | N | Partially — n = 45 per analysis. No rejection rule stated | Partially — n = 50 per analysis. No rejection rule stated | N — no result-level rule is stated. Reported values are "averages of repeated measurements of each sample when multiple analyses were possible", with no criterion for admitting or rejecting a measurement; the "any ratio outside 2σ was discarded" rule acts within a measurement and is recorded under Spike / Outlier Filtering Approach | Partially — n stated per reference material (n = 13–99 for terrestrial RMs over 10 months; n = 17–38 for eucrites and Colony; n = 32 and n = 37 for standard sessions). No rejection rule stated | Partially — "the mean ... of ten individual standard-bracketed sample analyses"; "Samples were typically analysed two to four times". No rejection rule stated | Partially — "Each sample was measured approximately 20 times". No rejection rule stated | N | N
Error Correlation Between Reported Quantities | The correlation coefficient between pairs of reported quantities whose uncertainties are not independent, together with the pair it applies to and how it was obtained. | Advanced | Basic | Numeric (dimensionless) / Text | e.g., 'rho(206Pb/238U, 207Pb/235U) = 0.83, from the data reduction software error propagation' ¦ 'rho = 0.31 (double-spike inversion covariance)' |  | 2026-08-27 | pair: reported property | Concordia and isochron regressions cannot be reconstructed without it. |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Constants and Reference Values Used | Physical constants and reference values used in data reduction to calculate the final reported quantity (e.g., decay constants for age calculation, standard isotope ratios, or other citable reference values used in a correction or calculation), together with their source. Distinct from the Group 6 reference-material fields, which document accepted values for specific calibration/validation materials rather than universal physical constants. Record "None" if no citable, revisable physical constants feed into this procedure's data reduction. | Basic | Editable | Text (free) | e.g., "λ238U = 1.55125×10⁻¹⁰ yr⁻¹, λ235U = 9.8485×10⁻¹⁰ yr⁻¹ (Jaffey et al. 1971); 238U/235U = 137.818 (Hiess et al. 2012)" ¦ "None" | Source: Core module | 2026-08-26 | (none) |  |  | 98Mo/96Mo = 1.453173 for internal normalization; 134Ba/136Ba = 0.3078 (Carlson et al. 2007) for the TIMS half | N | 57Fe/56Fe = 0.023095 and 57Fe/54Fe = 0.362549, "the certified ratios of IRMM-014" (Craddock and Dauphas, 2010) | N | 238U/235U = 137.818 (45), 18O/16O = 0.00205 (44), and α = 0.18 ± 0.02%/amu from repeat NBS-981 analyses; U decay constants of (47); Th/U[magma] = 2.8 ± 1.0 for the initial 230Th disequilibrium correction | 87Sr/88Sr = 0.085, "which is the terrestrial Sr ratio", used for the 87Sr interference correction; sensitivity tested at 0.0835 and 0.0885 | N | N | N | 94Zr/90Zr = 0.3381 and 91Zr/90Zr = 0.21798, both Minster & Ricard (1981) | N | N | N | N
6. Quality Control & Uncertainty | [6 empty] | 2026-08-26 |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Primary Calibration Standard Name | Name and reference material identifier of the primary reference material(s) against which the instrument is calibrated — converting raw signal intensities to concentrations, or anchoring an isotope ratio as the bracketing standard or zero-delta reference. Give the material name, its source or supplier, and a citation for the accepted values used. Where calibration instead uses the vendor's stored library or theoretical response factors rather than measured reference materials — 'standardless' or 'semi-quantitative' quantification — record that here, naming the library or model used. 'None' means no calibration was performed at all, which is a different answer. | Basic | Editable | Text (free) | e.g., IRMM-014 (Fe isotopes; Taylor et al. 1992) ¦ CRM-112a / SRM-960 (U isotopes) ¦ NIST SRM 981 (Pb isotopes) ¦ JMC Lyon Zn standard ¦ NIST SRM 3102a (Sb) ¦ In-house single-element standard solution: specify ¦ N/A ¦ None | Source: Composition QC module | 2026-09-01 | target species | Results calibrated against different published values for the same material are not directly comparable. |  | Alfa Aesar Mo solution standard | In-house S_Alfa and S_Spex 20 ppm S solutions, calibrated against IAEA-S-1, S-2, S-4 and NBS-123 | IRMM-524a | OL-REE series | ZrNIST | NIST SRM984 | UMd, DTM, LOsST and DROsS Os reference materials | DTM and LOsST | NIST SRM984 RbCl; BCR-2 as an alternative bracketing standard in some sessions | NIST SRM 3169 | IRMM-014, SRM979, DTS-2b | NIST SRM 3141a | NIST-SRM 3141a, NIST-SRM 976, JMC-Lyon | N
Calibration Standard Measurement Frequency | How often the primary calibration or bracketing standard is measured relative to unknown samples within a session. Specify the bracketing scheme where a fixed one is used, such as one standard measurement before and after each sample. | Basic | Read-Only | Text (free) | e.g., SSB: one standard before and one after every sample ¦ One standard every 3 samples ¦ Start and end of session only ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-31 | (none) |  |  | N | N | N | Every sample, spaced by 300 s rinsing | Every sample — "Each sample measurement was individually bracketed" | N | N | N | Every sample (bracketing), plus an external pure Rb solution "during each analytical session" | Each session — "The Zr standard material NIST SRM 3169 was analyzed in each session" | N | Every sample (bracketing) | N | N
Secondary Reference Materials | Quality-control reference material(s) measured as unknowns alongside samples to assess accuracy independently and to monitor drift. Give the material name, its source, and a citation for the accepted or reference values used for comparison. | Basic | Editable | Text (free) | e.g., BHVO-2 (Fe isotopes; Dauphas & Rouxel 2006 compilation) ¦ BCR-2 ¦ NIST SRM 960 or equivalent U standard ¦ NIST SRM 981 (Pb) ¦ AGV-2 ¦ N/A ¦ None | Source: Composition QC module | 2026-08-31 | defines: standard |  |  | BHVO-2, "several digestions of which were processed through the full analytical protocol and analyzed together with each set of samples" | Sch-M-2 anhydrite mineral standard; geological reference samples with known isotope compositions | BHVO-2 and BCR-2 | N | N | BHVO-2, BCR-2, BE-N, W-2, AGV-2, GSR-1, GS-N, G-A, G-3; DTS-2b and PCC-1 synthetic mixes; Allende | N | N | BCR-2, AGV-2, BHVO-2, GS-N and other terrestrial rocks | BHVO-2, BCR-2, AGV-1, SCo-1; eucrites Bouvante and Bereba; CO chondrite Colony | BHVO2 and DTS-2b, "processed alongside the samples" | BHVO-2 | BHVO-2 | N
Counting Statistics Error | Uncertainty predicted from counting statistics — the theoretical limit set by the Poisson distribution of the counts accumulated — for each reported quantity per analysis, with the sigma level stated. Derived from the counts on the target species together with those on any background or blank subtracted from it. Distinct from the scatter actually observed within a measurement or between repeated measurements, which is recorded separately. | Advanced | Basic | Text (free) | e.g., '2SE predicted from counting statistics over 45 cycles' ¦ '~0.6% (1 sigma), counting statistics plus propagated uncertainties' ¦ 'N/A' | Source: Composition QC module | 2026-09-01 | sample > sampling unit x reported property | Where a procedure reports both, agreement indicates the measurement is shot-noise limited, and a larger observed scatter indicates a further source of variance. |  | N | N | N | N | Internal uncertainty determined from counting statistics, used as the comparison against the external reproducibility adopted per determination; value not tabulated | N | N | N | N | N | N | N | N | N
Internal (Within-Measurement) Analytical Precision and Assessment Method | Precision of a single measurement, derived from the scatter of the cycles, sweeps or integrations that make it up, together with the method used to assess it. State the statistic (2SE, 2SD, 1s RSD), the number of cycles it is computed over, and the reported quantity it applies to. Distinct from Counting Statistics Error, which records the uncertainty predicted from the counts rather than the scatter observed; where a procedure reports both, record the observed value here and the predicted value there. | Advanced | Basic | Text (free) | e.g., '2SE = 2SD/sqrt(n), n = 45 cycles' ¦ '+/-0.001-0.002 (2SE) on 206Pb/204Pb' ¦ '~2.6% (2SE) on single-spot ages' ¦ 'N/A' | Source: ICP-MS module | 2026-08-26 | sample > sampling unit x reported property | This is the finest of the three precision levels the library records, below within-session and between-session precision: it describes the repeatability of one analysis rather than agreement between analyses, and is normally the smallest of the three. |  | N | N | N | N | Internal uncertainty determined from counting statistics; stated to be similar in magnitude to or slightly smaller than the external reproducibility adopted per determination | N | Within-run errors for individual analyses quoted as 2 standard errors of the mean, 2SE = 2SD/n^0.5, with n = 45 cycles | Within-run errors for individual analyses quoted as 2 standard errors of the mean, 2SE = 2SD/n^0.5, with n = 50 cycles | N | N | N | N | N | N
Within-Session Analytical Precision and Assessment Method | Precision of repeated measurements within a single analytical session and the method used to assess it. Report both the assessment method and the precision values. The assessment method must specify the reference material or standard measured, the number of replicates n, and the statistic reported (1s RSD, 2s RSD, 2SD, 2SE, 95% CI). Distinct from the internal precision of a single measurement, which derives from counting statistics over the cycles of that measurement rather than from repeated analyses. | Basic | Basic | Text (free) | e.g., 2SD = 0.03‰ on δ56Fe (n=10 BHVO-2); session ¦ 2σ = 0.004‰ on δ238U (n=9 CRM-112a brackets) ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-31 | standard x reported property |  |  | N | N | N | N | External reproducibility at 2 sigma of the spiked ZrNIST measurements from each run, adopted as the uncertainty on each determination and stated to be similar to or slightly larger than the internal counting-statistics uncertainty | N | Short-term reproducibility of standards analysed in a single analytical session, quoted as 2 standard deviations (2SD). Distinct from the within-run internal error, which the paper quotes separately as 2SE of the mean, 2SE = 2SD/n^0.5 with n = 45 for the Neptune and n = 50 for the Nu Plasma analyses | Short-term reproducibility of standards analysed in a single analytical session, quoted as 2 standard deviations (2SD). Distinct from the within-run internal error, which the paper quotes separately as 2SE of the mean, 2SE = 2SD/n^0.5 with n = 45 for the Neptune and n = 50 for the Nu Plasma analyses | N | N | N | N | N | N
Between-Session (Long-Term) Analytical Precision and Assessment Method | Precision of measurements across multiple analytical sessions over weeks to months — long-term or intermediate precision — and the method used to assess it. Report both the assessment method and the precision values, specifying the reference material, the number of measurements and sessions, the time span covered, and the statistic reported. | Advanced | Basic | Text (free) | e.g., 2SD = 0.030‰ on δ56Fe (n=23 sessions; BHVO-1; Dauphas et al. 2009) ¦ 2SD = 0.005‰ on δ238U (n > 50 analyses; CRM-112a) ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-31 | standard x reported property | Long-term precision is normally poorer than within-session precision and is the figure a data user should carry when comparing results from different sessions. |  | External reproducibility from repeated BHVO-2 measurements: ±0.14 for ε97Mo to ±0.39 for ε92Mo (2 s.d., n = 24); Ba ±0.13 for ε135Ba to ±0.31 for ε138Ba (2 s.d., n = 14) | "Long-term reproducibility of S isotope compositions is typically 0.20‰ and 0.45‰ (2σ) for solution and laser"; long-term reproducibility of in-house solution standards within ±0.2‰ | N | N | External reproducibility at 2σ of the spiked ZrNIST measurements from each run | N | N | N | "the long-term reproducibility was ±0.01‰ (n = 40)" from a pure Rb ICP-MS solution run as an external standard each session | Terrestrial RMs measured over 10 months (n = 13–99) give average 2SD of 0.3, 0.2 and 1.0 for ε91Zr, ε92Zr and ε96Zr; "The external precision estimated from the geological sample measurements integrates the uncertainty introduced by the chemical separation procedure and mass spectrometry" | N | N | N | N
Analytical Accuracy and Assessment Method | Offset between measured and accepted values for secondary reference materials, and the method used to assess it. Specify the reference material and the source of its accepted values, the number of analyses, and the quantities assessed. Report systematic biases and their likely causes. Express the offset in the form appropriate to what the procedure reports - percent relative bias for concentrations, or deviation in delta or ratio units for isotopic quantities. | Basic | Basic | Text (free) | e.g., % deviation from published δ56Fe for BHVO-1 compilation value ¦ Agreement with certified δ238U of IRMM-3636 spike solution ¦ Z-score relative to interlaboratory consensus ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-31 | standard x reported property |  |  | "The εiMo values obtained for BHVO-2 are indistinguishable from the Alfa Aesar standard, demonstrating that the Mo isotopic data are accurate" | Assessed against IAEA and NBS reference materials on the V-CDT scale and against geological reference samples with known compositions | N | Assessed in a dedicated "Assessment of data accuracy" section, using replicate matrix cuts and a processed geostandard | N | NIST SRM984 treated as a sample, plus synthetic DTS-2b+SRM984 and PCC-1+SRM984 mixes, "gave δ87Rb values of zero within error"; geostandards and Allende "yielded reproducible results that agree with literature data" | N | N | An aliquot of SRM984 passed through the full chemistry gave δ87Rb = 0.00 ± 0.03‰, "confirming that no isotope fractionation is caused by the Rb purification procedure"; Allende duplicate splits agreed at 0.12 ± 0.02‰ and 0.14 ± 0.04‰ | Terrestrial and meteorite reference materials measured repeatedly to verify data quality; doping tests with Ti, V, Cr, Mo, Hf and W showed "the observed trace levels have no effect on the accuracy of the Zr isotope data" | BHVO2 and DTS-2b processed alongside the samples | "The average d41K value for BHVO-2 was −0.448 ± 0.027‰ which is within error of its previously reported values, for example, −0.46 ± 0.09‰ (Wang et al., 2021)" | "To monitor data quality, the geostandard BHVO-2 was analysed alongside all sample analyses" | BHVO-2 and the Agua Zarcas (CM2) chondrite analysed alongside Bennu — "To verify the accuracy and reproducibility of these measurements, the terrestrial rock standard BHVO-2 and the Agua Zarcas (CM2) chondrite were analysed alongside the Bennu sample. The analytical uncertainties of 9 analyses of BHVO-2 are ±0.17 ε46Ti, ±0.09 ε48Ti and ±0.16 ε50Ti (2 s.d.)" (p.8). No accepted values or offsets are stated. The ±0.26 ε50Ti in the same paper is the LLNL procedure's (16 analyses of BCR-2 and BHVO-2, p.8), not this one
Detection Limit | Detection limit, one per reported concentration variable (one per target species, these being the same set). State the units and whether the values are procedure-typical estimates or session-specific measured values. The calculation method is recorded separately in Detection Limit Method. Record 'N/A' where the procedure reports no concentrations. | Advanced | Basic | Numeric + unit / Text |  | Source: Composition QC module | 2026-09-01 | reported property |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Detection Limit Method | Formula or approach used to calculate the detection limits, with a citation for the method where one exists. Must be consistent with the values reported in Detection Limit. | Basic | Read-Only | Controlled list / Text | 3σ blank ¦ 3σ background ¦ 3σ counting statistics ¦ 3× blank mean ¦ Poisson statistics ¦ N/A ¦ None | Source: Composition QC module | 2026-08-27 | reported property |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Limit of Quantification (LOQ) Method | Reference or description of the method used to calculate the limit of quantification (LOQ): the lowest concentration reliably measurable with acceptable precision and accuracy. Required when concentrations near the LOD are reported. | Advanced | Basic | Controlled list / Text | 10σ blank ¦ 10σ background ¦ 10× LOD ¦ 3.3× LOD ¦ S/N ≥ 10 ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-27 | reported property |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Goodness-of-Fit or Dispersion Statistic | The statistic reported to show whether scatter among the contributing individual results exceeds what analytical uncertainty alone predicts, together with its value. The procedure may still state an acceptance threshold, which belongs with the inclusion criteria. | N/A | Basic | Text (free) |  | Source: Aggregation module | 2026-09-15 | reported property | Answers whether a reported aggregate is defensible as a single population. The value cannot be known before the analysis. |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Additional Notes | Any procedure- or analysis-specific information not captured by a structured field anywhere in this TAPP — including anomalies, deviations from the registered procedure, instrument modifications, and supplementary context. Scope is the whole document, not Group 6: this is the last field of the TAPP and covers all six groups. Use sparingly; a structured field is preferred for anything that can be formally categorised. | Advanced | Advanced | Text (free) |  | Source: Core module | 2026-08-26 | (none) |  |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N

## Legends
Procedure-Level Tier | Description | Analysis-Level Tier | Description
Basic | Mandatory for procedure registration. Must be provided to register a valid procedure. | Read-Only | Directly imported from the registered procedure; cannot be changed by the analyst. Changing this value means running a different procedure. Also used for fields relevant only at procedure level — the value is inherited from the procedure record and shown read-only in the analysis form.
Advanced | Optional for procedure registration. Strongly recommended but not required. | Editable | Imported from the registered procedure but may be adjusted within procedure-defined bounds (e.g., daily tuning, minor software updates). The procedure registers the target or typical value; the analyst confirms or adjusts it. Cannot be left void if the procedure-level tier is Basic.
N/A | Not applicable at procedure level. This field captures analysis-level information only. | Basic | Mandatory user input at analysis time. Value comes from the analysis itself and cannot be pre-specified in the procedure.
|  | Advanced | Optional user input at analysis time. Recommended for complete and reproducible documentation.
Mode Column | Definition
Purpose
Y | This field applies to this analytical mode and should be reported.
N | This field does not apply to this analytical mode.
Data Type | What to enter
Controlled list | Choose one of the values listed in "Example / Allowed Content". The list is CLOSED — it is meant to cover every case. If your value genuinely is not there, that is a gap in the list worth reporting, not a reason to write your own.
Controlled list / Text | Choose a listed value AND qualify it in the same cell — the detail is expected, not optional. Typically what was corrected, which target species or phases it applied to, the method or equation used, or the source it came from. An answer the list cannot express is also valid here.
Numeric (unit) | A number in the unit named in brackets. Enter the number only — the unit is fixed by the field.
Numeric + unit | A number AND its unit, because the unit varies between procedures. Write both, e.g. '50 us' or '0.5 s'.
Text (free) | Free text. No controlled vocabulary applies.
N/A ¦ None | Offered by every controlled list. "N/A" means the field does not apply to this procedure; "None" means it applies but nothing was used. Prefer either over leaving the cell blank.
Keyed By (Rule 7) | What the field's value repeats over
(none) | Scalar — one value per procedure or analysis. The default.
acquisition pass | One value per pass over the sample with its own instrument settings.
defines: acquisition pass | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per pass over the sample with its own instrument settings.
defines: monitored property per target species | ENUMERATES the monitored property domain — the header of the child table, not a column in it — and does so once per target species. One value per chemical species determined, at whatever granularity the procedure determines it.
defines: preparation step | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per stage in sample preparation.
defines: reported property | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
defines: sample | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it.
defines: sample > sampling unit | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per subdivision of the physical sample that carries its own row — grain, spot, aliquot, phase, sub-volume.
defines: standard | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per reference material or reference database entry.
defines: target species | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per chemical species determined, at whatever granularity the procedure determines it.
monitored property
pair: reported property | One value per unordered PAIR. One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
preparation step | One value per stage in sample preparation.
reported property | One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
sample
sample > sampling unit x reported property | Cross-product, read as "for each sample, one value per reported property". One value per subdivision of the physical sample that carries its own row — grain, spot, aliquot, phase, sub-volume. + One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
standard x reported property | Cross-product, read as "for each standard, one value per reported property". One value per reference material or reference database entry. + One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
target species | One value per chemical species determined, at whatever granularity the procedure determines it.