## TAPP
Metadata Item | Description | Procedure-Level Tier | Analysis-Level Tier | Data Type | Example / Allowed Content | Comments | Last Update | Keyed By | Purpose | EDS Point Analysis | EDS Mapping | WDS Point Analysis | WDS Mapping | Literature Assessment | Ma+2015 ¦ Caltech GPS ¦ WDS Point Analysis (JEOL 8200) | Hu+2020 ¦ IGGCAS ¦ WDS Point Analysis (JEOL JXA-8100) | Liu+2016_UT ¦ Cameca SX100 ¦ WDS Mapping (U.Tennessee) | Liu+2016_Cal ¦ JEOL JXA-8200 ¦ WDS Point Analysis (Caltech GPS) | Ma+2017 ¦ JEOL 8200 ¦ WDS Point Analysis (Caltech GPS Analytical Facility) | Frank+2023 ¦ Cameca SX100 ¦ WDS Point Analysis (ARES JSC) | Broussard+2026 ¦ JEOL JXA-8200 ¦ WDS Mapping (WashU) | Seifert+2026 ¦ JEOL 8530 ¦ WDS Point Analysis (ARES JSC) | Pang+2016 ¦ JEOL JXA-8100 ¦ WDS Point Analysis (Nanjing U.) | McCoy+2025_SI ¦ JEOL 8530F+ ¦ WDS Point Analysis (Smithsonian) | McCoy+2025_UA ¦ Cameca SX-100 ¦ WDS Point Analysis (K-ALFAA U.Arizona) | Zega+2025 ¦ Cameca SX-100 Ultra ¦ WDS Point Analysis (K-ALFAA U.Arizona) | Barnes+2025 ¦ JEOL JXA-8230 ¦ WDS Point Analysis (CRPG Nancy) | Barnes+2025 ¦ Cameca SX100 ¦ WDS Point Analysis (NHM London) | Neuman+2025 ¦ WashU St. Louis ¦ WDS Mapping (JEOL JXA-8200)
1. Procedure Identification | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Procedure Name | A short descriptive name for this analytical procedure, including a version number. Should identify the instrument, the technique, and the scope of what is measured — target species, material, or feature. | Basic | Read-Only | Text (free) | e.g., 'EPMA-WDS Major Element Silicate Minerals' ¦ 'EPMA-WDS+EDS Minor Element Phosphates' | Source: Core module | 2026-09-01 | (none) |  | Y | Y | Y | Y |  | EPMA-WDS Major Element Silicates/Oxides, Tissint Mars Meteorite (Caltech GPS, JEOL 8200) | EPMA-WDS Major Element Silicates/Oxides, NWA 8657 Shergottite (IGGCAS, JEOL JXA-8100) | EPMA Major Element Silicates/Oxides+Mapping, Tissint (U. Tennessee, Cameca SX100) | EPMA-WDS Major Element Silicates/Oxides, Tissint (Caltech GPS, JEOL JXA-8200) | EPMA-WDS Major Element Silicates/Glasses, Zagami (Caltech GPS Analytical Facility, JEOL 8200) | EPMA Major/Minor Element Silicates+Oxides+Sulfides, CI Chondrite (ARES JSC, Cameca SX100) | EPMA-WDS Quantitative Mapping+Analysis, CI Chondrite Minerals (WashU, JEOL JXA-8200) | EPMA Major Element Apatite incl. Halogens, Bennu (ARES JSC, JEOL 8530 EMPA) | EPMA-WDS Major Element Silicates/Oxides, NWA 8003 Eucrite (Nanjing U., JEOL JXA-8100) | EPMA-WDS Carbonate+Silicate/Oxide Composition, Bennu (Smithsonian, JEOL 8530 F+ Hyperprobe) | EPMA-WDS Phosphate+Carbonate Composition, Bennu (K-ALFAA U.Arizona, Cameca SX-100) | EPMA Major Element Silicates/Sulfides/Oxides/Phosphates/Carbonates, Bennu (K-ALFAA, Cameca SX-100 Ultra) | EPMA Major Element Silicates/Oxides/Carbonates, Bennu Anhydrous Minerals (CRPG Nancy, JEOL JXA-8230) | EPMA Major/Minor Element Anhydrous Silicates, Bennu (NHM London, Cameca SX100) | EPMA-WDS quantitative compositional mapping, Apollo 17 core 73001 continuous thin sections (Washington University in St. Louis)
Technique | Top-level analytical technique identifier. | Basic | Read-Only | Controlled list | EPMA-WDS ¦ EPMA-EDS ¦ EPMA-WDS+EDS | Source: Core module | 2026-08-30 | (none) |  | Y | Y | Y | Y |  | EPMA-WDS | N | N | N | EPMA-WDS | N | EPMA-WDS | N | EPMA-WDS | N | EPMA-WDS | N | N | N | EPMA-WDS
Procedure Author | Person(s) or laboratory responsible for developing and registering this procedure. ORCID recommended for individuals. | Basic | Read-Only | Text (free) | e.g., 'Chi Ma, Caltech' ¦ 'ARES Geochemistry Lab, NASA JSC' | Source: Core module | 2026-07-24 | (none) |  | Y | Y | Y | Y |  | Chi Ma | Sen Hu | N | N | Chi Ma | N | N | Logan B. Seifert | N | T. J. McCoy | T. J. Zega | N | N | N | N (no individual operator named for the EPMA work)
Laboratory | Name of the laboratory or institution hosting the instrument. | Basic | Editable | Text (free) | e.g., 'Caltech GPS Division Analytical Facility' ¦ 'ARES, NASA Johnson Space Center' | Source: Core module | 2026-05-11 | (none) |  | Y | Y | Y | Y |  | Caltech GPS Division Analytical Facility | Institute of Geology and Geophysics, Chinese Academy of Sciences (IGGCAS), Beijing | Department of Earth and Planetary Sciences, University of Tennessee, Knoxville | Division of Geological and Planetary Sciences, Caltech | Division of Geological and Planetary Sciences Analytical Facility, Caltech | ARES, NASA Johnson Space Center | Washington University in St. Louis | ARES, NASA Johnson Space Center | State Key Laboratory for Mineral Deposits Research, Nanjing University | Smithsonian Institution, National Museum of Natural History | Kuiper-Arizona Laboratory for Astromaterials Analysis (K-ALFAA), University of Arizona | Kuiper-Arizona Laboratory for Astromaterials Analysis (K-ALFAA), University of Arizona | Centre de Recherches Pétrographiques et Géochimiques (CRPG), Nancy, France | Natural History Museum (NHM), London, UK | Washington University in St. Louis
Laboratory ID | Persistent identifier for the laboratory (e.g., ROR ID). | Advanced | Editable | URI / DOI | e.g., 'https://ror.org/05dxps055' (Caltech ROR) | Source: Core module | 2026-05-11 | (none) |  | Y | Y | Y | Y | [13 empty] | N | N | N
Procedure Start Date | First date this procedure configuration was used in production. | Basic | Read-Only | Date | YYYY-MM-DD | Source: Core module | 2026-07-24 | (none) |  | Y | Y | Y | Y | [13 empty] | N | N | N
Funding Source for Procedure Development | Grants and other funding sources that supported instrument acquisition, major upgrades, procedure development, and associated personnel time. Include grant numbers and funding agencies where applicable. | Advanced | Read-Only | Text (free) | e.g., 'NSF EAR-1234567' ¦ 'NASA 80NSSC20K0371' ¦ 'None' | Source: Core module | 2026-07-24 | (none) |  | Y | Y | Y | Y | [13 empty] | N | N | N
Procedure Reference(s) | DOI or URL for peer-reviewed publications or technical reports describing, validating, or benchmarking this procedure. | Advanced | Read-Only | URI / DOI | e.g., 'https://doi.org/10.1234/example' | Source: Core module | 2026-07-24 | (none) |  | Y | Y | Y | Y |  | Ma et al. 2015, Earth Planet. Sci. Lett. 422:194-205; doi:10.1016/j.epsl.2015.03.057 | Hu et al. 2020, Geochim. Cosmochim. Acta 278:185-198; doi:10.1016/j.gca.2019.06.012 | Liu et al. 2016, Meteorit. Planet. Sci.; doi:10.1111/maps.12726 | Liu et al. 2016, Meteorit. Planet. Sci.; doi:10.1111/maps.12726 | Ma et al. 2018, Meteorit. Planet. Sci. 53:50-61; doi:10.1111/maps.13000 (file dated 2017) | Frank et al. 2023, Meteorit. Planet. Sci. 58:1495-1511; doi:10.1111/maps.14083 | Broussard et al. 2026, Meteorit. Planet. Sci.; doi:10.1111/maps.70138 | Seifert et al. 2026, Meteorit. Planet. Sci.; doi:10.1111/maps.70167 | Pang et al. 2016, Sci. Rep. 6:26063; doi:10.1038/srep26063 | McCoy et al. 2025, Nature 637:320-325; doi:10.1038/s41586-024-08495-6 | McCoy et al. 2025, Nature 637:320-325; doi:10.1038/s41586-024-08495-6 | Zega et al. 2025, Nat. Geosci.; doi:10.1038/s41561-025-01741-0 | Barnes et al. 2025, Nat. Astron.; doi:10.1038/s41550-025-02631-6 | Barnes et al. 2025, Nat. Astron.; doi:10.1038/s41550-025-02631-6 | Neuman et al. 2025, J. Geophys. Res. Planets 130, e2024JE008556; doi:10.1029/2024JE008556 (section 2.6)
Procedure DOI | DOI or unique persistent identifier for the registered procedure used in this analysis, generated upon procedure submission to a data repository. If a DOI has been applied for but not yet minted, enter "pending". | N/A | Basic | URI / DOI | e.g., 'https://doi.org/10.xxxxx/procedure-id' ¦ 'pending' | Source: Core module | 2026-08-25 | (none) | Encourages formal procedure registration and ensures traceability. | Y | Y | Y | Y | [13 empty] | N | N | N
Session Identifier | Identifier for the analytical session this record describes — the laboratory's own run, sequence or batch identifier as generated by the instrument or acquisition software. Distinct from any persistent identifier a repository mints on submission. | N/A | Basic | Text (free) |  | Source: Core module | 2026-08-25 | (none) | The analysis record corresponds to one session, which may cover several samples, and this is the link back to the raw instrument files. | Y | Y | Y | Y | [15 empty] | N
Analyst | Name(s) of the analyst(s) who performed the analysis session. ORCID is recommended for persistent identification. | N/A | Basic | Text (free) | e.g., 'Sen Hu' | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y |  | Chi Ma | Sen Hu | N | N | Chi Ma | David R. Frank | N | Logan B. Seifert | N | N | N | N | N | N | N
Analysis Start Date | Date on which the analytical session began. For sessions spanning multiple days, use the date of the first session. Format: YYYY-MM-DD. | N/A | Basic | Date | YYYY-MM-DD | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y | [13 empty] | N | N | N
Analysis End Date | Date on which the analytical session ended. May equal Analysis Start Date for single-day sessions. Format: YYYY-MM-DD. | N/A | Basic | Date | YYYY-MM-DD | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y | [13 empty] | N | N | N
Funding Source for Analysis | Grants and other funding sources that supported the specific sample analyses and data collection reported here. Include grant numbers and funding agencies. Distinct from procedure development funding, which may have been awarded years earlier. | N/A | Basic | Text (free) | e.g., 'NSF EAR-2345678' ¦ 'NASA 80NSSC23K0415' ¦ 'None' | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y | [13 empty] | N | N | N
Coupled Technique(s) | Other analytical techniques applied to the same sample(s) whose results are intended to be interpreted together with data from this procedure. Document coupling with any technique whose results are functionally linked to this dataset — providing calibration inputs, complementary spatial context, or required companion measurements. Name the instrument, laboratory or purpose alongside each technique where known, and join multiple techniques with '; '. Enter "None" if no coupling is intended. | Advanced | Editable | Text (free) | e.g., 'EPMA (Brown University CAMECA SX-100)' ¦ 'FIB-SEM (Crossbeam 540) for 3D tomography; EDS for mineral analysis' ¦ 'NanoSIMS; STXM; Raman' ¦ 'None' ¦ 'N/A' | Source: Core module | 2026-08-31 | (none) |  | Y | Y | Y | Y |  | SEM (Carl Zeiss 1550VP FE-SEM, BSE imaging); EBSD (HKL system on ZEISS 1550VP); synchrotron XRD; micro-Raman | SEM-EDS (FEI Nova NanoSEM 450); Raman spectroscopy | SEM (BSE imaging); petrographic microscopy; LA-ICP-MS (Agilent 7500ce, Virginia Tech) | SEM (BSE imaging); petrographic microscopy; LA-ICP-MS (Agilent 7500ce, Virginia Tech) | SEM (Carl Zeiss 1550VP FE-SEM, BSE imaging); EBSD; synchrotron XRD; micro-Raman | Petrographic microscopy; SEM-BSE (JEOL 5900LV); EPMA X-ray mapping; Cameca ims1280 ion microprobe (O isotopes; 26Al-26Mg); FIB-TEM | Powder XRD (Rigaku MiniFlex 600); ICP-MS (Thermo Fisher iCAP Qc, WashU); K isotope MC-ICP-MS (Neptune Plus, WashU); CO2 laser-fluorination O isotope MS (U. New Mexico); AMS (PRIME Lab, Purdue) | SEM-EDS; SIMS (Cameca ims 1280); TEM-EDS | SEM (BSE imaging); petrographic microscopy | SEM-EDS (NHM London; Smithsonian; JSC); TEM-EDS/EELS; FIB-SEM; ToF-SIMS; XRD; XANES | SEM-EDS (NHM London; Smithsonian; JSC); TEM-EDS/EELS; FIB-SEM; ToF-SIMS; XRD; XANES | SEM-EDS; TEM-EDS/EELS; FIB-SEM; XRD; XANES | SEM-BSE (JEOL JSM-6510, 15 kV, 3 nA); SEM-EDS (multi-element mapping); SIMS (CAMECA IMS 1270 E7, CRPG); NanoSIMS (K-ALFAA); ICP-MS; MC-ICP-MS; noble gas MS | SEM-EDS (Zeiss EVO 15LS + Oxford X-Max80, 20 kV, 3 nA); NanoSIMS (OU); SIMS (CAMECA ims-1280-HR, Hokkaido); laser fluorination O isotopes (OU) | BSE mosaic imaging (same JEOL JXA-8200); QEMSCAN (FEI QUANTA 650 FEG-SEM, Univ. Manchester); optical microscopy (Keyence VHX 7000, NASA JSC); micro-XCT (custom NSI instrument, UTCT)
Coupling Description | Description of how this procedure is coupled with the technique(s) listed above. Include: (1) the functional relationship — what data or context flows between techniques, or how results are combined (e.g. which output from the coupled technique serves as input to data reduction for this technique); and (2) the analytical sequence — which technique is performed first and why (e.g. non-destructive before destructive). Required when Coupled Technique(s) is not "None". | Advanced | Editable | Text (free) | e.g., 'EPMA WDS major element concentrations used as internal standard for LA-ICP-MS trace element data reduction. EPMA performed first because LA-ICP-MS ablation is destructive.' | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | Single-element and RGB composite X-ray maps are compared with BSE and optical image mosaics (reflected light, plane polarized, crossed polars) to discriminate crystalline from glassy phases; an Al-Mg-Fe RGB composite X-ray map is used to discriminate feldspathic (red) from ferromagnesian (green and blue) phases
Coupled Procedure DOI | Registered procedure DOI for the coupled technique named above. If the coupled procedure has not yet been registered, enter the DOI of a publication describing the coupled method, or "pending". Enter "None" if no coupling is planned. | N/A | Advanced | URI / DOI | e.g., 'https://doi.org/10.xxxxx/coupled-procedure' ¦ 'pending' ¦ 'None' | Source: Core module | 2026-08-25 | (none) | Provides a stable, citable link to the companion method independent of whether a dataset has been deposited. | Y | Y | Y | Y | [13 empty] | N | N | N
Coupled Dataset or Publication Reference | DOI or other persistent identifier for the co-registered dataset or publication where both datasets are reported together. Accepts: a dedicated dataset DOI (if separately deposited), a shared dataset DOI (if co-submitted in the same package), or a publication DOI. Use "same submission" if the coupled dataset is included in this data package, or "pending" if not yet assigned. If coupling is documented through a shared sample identifier only, that information is already captured in Sample Persistent Identifier (Group 2). | N/A | Advanced | URI / DOI / Text | e.g., 'https://doi.org/10.xxxxx/dataset' ¦ 'same submission' ¦ 'pending' ¦ 'None' | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y | [13 empty] | N | N | N
2. Samples | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Target Material | General description of the material type(s) this procedure is designed to analyse. | Basic | Read-Only | Controlled list / Text | Silicate mineral ¦ Silicate glass ¦ Oxide ¦ Sulfide ¦ Carbonate ¦ Phosphate ¦ Metal or alloy ¦ N/A ¦ None | Source: Core module | 2026-08-31 | (none) | Used for discoverability and procedure matching, and because the material type constrains sample preparation, calibration and matrix-matching requirements. | Y | Y | Y | Y |  | Silicate mineral (tissintite clinopyroxene, plagioclase, maskelynite) ¦ Oxide ¦ Glass (melt pocket) | Silicate mineral (coesite, pyroxene, feldspar) ¦ Oxide ¦ Sulfide | Silicate mineral (olivine, pyroxene, maskelynite) ¦ Oxide (chromite, ulvospinel, ilmenite) ¦ Sulfide ¦ Phosphate (merrillite) ¦ Glass (melt pocket) | Silicate mineral (olivine, pyroxene, maskelynite) ¦ Oxide ¦ Sulfide ¦ Phosphate ¦ Glass | Silicate mineral (liebermannite, lingunite, maskelynite, augite, pigeonite) | Silicate mineral ¦ Oxide ¦ Sulfide ¦ Phosphate (CI chondrite phases) | Phyllosilicate (matrix) ¦ Oxide (magnetite, ilmenite) ¦ Sulfide (pyrrhotite, pentlandite) ¦ Carbonate (dolomite, magnesite) ¦ Phosphate (Ca phosphate, Na-Mg hydrous phosphate) | Phosphate (apatite) | Silicate mineral ¦ Oxide (eucrite phases) | Carbonate ¦ Oxide ¦ Silicate mineral (Bennu evaporite and host phases) | Phosphate (Mg,Na phosphate) ¦ Carbonate (Bennu evaporite phases) | Silicate mineral ¦ Sulfide ¦ Oxide ¦ Phosphate ¦ Carbonate (Bennu samples) | Silicate mineral (olivine, pyroxene) ¦ Oxide ¦ Carbonate (anhydrous minerals in Bennu aggregate particles) | Silicate mineral (olivine, pyroxene) in Bennu aggregate particles | Lunar regolith (Apollo 17 double drive tube, lower section 73001), as continuous thin sections
Sample Preparation Method | The form in which the sample is presented to the instrument, and the preparation that brought it to that form — for example mounting, sectioning, polishing, coating, crushing, fusion, or extraction of an electron-transparent section. Record 'None' where the material is analysed as received. | Basic | Editable | Controlled list / Text | Polished thin section ¦ Polished thick section ¦ Polished block or epoxy mount ¦ Grain mount ¦ Fragments or particles on conductive tape ¦ Bulk specimen, polished face ¦ Acid-dissolution residue ¦ N/A ¦ None | Source: Core module | 2026-08-27 | (none) |  | Y | Y | Y | Y |  | Polished thin section; carbon coating N | Polished thick section (NWA 8657); carbon coating N | Polished thin sections (coating type N) | Polished thin sections (coating type N) | Polished thin section USNM 7619 (coating type N) | N | Fragments mounted and dry-polished in a petrographic thin section; carbon coating N | Fragments embedded in epoxy; dry-polished with diamond powder; one mount ion-polished before carbon coating | Polished thin section; carbon coating N | Ir-coated specimens mounted on Ir-coated Parafilm | Polished section; 20 nm carbon coat | N | Aggregate particles (<1 mm) mounted in epoxy at Université Côte d'Azur; polished; carbon coated (thickness N) | Mounted in resin blocks; polished at NHM London; fragmented during polishing (P1, P2); initial carbon coat for SEM/EPMA (thickness N); additional coat added after for SIMS (total ~30 nm) [N5,N6] | Core extruded and dissected in 0.5 cm depth intervals; the remaining material impregnated with epoxy to create a continuous thin section set of the entire core; sections are 50 x 25 mm. Carbon coating N
Sample Name | Name or identifier of each sample analysed in this session, as used in the laboratory — a sample mount, section or aliquot counts as one entry where that is what the laboratory tracks. Should match the identifier used in associated publications or data tables. | N/A | Basic | Text (free) | e.g., 'NWA 8657 polished thin section 1' ¦ 'Tissint UT1' ¦ 'Ivuna MZ2' | Source: Core module | 2026-08-25 | defines: sample | The analysis record corresponds to one session and may cover several samples. | Y | Y | Y | Y |  | Tissint Mars meteorite | NWA 8657 shergottite | Tissint thin sections UT1, UT2, UT3; Tata-1-C1 to C3; Tata-2-C1 to C3; Tata-3-C1 to C3; Tissint-B | Tissint thin sections UT1, UT2, UT3; Tata-1-C1 to C3; Tata-2-C1 to C3; Tata-3-C1 to C3; Tissint-B | Zagami USNM 7619 | Ivuna CI chondrite, section MZ2 | OC002 LAB24-2 (10-11 fragments of Oued Chebeika 002) | OREX-803079-0 and OREX-803080-0 (Bennu OSIRIS-REx samples) | NWA 8003 eucrite | Bennu OSIRIS-REx samples (OREX-8#####-###) | Bennu OSIRIS-REx samples (OREX-8#####-###) | Bennu OSIRIS-REx samples (OREX-5/8#####-###) | OREX-800045-103 and OREX-800045-107 | OREX-501054-0 and OREX-501059-0 (particles P1, P2) | 73001,6014-73001,6021
Sampling Unit | The physical subdivision of the sample to which one row of reported values corresponds — the unit that is analysed and reported, as distinct from the sample as a whole. Where units nest (e.g. confined tracks within grains), state both levels. | Basic | Basic | Controlled list / Text | Whole sample ¦ Aliquot ¦ Grain ¦ Spot ¦ Analysis point ¦ Phase ¦ Sub-volume ¦ Region of interest ¦ N/A ¦ None | Source: Core module | 2026-08-30 | defines: sampling unit |  | Y | Y | Y | Y | [15 empty] | Analysis point - each map pixel is a fully quantitative analysis (1,024 x 1,024 per stage map; five stage maps per thin section; 20 x 10^6 analyses across all slides)
Sample Persistent Identifier | Globally unique, persistent identifier for each sample listed in Sample Name. IGSN (International Geo Sample Number) is the recommended standard for geological and cosmochemical samples, as used by Astromat, EarthChem and SESAR. Where a sample and its sub-samples are separately registered, record the identifier at the level actually analysed. | Advanced | Advanced | URI / IGSN | e.g., 'IGSN:AU1234567' | Source: Core module | 2026-08-25 | sample |  | Y | Y | Y | Y | [10 empty] | OREX-8#####-### (NASA curatorial ID, not IGSN) | OREX-8#####-### (NASA curatorial ID, not IGSN) | OREX-5/8#####-### (NASA curatorial ID, not IGSN) | OREX-800045-103; OREX-800045-107 (NASA curatorial IDs, not IGSN) | OREX-501054-0; OREX-501059-0 (NASA curatorial IDs, not IGSN) | N (NASA curation sample numbers given as names; not stated as persistent identifiers)
Sampling Unit Selection Criteria | The rules governing which sampling unit(s) within a sample are selected for analysis, and why. Covers the criteria applied when choosing grains, aliquots, spots, or a region of interest — size, morphology, clarity, freedom from inclusions or alteration, phase identity, or spatial position. Distinct from Target Material, which states the material type the procedure is designed for, and from Sampling Unit, which names the kind of subdivision one row of reported values corresponds to: this field states how, within such a sample, the unit actually analysed is picked out. | Basic | Editable | Text (free) |  | Source: Sampling Unit Selection module | 2026-09-01 | (none) |  | Y | Y | Y | Y | [15 empty] | N - continuous thin sections of the whole core; five stage maps per section, no selection rule stated
Pre-Analysis Imaging and Screening | Imaging or other characterisation performed before the measurement in order to select or locate the sampling unit to be analysed, including the technique, instrument and settings used, and how individual analyses are linked back to the images. Distinct from any imaging the procedure performs as its own measurement. Where the imaging is performed on a separate instrument, it should also be recorded in the Group 1 coupling fields. | Advanced | Editable | Text (free) |  | Source: Sampling Unit Selection module | 2026-09-01 | sample |  | Y | Y | Y | Y | [15 empty] | BSE mosaic: approximately 325 backscattered-electron images collected with the JEOL guide-net mapping software at 15 kV, 2 nA probe current and 70x magnification, stitched with the ImageJ Fiji grid-collection stitching plug-in (Donovan et al., 2021) into a 20k x 5k pixel mosaic at ~1.5 um pixel resolution
3. Instrument & Software | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Instrument Manufacturer | Manufacturer of the instrument that performs the measurement, recorded as a controlled value. Where a procedure couples a sample-introduction system to an analysing instrument, this records the analysing instrument. Instrument Model gives the specific designation. | Basic | Read-Only | Controlled list | JEOL ¦ Cameca ¦ Unknown ¦ N/A ¦ None | Source: Core module | 2026-08-30 | (none) | Recording it as a controlled value lets procedures be found by vendor. | Y | Y | Y | Y |  | JEOL | JEOL | Cameca | JEOL | JEOL | Cameca | JEOL | JEOL | JEOL | JEOL | Cameca | Cameca | JEOL | Cameca | JEOL
Instrument Model | Model designation of the instrument that performs the measurement, including any generation or configuration suffix. Conventionally written with the manufacturer name included; Instrument Manufacturer records the vendor separately, as a controlled value, so that procedures remain findable by vendor. | Basic | Read-Only | Text (free) | e.g., JXA-8530F Plus ¦ JXA-8230 ¦ JXA-8200 ¦ JXA-8100 ¦ SXFiveTactis ¦ SXFive ¦ SX100 ¦ iHP200F ¦ Unknown ¦ N/A ¦ None | Source: Core module | 2026-08-31 | (none) |  | Y | Y | Y | Y |  | JEOL 8200 (stated as "JEOL 8200 electron microprobe"; no JXA prefix stated) | JXA-8100 (stated as "JEOL JXA-8100") | SX100 (stated as "Cameca SX100") | JXA-8200 (stated as "JEOL JXA-8200") | JEOL 8200 (stated as "JEOL 8200 electron microprobe"; no JXA prefix in paper) | SX100 (stated as "Cameca SX100") | JXA-8200 (stated as "JEOL JXA-8200 electron microprobe") | JEOL 8530 EMPA (stated as "JEOL 8530 EMPA at NASA JSC"; no F suffix or JXA prefix stated) | JXA-8100 (stated as "JEOL 8100") | JXA-8530F Plus (stated as "JEOL 8530 F+ Hyperprobe Field Emission Electron Probe Microanalyzer") | SX-100 (stated as "Cameca SX-100 electron microprobe located at K-ALFAA") | SX-100 Ultra (stated as "SX-100 Ultra electron microprobe in the K-ALFAA") | JXA-8230 (stated as "JEOL JXA-8230 electron microprobe analyser (EPMA)") | SX100 (stated as "CAMECA SX100 electron microprobe") | JEOL JXA-8200
Electron Source | Type of electron gun used in the instrument. | Basic | Read-Only | Controlled list | Cold-FEG ¦ Schottky FEG (X-FEG) ¦ Schottky FEG (standard) ¦ Field emission gun (FEG) — subtype not specified ¦ LaB6 / CeB6 ¦ Tungsten (W) ¦ Unknown ¦ N/A ¦ None |  | 2026-08-30 | (none) |  | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | Field Emission (FEG) | N | N | N | N | N
Acquisition Software | Instrument control, automation and data acquisition software used to collect the raw data, including version number. Distinct from Data Processing Software(s), which covers everything applied to the data after acquisition. | Basic | Editable | Text (free) | e.g., 'Probe for EPMA v12.9.5' ¦ 'JEOL PC-EPMA v3.0' ¦ 'Cameca PeakSight 6.2' | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y |  | Probe for EPMA (Probe Software, Inc.) | N | N | N | Probe for EPMA (Probe Software, Inc.) | N | Probe for EPMA microanalysis software | N | N | N | N | N | N | N | JEOL guide-net mapping software (BSE mosaic); N for the WDS stage maps
Data Processing Software(s) | All software applied to the data after acquisition in order to produce the reported quantities, including version numbers. List every package used. Distinct from Acquisition Software, which controls the instrument and collects the raw data. | Basic | Editable | Text (free) | e.g., 'Probe for EPMA (built-in)' ¦ 'CalcZAF v12' ¦ 'CITZAF (Armstrong 1995)' ¦ 'XMapTools 4.1 (for mapping)' | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y |  | CITZAF correction procedure (Armstrong 1995) | Bence-Albee method | N | N | CITZAF correction procedure (Armstrong 1995) | N | Probe for EPMA (CITZAF matrix correction, Armstrong 1995); CalcImage and Quantitative Microanalysis Explorer web-based tool (for stage mapping) | N | N | N | N | XMapTools (for phase maps) | N | N | Probe Software CalcImage and Probe for EPMA (full Phi(rho-z) correction at each pixel); MATLAB routines generating 32-bit floating point .tiff quantitative maps; Fiji and MATLAB for stitching; ENVI input image stacks
WDS Spectrometer Configuration | Number, type, and crystal range of WDS spectrometers on the instrument. Include manufacturer, model, and crystal range. For SEM-WDS configurations (third-party WDS on a non-EPMA platform), include WDS manufacturer and model. | Advanced | Read-Only | Text (free) | e.g., '5 Rowland-circle WDS spectrometers (JEOL), crystal range LDE1-LIFH' ¦ '4 WDS spectrometers (Cameca)' |  | 2026-08-12 | (none) |  | N | N | Y | Y |  | N | N | N | N | N | N | 5 wavelength-dispersive spectrometers (JEOL) | N | N | LIFL (Fe Ka, Mn Ka); TAPL (Mg Ka); PETL (Ca Ka) — partial; full config N | N | N | 5 wavelength-dispersive spectrometers (JEOL) | N | Fixed wavelength-dispersive spectrometers; count not stated (five elements collected per pass)
EDS Detector Configuration | EDS detector type, manufacturer, number of detector elements, active area and solid angle, window type, and geometry (take-off angle, position). List multiple detectors separately. Record 'N/A' where the procedure has no EDS detector. | Advanced | Read-Only | Text (free) | e.g., 'Oxford Ultim Max 100 SDD, 100 mm2, polymer window, 35 deg take-off angle' ¦ 'JEOL (e2v/Gresham) SDD, Be window' |  | 2026-08-27 | (none) |  | Y | Y | N | N |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A - WDS mapping; no EDS used for the EPMA work
4. Measurement Information | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Analytical Mode | Primary analytical mode(s) executed under this procedure. Records one value for a single-mode procedure and all applicable modes for a multi-mode one. Applies even where a procedure defines no mode flag columns. | Basic | Read-Only | Controlled list | EDS Point Analysis ¦ EDS Mapping ¦ WDS Point Analysis ¦ WDS Mapping | Source: Core module | 2026-09-09 | (none) | Serves as the procedure-level declaration of measurement type, distinct from the mode flag columns, which state per-field applicability. Ensures every procedure record is self-describing and comparable across the library. | Y | Y | Y | Y | [15 empty] | WDS Mapping
Beam Mode | Whether the electron beam was operated as a stationary focused spot, defocused to a specified diameter, or rastered over a small area during a single-point analysis. Must be consistent with the Beam Diameter and Beam Raster Dimensions fields. For mapping, beam scanning is controlled by Step Size / Pixel Size and Stage Scan vs. Beam Scan instead. | Basic | Read-Only | Controlled list / Text | Focused ¦ Defocused ¦ Rastered ¦ N/A ¦ None |  | 2026-08-30 | sample > sampling unit |  | Y | Y | Y | Y |  | Focused (stated: "beam in focused mode") | Focused | Focused (olivine, pyroxene, Fe-Ti-Cr oxides); Defocused 5-10 µm (maskelynite, phosphate, sulfide, glass) | Focused (olivine, pyroxene, Fe-Ti-Cr oxides); Defocused 5-10 µm (maskelynite, phosphate, sulfide, glass) | Focused (stated: "beam in focused mode") | Focused (point analysis); mapping beam mode N | N | Focused (2 µm probe size stated) | Focused (most phases); Defocused 2-5 µm (plagioclase and polymorphs) | Focused (1 µm, silicates/oxides); Focused (5 µm, carbonates) | Focused (1 µm) | Focused (silicates, sulfides, oxides); Defocused 2 µm (phosphates, carbonates) | Focused (1 µm, point analysis); Rastered 5×5 µm² for carbonates | Focused (1 µm) | Fixed 10 um beam (stated 'a fixed 10 um electron beam')
Accelerating Voltage | Electron beam accelerating voltage in kilovolts (kV). Justify any deviation from the standard operating voltage. | Basic | Editable | Numeric (kV) | e.g., 15 ¦ 20 ¦ 10 ¦ 7.5 |  | 2026-08-27 | (none) |  | Y | Y | Y | Y |  | 15 kV | 15 kV | 15 kV | 15 kV | 15 kV | 20 kV | 15 kV | 15 kV | 15 kV | 15 kV | 15 kV | 15 kV | 20 kV | 20 kV | 15 kV (stage maps and BSE mosaic)
Beam Current | Probe current in nanoamperes (nA). Often varies by phase type or target species; record the procedure-standard value(s). | Basic | Editable | Numeric (nA) | e.g., 5 ¦ 10 ¦ 20 ¦ 50 ¦ 100 |  | 2026-09-01 | sample > sampling unit |  | Y | Y | Y | Y |  | 5 nA | 10 nA | 20 nA (olivine, pyroxene, Fe-Ti-Cr oxides); 10 nA (maskelynite, phosphate, sulfide, glass); 200 nA (olivine megacryst Ka maps) | 20 nA (olivine, pyroxene, Fe-Ti-Cr oxides); 10 nA (maskelynite, phosphate, sulfide, glass) | 5 nA | 20 nA | 25 nA (point analysis) | 20 nA | 20 nA (focused, most phases); 20 nA defocused (plagioclase and polymorphs) | 10 nA (carbonates and silicates/oxides) | 8 nA (phosphates); carbonate conditions N | 20 nA (silicates, sulfides, oxides); 8 nA (phosphates); 4 nA (carbonates) | 10 nA | N | 100 nA probe current (stage maps); 2 nA (BSE mosaic)
Beam Diameter | Diameter of the electron beam in micrometers. 0 indicates a fully focused beam. Document defocused diameter when used to minimize beam damage or improve spatial averaging for beam-sensitive phases. | Basic | Editable | Numeric (µm) / Text | e.g., 0 (focused) ¦ 1 ¦ 2 ¦ 5 ¦ 10 ¦ 20 |  | 2026-08-30 | sample > sampling unit |  | Y | Y | Y | Y |  | Focused (exact diameter N) | Focused (exact diameter N) | 1-2 µm (olivine, pyroxene, Fe-Ti-Cr oxides); 5-10 µm defocused (maskelynite, phosphate, sulfide, glass) | 1-2 µm (olivine, pyroxene, Fe-Ti-Cr oxides); 5-10 µm defocused (maskelynite, phosphate, sulfide, glass) | Focused (exact diameter N) | 1 µm (focused) | N | 2 µm (stated as "2 µm probe size") | Focused (exact diameter N); 2-5 µm defocused (plagioclase and polymorphs) | 5 µm (carbonates); 1 µm (silicates/oxides) | 1 µm | Focused (silicates, sulfides, oxides); 2 µm defocused (phosphates, carbonates) | 1 µm (point analysis); 5×5 µm² raster area for carbonates | 1 µm (focused) | 10 um (fixed)
Beam Raster Dimensions | Dimensions of the small area over which the beam is rastered at a single analysis point, reported as width × height in µm. Applicable when Beam Mode = Rastered; defines the effective spatial footprint of the measurement. Not applicable when mapping. | Advanced | Editable | Numeric pair (µm x µm) | e.g., 5x5 ¦ 10x10 ¦ None |  | 2026-08-30 | sample > sampling unit | Averages over a coarse-grained or beam-sensitive phase by distributing dose across a larger area. | Y | N | Y | N |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 5×5 µm² for carbonates | N/A | N/A - stage scan, not beam scan
Beam Damage Minimization | Measures taken to minimize beam damage, particularly volatilization or migration of Na, K, F, and Cl in hydrous minerals, glasses, feldspars, phosphates, and carbonates. Document approach, beam conditions used, and phases for which it was applied. | Advanced | Editable | Text (free) | e.g., 'Na measured first with 10 um defocused beam at 5 nA' ¦ 'Time-resolved acquisition extrapolated to t=0 for Na and K' ¦ 'None required (anhydrous minerals)' |  | 2026-08-12 | sample > sampling unit |  | Y | Y | Y | Y |  | N | N | Defocused beam 5-10 µm for maskelynite, phosphate, sulfide, and glass | Defocused beam 5-10 µm for maskelynite, phosphate, sulfide, and glass | Low beam current (5 nA); Na diffusion away from beam still observed in liebermannite | N | N | 2 µm probe used for all analyses; Durango apatite tested at 10 µm and 3 µm spot sizes to assess halogen volatilization; no significant loss found under adopted conditions | N | N | N | Defocused 2 µm beam for phosphates (8 nA) and carbonates (4 nA) | N | N | N
Drift Correction | Method used to monitor and correct for instrument drift (beam current drift, spectrometer drift) during the analytical session. | Advanced | Editable | Text (free) | e.g., 'Primary standards re-measured at start, midpoint, and end of session; calibration linearly interpolated' ¦ 'Faraday cup beam current verified every 30 min' ¦ 'None' |  | 2026-05-11 | (none) |  | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Target Species | The chemical species this procedure is designed to determine, recorded at whatever resolution the chemistry is resolved — element(s) for this technique; valence species where a procedure resolves oxidation state; compounds where it resolves molecules. Isotopes are not target species: isotopes of an element are the same chemical species. The subset actually measured may be narrower than the registered suite where species were excluded due to interferences or scope reduction. | Basic | Editable | Text (free) | e.g., 'Si, Ti, Al, Fe, Mn, Mg, Ca, Na, K, Cr, Ni, P, S, Cl, F' ¦ 'Si, Al, Fe, Mg, Ca, Na, K' | Source: Target Species module | 2026-09-01 | defines: target species |  | Y | Y | Y | Y |  | Si, Ti, Al, Cr, Fe, Mn, Mg, Ca, Na, K (from standards list; P not in standard suite) | Si, Ti, Al, Cr, Fe, Mn, Mg, Ca, Na, K | Si, Ti, Al, Mg, Ca, Fe, Mn, Cr, Ni, Na, K, P (from detection limits table) | Si, Ti, Al, Mg, Ca, Fe, Mn, Cr, Ni, Na, K, P (from detection limits table; conditions shared with UT) | Si, Al, K, Na, Ca, Fe, Mg, Ti, Cr, Mn (stated via standards sentence and detection limit list) | Si, Al, Ti, K, Na, Fe, Mg, Ca, S, Mn, Cr, Ni, P, V | F (explicitly stated); full analyte list N in methods | P, F, Cl, Ca, Mn, Fe, Na, Mg, Si, S | Si, Ti, Al, Cr, Fe, Mn, Mg, Ca, Na, K, P (from standards/methods text) | Fe, Mn, Mg, Ca (carbonates); Fe, Mn, Mg, Ca, Ni, Cr, Al, Si, Ti, K, Na (silicates/oxides) | Na, Si, Mg, Ca, Mn, P, S, Fe (carbonates); F, P, Ca, Si, Mg, Fe, Al, S, K, Cl (phosphates) | N | Session 1: Al, Ti, Ca, Cr, Mn, Ni, Mg, Fe, Si; Session 2: Na, K, Al, Ti, Ca, Cr, Mn, Ni, Mg, Fe, Si | N | Mg, Al, Fe, Ca, Ti, Na, Si, Mn, K, Cr (ten elements, collected in two passes)
Reported Variables and Units | The final variable(s) this procedure reports and their units — distinct from the fields recording what was *acquired* rather than what is reported. Record every reported variable, including intermediate quantities reported alongside final ones (e.g. both the 206Pb/238U ratio and the 206Pb/238U date). Where a reported variable is a nominal property with no magnitude (e.g. a mineral identification), record the variable and give the unit as 'N/A — nominal property'. | Basic | Read-Only | Text (free) | e.g., 'Element concentration (ppm); oxide (wt%)' ¦ 'd56Fe (permil vs IRMM-014); d57Fe (permil vs IRMM-014)' ¦ '206Pb/238U date (Ma); 207Pb/206Pb date (Ma)' ¦ 'Mineral identity (N/A — nominal property); match score (0-1)' | Source: Core module | 2026-08-25 | defines: reported property | A procedure may acquire many channels and report a small number of derived quantities; without this field a data consumer cannot tell which. | Y | Y | Y | Y | [15 empty] | Quantitative element and oxide wt.% maps; cation stoichiometry; derived mineral endmember maps (32-bit floating point .tiff)
EPMA Technique per Target Species | Whether the measurement was made by WDS or EDS. Applies where a procedure uses both WDS and EDS. | Basic | Read-Only | Controlled list | WDS ¦ EDS ¦ N/A ¦ None |  | 2026-09-01 | target species |  | Y | Y | Y | Y |  | WDS (all analytes) | N | N | N | WDS (all analytes) | N | WDS (all analytes) | N | WDS (all analytes) | N | WDS (stated for phosphate analyses) | N | N | N | WDS (all ten elements)
X-ray Line | X-ray emission line measured. | Basic | Read-Only | Controlled list | Ka ¦ Kb ¦ La ¦ Lb ¦ Lg ¦ Ma ¦ Mb ¦ N/A ¦ None |  | 2026-09-10 | monitored property | Line choice affects sensitivity, matrix correction accuracy, and susceptibility to peak overlap and spectral interference. | Y | Y | Y | Y |  | SiKα, AlKα, CaKα (anorthite); NaKα (albite); FeKα (fayalite); MgKα (forsterite); MnKα (Mn2SiO4); TiKα (TiO2); CrKα (Cr2O3); KKα (microcline) | N | N | N | SiKa, AlKa, KKa (microcline); CaKa (anorthite); NaKa (albite); FeKa (fayalite); MgKa (forsterite); TiKa (TiO2); CrKa (Cr2O3); MnKa (Mn-olivine) | N | N | N | N | N | N | N | N | N | N
Diffracting Crystal | Analyzing crystal (monochromator). | Basic | Read-Only | Controlled list | LIF ¦ LLIF ¦ LIFL ¦ LIFH ¦ PET ¦ LPET ¦ PETL ¦ PETH ¦ PETJ ¦ TAP ¦ TAPL ¦ TAPH ¦ LTAP ¦ LDE1 ¦ LDE2 ¦ LDE3 ¦ RAP ¦ KAP ¦ ADP ¦ Unknown ¦ N/A ¦ None |  | 2026-09-10 | monitored property | Crystal choice determines the detectable wavelength range and dispersion. | N | N | Y | Y |  | N | N | N | N | N | N | LDE1 (for F); other analytes N | N | N | LIFL (Fe, Mn); TAPL (Mg); PETL (Ca) — partial; other analytes N | N | N | N | N | N
WDS Spectrometer Channel | WDS spectrometer position(s) assigned to each target species, one entry per assignment. An target species may be assigned to more than one spectrometer with intensities aggregated (aggregate intensity counting), and one spectrometer serves several target species across a run, so the assignment — not the target species — is the unit carrying the spectrometer setup. | Advanced | Read-Only | Text (free) | e.g., 'Si=Sp1; Ti=Sp2; Cr=Sp2+Sp3 (aggregate intensity counting); Fe=Sp3' ¦ 'Mg=Sp1; Ca=Sp4; Na=Sp5' |  | 2026-09-10 | defines: monitored property per target species | Different spectrometers may have different crystal configurations. | N | N | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N - spectrometer-to-element assignments not stated
Sequence | Order in which spectrometer assignments are acquired, and — where the element suite exceeds the number of spectrometers — the passes the acquisition is divided into. Within a single pass all assigned spectrometers collect simultaneously, including at every pixel in X-ray mapping; a suite larger than the spectrometer count therefore requires the acquisition to be run more than once, each pass covering a different subset of elements. | Advanced | Read-Only | Integer | e.g., 1 ¦ 2 ¦ 3 ¦ 4 |  | 2026-09-10 | monitored property | Relevant for minimizing beam damage (volatile elements measured first) and for sequential multi-channel setups. | N | N | Y | N |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | Two passes per stage map: pass 1 = Mg, Al, Fe, Ca, Ti; pass 2 = Na, Si, Mn, K, Cr
Proportional Counter / Detector | Type of detector used. | Advanced | Read-Only | Controlled list | High Pressure P10 Flow ¦ Low Pressure P10 Flow ¦ Xe Sealed ¦ SDD ¦ Si(Li) ¦ Unknown ¦ N/A ¦ None |  | 2026-09-10 | monitored property | Affects sensitivity and count rate linearity. | N | N | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
WDS PHA Setting | Pulse height analyzer (PHA) setting for the WDS detector. Integral mode accepts all pulses above a threshold; Differential mode selects a narrow energy window to reject higher-order reflections and escape peaks. | Advanced | Read-Only | Controlled list | Integral ¦ Differential ¦ Unknown ¦ N/A ¦ None |  | 2026-09-10 | monitored property |  | N | N | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Peak Counting Time | Time spent counting X-ray intensity at the peak position, in seconds. Adjustments stay within procedure-defined bounds. | Basic | Editable | Numeric (s) | e.g., 10 ¦ 20 ¦ 30 ¦ 60 |  | 2026-09-10 | monitored property |  | N | N | Y | N |  | N | N | N | 20 s on-peak | N | 10-50 s | N | N | N | N | N | 20 s (silicates, sulfides, oxides; phosphates); 10 s (carbonates) | 200 ms total (peak + background) for Al, Ti, Ca, Mn, Cr; 20 ms total (peak + background) for Mg, Fe, Si | N | 25 msec dwell per pixel, all on-peak - the MAN background correction 'allows all map collection time to be dedicated to on-peak X-ray measurement'
Background Counting Time | Total time spent counting at off-peak background position(s) in seconds, summed across all background positions. | Basic | Editable | Numeric (s) | e.g., 5 ¦ 10 ¦ 20 |  | 2026-09-10 | monitored property |  | N | N | Y | N |  | N | N | N | 10 s each on high and low background positions (20 s total) | N | N | N/A | N | N | N | N | 10 s each on high and low positions (silicates, sulfides, oxides; phosphates); 5 s on each position (carbonates) | N/A | N | N/A - MAN (mean atomic number) background calibration; no off-peak counting
Background Position(s) | Location(s) of off-peak background measurement(s) relative to the peak, in mm or sin-theta, and whether on the high- or low-energy side. | Advanced | Editable | Text (free) | e.g., '+5 mm (High), -5 mm (Low)' ¦ 'High only (+4 mm)' ¦ 'MAN method (no off-peak positions)' |  | 2026-09-10 | monitored property |  | N | N | Y | N |  | N | N | N | High and low off-peak positions (2-point; offset in mm or sin-theta N) | N | N | N/A | N | N | N | N | High and low off-peak positions (2-point; positions in mm or sin-theta N) | N | N | N/A - MAN background calibration; no off-peak positions
EDS Live Time per Point or Pixel | EDS spectral acquisition live time per analysis point in seconds. Previously referred to as "EDS Acquisition Time" in this TAPP and commonly used under that name in EPMA and SEM-EDS contexts. Renamed to align with TEM-EDS usage, where the per-point vs. per-pixel distinction (point/line mode vs. spectrum image) is explicit. In EPMA, acquisition is always per point. | Basic | Editable | Numeric (s) | e.g., 30 ¦ 60 ¦ 120 |  | 2026-08-11 | (none) |  | Y | N | N | N |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
EDS Acquisition Mode | Spatial acquisition sub-strategy for EDS measurements: stationary-beam point acquisition, line scan (beam stepped along a transect at defined intervals), or area map / spectrum image (beam rastered over a pixel grid). Specifies how the beam is positioned during data collection within the declared analytical mode. Record 'N/A' where the procedure has no EDS detector. 'Point' covers what the literature also calls spot or point-spectrum analysis. 'Map' and 'Spectrum image' are distinct acquisitions, not synonyms: a map may retain element intensities alone, whereas a spectrum image retains a full spectrum at every pixel and can be requantified afterwards — record which was acquired. Where more than one mode was used, join them with '; ' rather than looking for a combined member. | Basic | Read-Only | Controlled list | Point ¦ Line scan ¦ Map ¦ Spectrum image ¦ N/A ¦ None |  | 2026-08-31 | (none) | Particularly important when a procedure includes linescans as a distinct acquisition approach not fully captured by the mode flag columns. | Y | Y | N | N | [15 empty] | N/A
Dwell Time per Pixel | Time spent acquiring X-ray signal at each pixel during X-ray mapping, in milliseconds. For WDS: one value per spectrometer assignment per pixel. For EDS: total live-time per spectrum per pixel, a single value. | Basic | Editable | Numeric + unit | e.g., '10 ms' ¦ '100 ms' ¦ '0.5 s' |  | 2026-09-10 | monitored property |  | N | Y | N | Y |  | N/A | N/A | ~0.5 s per step (olivine megacryst Ka maps; instrument ambiguous) | N/A | N | N | N | N/A | N/A | N/A | N/A | N | N/A | N/A | 25 msec
Step Size / Pixel Size | Distance between adjacent measurement points in the X-ray map in micrometers, defining the spatial resolution. Report both X and Y step if they differ. | Basic | Editable | Numeric (µm) / Text | e.g., 0.5 ¦ 1 ¦ 2 ¦ 5 ¦ 10; or '1 um (X) x 2 um (Y)' |  | 2026-08-26 | (none) |  | N | Y | N | Y |  | N/A | N/A | 8-12 µm (BSE + Ca/Al/Fe/Mg Ka phase maps at UT); 2 µm (olivine megacryst Ka maps; instrument ambiguous) | N/A | N | N | N | N/A | N/A | N/A | N/A | N | N/A | N/A | 9.5 um (stage maps); ~1.5 um per pixel (BSE mosaic)
Map Dimensions | Number of pixels in the X-ray map in the X and Y directions. Based on the area of interest and selected step size. | N/A | Basic | Numeric pair (pixels x pixels) | e.g., 256x256 ¦ 512x512 ¦ 1024x768 |  | 2026-08-27 | (none) |  | N | Y | N | Y |  | N/A | N/A | N | N/A | N | N | N | N/A | N/A | N/A | N/A | N | N/A | N/A | 1,024 x 1,024 pixels per stage map; five stage maps per thin section
Map Area | Physical extent of the mapped region, given either as width × height in µm or as a total area in µm² or mm², and equal to (map width in pixels × step size) × (map height in pixels × step size). Complements the map's pixel-grid dimensions by recording the physical scale of the mapped region. | N/A | Basic | Numeric + unit | e.g., '100 × 100 µm' ¦ '500 × 300 µm' ¦ '1000 × 500 µm' ¦ '0.2 mm²' |  | 2026-08-27 | (none) | Based on the sample feature or region of interest, so it is chosen to cover the target. | N | Y | N | Y |  | N/A | N/A | N | N/A | N | N | N | N/A | N/A | N/A | N/A | N | N/A | N/A | N - not stated directly (1,024 pixels at 9.5 um step corresponds to ~9.7 mm per side)
Stage Scan vs. Beam Scan | For mapping modes, whether the map was acquired by moving the stage while the beam is held fixed (stage scan), or by deflecting the beam across the field while the stage is stationary (beam scan). | Advanced | Read-Only | Controlled list | Stage scan ¦ Beam scan ¦ Combined stage and beam scan ¦ Unknown ¦ N/A ¦ None |  | 2026-08-30 | (none) | Stage scan preserves beam geometry and spot size across the full map but is slower; beam scan is faster but may introduce geometric distortion at large deflections. | N | Y | N | Y |  | N/A | N/A | N | N/A | N | N | Stage scan | N/A | N/A | N/A | N/A | N | N/A | N/A | Stage scan (stated 'EPMA stage maps')
5. Data Processing | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Matrix Correction Method | X-ray matrix correction algorithm applied during quantitative EDS or WDS data reduction. For X-ray mapping, applies when raw count maps are converted to quantitative concentration maps. Where the k-factors or calibration constants themselves came from — measured standards, a vendor library, or theoretical cross-sections — is a separate question answered by this technique's calibration-standard field, not here; a procedure may be both absorption-corrected and standardless. | Basic | Read-Only | Controlled list / Text | PAP (Pouchou & Pichoir Full) ¦ XPP (Simplified PAP) ¦ PhiRhoZ Bastin (EPQ-91) ¦ Love-Scott I ¦ Love-Scott II ¦ Armstrong / Love-Scott ¦ ZAF ¦ CITZAF (Armstrong 1995) ¦ Bence-Albee ¦ Unknown ¦ N/A ¦ None |  | 2026-08-30 | (none) |  | Y | Y | Y | Y |  | CITZAF (Armstrong 1995) | Bence-Albee | N | N | CITZAF (Armstrong 1995) | N | CITZAF (Armstrong 1995) | N | ZAF | N | N | N | N | N | Full Phi(rho-z) correction applied at each pixel, of the form C = k x ZAF, where ZAF is the compositionally dependent correction for atomic number, X-ray absorption and characteristic fluorescence in both sample and standard
Mass Absorption Coefficients (MACs) | Database of mass absorption coefficients used in the matrix correction. | Basic | Read-Only | Controlled list | LINEMU ¦ CITZMU ¦ MCMASTER ¦ MAC30 ¦ MACJTA ¦ FFAST ¦ Unknown ¦ N/A ¦ None |  | 2026-08-30 | (none) | MAC database choice affects accuracy particularly for light elements (B, C, N, O, F, Na). | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
X-ray Background Correction Method | Method used to estimate and subtract background X-ray intensity beneath the peak. For WDS: typically 2-point off-peak linear interpolation or Mean Atomic Number (MAN) background model. For EDS: spectral background fitting or top-hat filter applied during spectral processing. | Basic | Editable | Controlled list | 2-point off-peak linear ¦ 2-point off-peak exponential ¦ 2-point off-peak polynomial ¦ 2-point off-peak (interpolation not specified) ¦ 1-point high with slope factor ¦ 1-point low with slope factor ¦ Mean Atomic Number (MAN) ¦ EDS background fit ¦ Top-hat filter ¦ N/A ¦ None |  | 2026-09-10 | monitored property |  | Y | Y | Y | Y |  | N | N | N | Two-point off-peak (high and low positions measured; interpolation method not named) | N | N | Mean Atomic Number (MAN) (most analytes); polynomial fit on representative phases (for F, LDE1 crystal) | N | N | N | N | Two-point off-peak (high and low positions; 10 s each for silicates/sulfides/oxides and phosphates; 5 s each for carbonates) | N | N | MAN (mean atomic number) background calibration, using EPMA standards spanning a range of average atomic number Z
Time-Dependent Intensity Correction | Type of time-dependent intensity (TDI) correction applied to compensate for beam-induced volatilization or migration of sensitive elements (e.g., Na, K, F in glasses, feldspars, carbonates). | Advanced | Editable | Controlled list | None ¦ Linear ¦ Exponential ¦ Double exponential ¦ Unknown ¦ N/A |  | 2026-09-01 | target species | Most commonly applied in point analysis; uncommon for X-ray mapping. | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Target Species Estimation Method | Whether elemental concentrations were calculated directly from measured X-ray intensities, or estimated by cation stoichiometry (e.g., oxygen calculated from cation proportions in silicates; carbon from stoichiometry in carbonates). | Advanced | Read-Only | Controlled list / Text | Direct (measured) ¦ Stoichiometry from cations ¦ By difference ¦ Unknown ¦ N/A ¦ None |  | 2026-09-01 | target species |  | Y | Y | Y | Y |  | N | N | N | N | N | N | Stoichiometry from cations (O by stoichiometry) | Ketcham (2015) method (13-anion basis); OH by difference from 1-F-Cl=OH | N | N | N | N | N | N | N
Halogen Correction on Oxygen | Whether oxygen content was adjusted to account for halogen substitution (F and/or Cl replacing OH) in halogen-bearing phases such as apatite, amphibole, and mica, where oxygen is calculated by stoichiometry. | Advanced | Editable | Controlled list / Text | e.g., 'Yes — F and Cl substitution in apatite; O calculated as 1-F-Cl=OH' ¦ 'No' ¦ N/A ¦ None |  | 2026-08-30 | (none) |  | Y | Y | Y | Y |  | N/A | N/A | N/A | N/A | N/A | N/A | N | Yes (F and Cl substitution in apatite; 1-F-Cl=OH) | N/A | N/A | N | N | N/A | N/A | N/A
WDS Dead Time Correction | Method used to correct for WDS proportional counter dead time at high count rates. Unlike EDS dead time — which is hardware-managed and reported as a session QC percentage (see EDS Dead Time) — WDS dead time correction is a user-selectable algorithm in the data reduction software. No separate measured WDS dead time value is reported; the correction is applied transparently during intensity-to-concentration conversion. Record the algorithm here and any instrument-specific constant alongside it — 'Default constant (manufacturer) — 3 µs, Cameca'. The instrument vendor itself is recorded by Instrument Manufacturer, not by this field's allowed values. | Basic | Read-Only | Controlled list / Text | Default constant (manufacturer) ¦ Adjusted constant ¦ Logarithmic ¦ High-precision (Probe for EPMA) ¦ Super-precision (Probe for EPMA) ¦ Unknown ¦ N/A ¦ None |  | 2026-08-31 | (none) | Dead time errors are most significant for major elements with high count rates (e.g., Si, Fe, Ca). | N | N | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
EDS Spectral Processing Type | Method used to process EDS spectra and extract net peak intensities from raw spectral data. Applied before quantification (see Matrix Correction Method). Common approaches include background fitting and subtraction followed by peak integration, and filter fit or Gaussian deconvolution for overlapping peaks. | Advanced | Read-Only | Controlled list | Background fitting ¦ Peak deconvolution ¦ Top-hat filter ¦ Gaussian fit ¦ Filter fit ¦ Unknown ¦ N/A ¦ None |  | 2026-08-30 | (none) |  | Y | Y | N | N |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Blank Correction | Method and reference material(s) used to determine and subtract blank signal contributions (e.g., carbon coat contribution to C signal, or background contamination for trace elements). | Advanced | Editable | Text (free) | e.g., 'Blank C signal subtracted using carbon-coated polished glass reference' ¦ 'None' |  | 2026-09-01 | target species |  | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Normalization / Standards-Based Correction | Post-acquisition normalization applied to the reported data beyond the primary calibration — for example correction to a reference value derived from secondary reference materials, or correction for a systematic bias those materials reveal. Record 'None' if no additional normalization is applied. | Advanced | Editable | Text (free) | e.g., 'Normalization to accepted values of Kakanui kaersutite run as secondary standard' ¦ 'None' | Source: Composition QC module | 2026-08-27 | reported property |  | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | k-ratio against the calibration standard: k = (P-B)smp / (P-B)std, the background-corrected relative peak X-ray intensity at each pixel compared to the calibration standard
Calibration Factor and Determination Method | An externally-calibrated factor that converts the measured quantity into the reported quantity, how it was determined, and its uncertainty. Applies where the conversion depends on a factor calibrated against a reference of independently known value, rather than on the instrument response alone. Distinct from the fields that name the calibration material and that state which approach applies to which target species, where the technique has them: this field records the resulting factor itself. | Basic | Editable | Text (free) |  | Source: Calibration Factor module | 2026-09-01 | reported property |  | Y | Y | Y | Y | [15 empty] | N
Procedural Blank Level | The measured level of the analytical blank in the session, and — where the reported quantity is a ratio — its composition, since a blank subtracted from a ratio biases the result unless its own composition is known. Companion to the blank correction method. | N/A | Basic | Text (free) |  | Source: Blank module | 2026-09-01 | target species | Follows the criterion-versus-measurement split the library applies wherever a procedure sets a threshold and an analysis reports a value against it. | Y | Y | Y | Y | [15 empty] | N/A
Analysis Inclusion and Rejection Criteria | The rules determining which individual analyses contribute to a reported aggregate result, together with the outcome of applying them: how many analyses were acquired, how many were included, and on what grounds any were excluded. Distinct from within-analysis outlier filtering, which removes anomalous points inside a single analysis: this field decides which whole analyses enter the reported value. | Basic | Basic | Text (free) |  | Source: Aggregation module | 2026-08-25 | (none) | Criteria and outcome are combined in one field because neither is interpretable without the other, following the precision/accuracy precedent. | Y | Y | Y | Y | [15 empty] | N
Constants and Reference Values Used | Physical constants and reference values used in data reduction to calculate the final reported quantity (e.g., decay constants for age calculation, standard isotope ratios, or other citable reference values used in a correction or calculation), together with their source. Distinct from the Group 6 reference-material fields, which document accepted values for specific calibration/validation materials rather than universal physical constants. Record "None" if no citable, revisable physical constants feed into this procedure's data reduction. | Basic | Editable | Text (free) | e.g., "λ238U = 1.55125×10⁻¹⁰ yr⁻¹, λ235U = 9.8485×10⁻¹⁰ yr⁻¹ (Jaffey et al. 1971); 238U/235U = 137.818 (Hiess et al. 2012)" ¦ "None" | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y | [15 empty] | N
6. Quality Control & Uncertainty | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Primary Calibration Standard Name | Name and reference material identifier of the primary reference material(s) against which the instrument is calibrated — converting raw signal intensities to concentrations, or anchoring an isotope ratio as the bracketing standard or zero-delta reference. Give the material name, its source or supplier, and a citation for the accepted values used. Where calibration instead uses the vendor's stored library or theoretical response factors rather than measured reference materials — 'standardless' or 'semi-quantitative' quantification — record that here, naming the library or model used. 'None' means no calibration was performed at all, which is a different answer. | Basic | Editable | Text (free) | e.g., 'Smithsonian anorthite (Si Ka, Al Ka, Ca Ka)' ¦ 'Amelia albite (Na Ka)' ¦ 'Synthetic fayalite (Fe Ka)' ¦ 'SrF2 (F Ka)' | Source: Composition QC module | 2026-09-01 | target species | Results calibrated against different published values for the same material are not directly comparable. | Y | Y | Y | Y |  | Anorthite (SiKα, AlKα, CaKα); albite (NaKα); fayalite (FeKα); forsterite (MgKα); Mn2SiO4 (MnKα); TiO2 (TiKα); Cr2O3 (CrKα); microcline (KKα) | Natural kaersutite (Si, Mg, Fe); jadeite (Na, Al); bustamite (Ca, Mn); K-feldspar (K); synthetic rutile (Ti); Cr2O3 (Cr) | N | N | Asbestos microcline (SiKa, AlKa, KKa); synthetic anorthite (CaKa); Amelia albite (NaKa); synthetic fayalite (FeKa); synthetic forsterite (MgKa); synthetic TiO2 (TiKa); synthetic Cr2O3 (CrKa); synthetic Mn-olivine (MnKa) | Kakanui kaersutite (Si, Al, Ti, K, Na, Fe, Mg, Ca); Canyon Diablo troilite (S); rhodonite (Mn); chromium metal (Cr); nickel metal (Ni); apatite (P); vanadium metal (V) | Natural and synthetic minerals routinely used in analytical facility (specific names N); synthetic F-phlogopite (for F, LDE1 crystal) | SrF2 (F); albite (Na); olivine (Mg, Fe, Si); quartz (Si); apatite (Ca, P); barite (S); tugtupite (Cl); rhodonite (Mn); ilmenite (Fe, Ti) | Natural and synthetic mineral standards (specific names N) | Carbonates: magnetite USNM 114887 (Fe,Mn), calcite USNM 13621 (Ca), dolomite USNM 10057 (Mg), siderite R-2460, rhodonite; Silicates/oxides: chromite USNM 117075, ilmenite USNM 96189, magnetite USNM 114887, manganite USNM 157872, bytownite R-2912, forsterite P140, San Carlos olivine USNM 111312 (Fo90), Springwater olivine USNM 2566 (Fo83) | Phosphates: fluorapatite (F, P, Ca), Fo92 olivine (Si, Mg), rhodonite (Mn), fayalite (Fe), anorthite (Al), baryte (S), K-feldspar (K), scapolite (Cl); Carbonates: albite (Na), Fo olivine (Si), dolomite (Mg), calcite (Ca), Mn carbonate (Mn), apatite (P), baryte (S), fayalite (Fe) | Well-characterized natural and synthetic materials (specific names N) | Springwater olivine (Mg, Si); fayalite (Fe); wollastonite (Ca); albite (Na, Al); orthoclase (K); rutile (Ti); Ni metal (Ni); chromite (Cr); rhodochrosite (Mn) | N | N - 'EPMA standards having a range of average atomic number Z' are used for the MAN background calibration; individual standards not named
Secondary Reference Materials | Quality-control reference material(s) measured as unknowns alongside samples to assess accuracy independently and to monitor drift. Give the material name, its source, and a citation for the accepted or reference values used for comparison. | Basic | Editable | Text (free) | e.g., 'Kakanui kaersutite (Si, Al, Ti, Fe, Mg, Ca, Na, K; n=5; within 1% of GeoReM values)' ¦ 'Durango apatite (F, Cl, P, Ca; Wudarska et al. 2021)' ¦ 'USNM San Carlos olivine 111312 (Mg, Fe, Si)' | Source: Composition QC module | 2026-08-27 | defines: standard |  | Y | Y | Y | Y |  | N | N | N | N | Feldspar standards run as unknowns (material names N beyond what is listed above) | USNM San Carlos olivine (Fo90); Kakanui kaersutite | Smithsonian Microbeam standards (specific materials and values N) | N | N | Carbonates: calcite, dolomite, rhodochrosite; Silicates/oxides: magnetite, San Carlos olivine USNM 111312, Springwater olivine USNM 2566 | N | N | N | N | N
X-ray Line Overlap Corrections Applied | Whether a spectral interference correction was applied. Common interferences include Ti Kb on V Ka, Cr Kb on Mn Ka, and Ba La on Ti Ka. | Basic | Read-Only | Controlled list / Text | Yes — specify the overlapping lines and the correction method ¦ No ¦ N/A ¦ None |  | 2026-09-01 | target species |  | Y | Y | Y | Y |  | N | Yes | N | N | N | N | N | N | N | N | N | N | N | N | N
Interfering Elements | Element(s) whose X-ray lines overlap with the measured peak, requiring a correction. | Advanced | Read-Only | Text (free) | e.g., 'Ti Kb overlaps V Ka' ¦ 'Cr Kb overlaps Mn Ka' ¦ 'Ba La overlaps Ti Ka' |  | 2026-09-01 | target species |  | Y | Y | Y | Y |  |  | Cr Kβ on Mn Kα | [10 empty] | N | N | N
Interference Correction Standard | Reference material used to quantify and calibrate the interference correction. | Advanced | Read-Only | Text (free) | e.g., 'Synthetic TiO2 (for Ti Kb on V correction)' ¦ 'Chromite USNM 117075 (for Cr Kb on Mn correction)' |  | 2026-09-01 | target species |  | Y | Y | Y | Y | [13 empty] | N | N | N
Detection Limit | Detection limit, one per reported concentration variable (one per target species, these being the same set). State the units and whether the values are procedure-typical estimates or session-specific measured values. The calculation method is recorded separately in Detection Limit Method. Record 'N/A' where the procedure reports no concentrations. | Advanced | Basic | Numeric + unit / Text | e.g., 'SiO2: 0.02 wt% (3-sigma, 20 s peak, 10 s bg)' ¦ 'Na2O: 0.03 wt%' ¦ 'F: 0.04 wt%' | Source: Composition QC module | 2026-09-01 | reported property |  | Y | Y | Y | Y |  | K: 0.02 wt%; Cr: 0.05 wt%; Mn: 0.06 wt% (Table 1 footnote; other elements N) | K2O: 0.01 wt% (lowest); MnO: 0.06 wt% (highest); full per-element values stated in paper | <0.03 wt% for SiO2, TiO2, Al2O3, MgO, CaO; <0.05-0.1 wt% for FeO, MnO, Cr2O3, NiO, Na2O, K2O, P2O5 (stated as "typical") | <0.03 wt% for SiO2, TiO2, Al2O3, MgO, CaO; <0.05-0.1 wt% for FeO, MnO, Cr2O3, NiO, Na2O, K2O, P2O5 (stated as shared conditions with UT) | 0.05 Si, 0.04 Ti, 0.06 Al, 0.06 Fe, 0.02 Mg, 0.02 Ca, 0.03 Na, 0.02 K, 0.05 Cr, 0.06 Mn (all wt%) | Al2O3, K2O, CaO: 0.03-0.04 wt%; Na2O, MgO, SiO2, FeO, MnO: 0.05 wt%; P2O5, SO2, TiO2, V2O3, Cr2O3, NiO: 0.06-0.09 wt% (stated as "typically") | N | N | Better than 0.02 wt% (as stated in paper) | N | N | N | Mg: 0.025 wt%; Fe: 0.025 wt%; Si: 0.05 wt%; K: 0.05 wt%; Na: 0.05 wt%; Ca: 0.005 wt%; Al: 0.02 wt%; Ti: 0.005 wt%; Cr: 0.015 wt%; Mn: 0.008 wt% | Transition metals: ~250 ppm (stated as "typical") | 0.1-0.2 element wt.% for all elements in this map set
Detection Limit Method | Formula or approach used to calculate the detection limits, with a citation for the method where one exists. Must be consistent with the values reported in Detection Limit. | Basic | Read-Only | Controlled list / Text | 3σ blank ¦ 3σ background ¦ 3σ counting statistics ¦ 3× blank mean ¦ Poisson statistics ¦ N/A ¦ None | Source: Composition QC module | 2026-08-27 | reported property |  | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N - attributed to the MAN background correction dedicating all map collection time to on-peak measurement, 'which improves precision and detection limits'
Analytical Precision | Reproducibility of repeated measurements on the same or equivalent reference material, expressed as 1-sigma relative standard deviation (%). Include reference material name, number of analyses (n), and the measured value. | Advanced | Basic | Text (free) | e.g., 'SiO2: 0.5% 1-sigma RSD (n=20, Kakanui kaersutite)' ¦ 'FeO: 1.2% 1-sigma RSD (n=15, San Carlos olivine)' |  | 2026-08-27 | standard x reported property |  | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Analytical Accuracy | Offset between measured and accepted reference values for secondary standards, expressed as percent relative bias. Include reference material, reference value source, and the measured value. | Advanced | Basic | Text (free) | e.g., 'SiO2: +0.5% relative to GeoReM preferred value (Kakanui kaersutite)' ¦ 'CaO: -1.2% relative to accepted value' |  | 2026-08-27 | standard x reported property |  | Y | Y | Y | Y |  | N | N | N | N | 1-2% for Si, Al, Ca, Na, and K (based on feldspar standards run as unknowns) | N | N | N | N | N | N | N | N | N | N
Counting Statistics Error | Uncertainty predicted from counting statistics — the theoretical limit set by the Poisson distribution of the counts accumulated — for each reported quantity per analysis, with the sigma level stated. Derived from the counts on the target species together with those on any background or blank subtracted from it. Distinct from the scatter actually observed within a measurement or between repeated measurements, which is recorded separately. | Advanced | Basic | Text (free) | e.g., 'SiO2: +-0.04 wt% (1-sigma)' ¦ 'Na2O: +-0.03 wt% (1-sigma)' | Source: Composition QC module | 2026-09-01 | sample > sampling unit x reported property | Where a procedure reports both, agreement indicates the measurement is shot-noise limited, and a larger observed scatter indicates a further source of variance. | Y | Y | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
EDS Dead Time | Percent dead time reported by the EDS detector during the session — the fraction of total acquisition time the detector spent processing rather than counting. This field documents the resulting percentage as a session QC metric. Unlike WDS dead time (see WDS Dead Time Correction), no user-selectable correction algorithm is required. | N/A | Basic | Numeric (%) | e.g., 10% ¦ 25% ¦ <40% |  | 2026-08-27 | (none) | Values above ~40% indicate excessive count rate and may degrade spectral quality and quantitative accuracy. EDS dead time correction is managed automatically by the detector electronics. | Y | Y | N | N |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Goodness-of-Fit or Dispersion Statistic | The statistic reported to show whether scatter among the contributing analyses exceeds what analytical uncertainty alone predicts, together with its value. The procedure may still state an acceptance threshold, which belongs with the inclusion criteria. | N/A | Basic | Text (free) |  | Source: Aggregation module | 2026-08-25 | reported property | Answers whether a reported aggregate is defensible as a single population. The value cannot be known before the analysis. | Y | Y | Y | Y | [15 empty] | N
Additional Notes | Any procedure- or analysis-specific information not captured by a structured field anywhere in this TAPP — including anomalies, deviations from the registered procedure, instrument modifications, and supplementary context. Scope is the whole document, not Group 6: this is the last field of the TAPP and covers all six groups. Use sparingly; a structured field is preferred for anything that can be formally categorised. | Advanced | Advanced | Text (free) | Free text | Source: Core module | 2026-08-25 | (none) |  | Y | Y | Y | Y |  | Ma et al. 2015, Earth Planet. Sci. Lett. — tissintite discovery paper (Tissint Mars meteorite). Instrument stated as "JEOL 8200 electron microprobe" (no JXA prefix). WDS explicitly stated ("WDS: 15 kV; 5 nA; beam in focused mode"). Point analysis only; no X-ray mapping reported. Probe for EPMA stated; CITZAF correction procedure (Armstrong 1995). Full standard suite with X-ray lines given. Detection limits: K=0.02, Cr=0.05, Mn=0.06 wt% from Table 1 footnote. Caltech GPS Division Analytical Facility. | Hu et al. 2020, Geochim. Cosmochim. Acta — coesite in NWA 8657 shergottite. JEOL JXA-8100 at IGGCAS; 15 kV, 10 nA; point analysis WDS only. Matrix correction: Bence-Albee (not PAP). Full primary standard suite stated (kaersutite, jadeite, bustamite, K-feldspar, rutile, Cr2O3). Mn Kα / Cr Kβ interference correction applied. Detection limits 0.01-0.06 wt% stated per element. Analytical software not stated. | Liu et al. 2016, Meteorit. Planet. Sci. — Tissint mineral chemistry. Protocol 1 of 2: University of Tennessee Cameca SX100. Same paper also uses Caltech GPS JXA-8200 (see Liu+2016_Cal column). Point analysis AND X-ray mapping performed at UT. Specific mapping: BSE + Ca/Al/Fe/Mg Ka maps (15 kV, 20 nA, step 8-12 µm). Olivine megacryst mapping (15 kV, 200 nA, step 2 µm, dwell ~0.5 s) described as "using the EMP" — instrument ambiguous (may be UT or Caltech instrument). Standards, matrix correction, and software not stated. | Liu et al. 2016, Meteorit. Planet. Sci. — Tissint mineral chemistry. Protocol 2 of 2: Caltech GPS Division JEOL JXA-8200. Point analysis only (no mapping attributed to Caltech instrument). Conditions stated jointly for UT and Caltech instruments. Standards, matrix correction, and software not stated for EPMA. | Ma et al. 2018, Meteorit. Planet. Sci. 53:50-61 (file dated 2017) — liebermannite (KAlSi3O8) discovery from Zagami. Instrument stated as "JEOL 8200 electron microprobe" (no JXA prefix in text). WDS explicitly stated ("WDS: 15 kV, 5 nA"). Probe for EPMA; CITZAF correction (Armstrong 1995) — NOT PAP. Full standard suite and X-ray lines stated. K-mapping by EPMA also performed (used for mineral identification) but mapping conditions (step size, dwell time, current) not stated. Na diffusion observed during analysis despite low 5 nA beam current. Detection limits stated (per-element wt% values). Analytical accuracy: 1-2% for Si, Al, Ca, Na, K (feldspar standards as unknowns). Caltech GPS Division Analytical Facility. | Frank et al. 2023, Meteorit. Planet. Sci. 58:1495-1511 — CAI in Ivuna CI chondrite. ARES NASA JSC. Instrument stated as "Cameca SX100 electron microprobe at ARES, Johnson Space Center" — NOT JEOL JXA-8530F as in v2 header. Accelerating voltage 20 kV (not 15 kV). Both point analysis (20 kV, 20 nA, 1 µm focused) and X-ray mapping performed. X-ray mapping described but conditions (step size, dwell time, mapping beam mode) N. WDS not explicitly stated. Matrix correction and background correction method N. Peak counting time 10-50 s. Primary standard suite fully documented. Secondary standards: USNM San Carlos olivine (Fo90); Kakanui kaersutite. Detection limits stated per element group. | Broussard et al. 2026, Meteorit. Planet. Sci. — OC002 CI chondrite links Bennu and Ryugu. Washington University in St. Louis. Instrument stated as "JEOL JXA-8200 electron microprobe" — NOT JXA-8230 as in v2 header. WDS explicitly stated ("wavelength-dispersive quantitative compositional mapping and analysis"). CITZAF matrix correction (Armstrong 1995) — NOT PAP or XPP. MAN background for most analytes; polynomial fit for F via LDE1 crystal. Both point analysis (15 kV, 25 nA) and quantitative stage mapping performed. O by stoichiometry from cations. F is the only explicitly named analyte in methods; full list N. EDS spectrometer present but not used for quantitative analyses. Smithsonian Microbeam standards as secondary QC. No peak counting time, beam diameter, detection limits, or interference corrections stated. | Seifert et al. 2026, Meteorit. Planet. Sci. — apatite in Bennu OSIRIS-REx samples. ARES NASA JSC. Instrument stated as "JEOL 8530 EMPA at NASA JSC" (no "JXA", no "F", no "+" suffix stated in paper). Analytical conditions: 15 kV, 20 nA, 2 µm probe size. Previous v2 values of 10/40-100 nA and 10 µm beam were WRONG — those were Durango apatite test conditions used to assess beam damage, not the actual protocol. Analytes: P, F, Cl, Ca, Mn, Fe, Na, Mg, Si, S. Apatite stoichiometry by Ketcham (2015) method (13-anion basis; OH by difference). Halogen correction on O: Yes. Primary standards: SrF2, albite, olivine, quartz, apatite, barite, tugtupite, rhodonite, ilmenite. Sample preparation: fragments embedded in epoxy, dry-polished, ion-polished (one mount), carbon coated. 14 total analyses performed. | Pang et al. 2016, Sci. Rep. 6:26063 — NWA 8003 eucrite, Nanjing University. JEOL JXA-8100 (stated as "JEOL 8100"). WDS explicitly stated ("JEOL 8100 WDS"). ZAF matrix correction (NOT "ZAF or PAP" as in v2; paper states ZAF). Focused beam (20 nA) for most phases; defocused 2-5 µm for plagioclase and polymorphs. Natural and synthetic mineral standards (specific names N). Detection limit better than 0.02 wt% (as stated). Analytical software not stated. | McCoy et al. 2025, Nature 637:320-325 — Bennu evaporites. Protocol 1 of 2: Smithsonian Institution JEOL 8530 F+ Hyperprobe (Field Emission). Ir-coated specimens mounted on Ir-coated Parafilm. Carbonate analyses: 15 kV, 10 nA, 5 µm spot; LIFL (Fe,Mn), TAPL (Mg), PETL (Ca). Silicate/oxide analyses: 15 kV, 10 nA, 1 µm spot; broader standard suite. Both primary and secondary standard suites fully documented with USNM catalog numbers. Acquisition software and matrix correction method N. WDS not explicitly stated in text (crystal designations LIFL/TAPL/PETL confirm WDS use). | McCoy et al. 2025, Nature 637:320-325 — Bennu evaporites. Protocol 2 of 2: U. Arizona K-ALFAA Cameca SX-100. 20 nm carbon coat. WDS explicitly stated for phosphate analyses. Mg,Na phosphate analyses: 15 kV, 8 nA, 1 µm. Carbonate analyses at K-ALFAA also mentioned; conditions N. Full primary standard suite documented for phosphates and carbonates. Acquisition software and matrix correction method N. | Zega et al. 2025, Nat. Geosci. — mineralogical evidence for hydrothermal alteration of Bennu. K-ALFAA, University of Arizona. Instrument stated as "SX-100 Ultra electron microprobe in the K-ALFAA". IMPORTANT: v2 had "no protocol details reported" — this was WRONG. The paper provides detailed EPMA conditions: X-ray maps and BSE images: 15 kV, 20 nA. Silicates/sulfides/oxides: 15 kV, 20 nA, focused, 20 s peak, 10 s/bg each side. Phosphates: 15 kV, 8 nA, 2 µm defocused, 20 s peak, 10 s/bg each side. Carbonates: 15 kV, 4 nA, 2 µm, 10 s peak, 5 s/bg each side. Standards: "well-characterized natural and synthetic materials" (specific names N). Phase maps generated using XMapTools. WDS and matrix correction NOT explicitly stated in paper. | Barnes et al. 2025, Nat. Astron. — variety and origin of accreted materials in Bennu. Protocol 1 of 2: CRPG Nancy, JEOL JXA-8230. Instrument has 5 WDS spectrometers + 1 SDD EDS; per-analyte technique (WDS vs. EDS) not stated. Two analytical sessions: session 1 (no Na, K); session 2 (with Na, K). Counting times are stated as total peak + background combined: 200 ms for minor elements (Al, Ti, Ca, Mn, Cr) and 20 ms for major elements (Mg, Fe, Si) — unusually short, possibly per-pixel for fast mapping mode. Full primary standard suite stated with element assignments. Full per-element detection limits stated. Matrix correction method not stated. Sample preparation done at Université Côte d'Azur (not at CRPG). Beam current not stated for NHM protocol; 3 nA mentioned in text is for SEM-EDS (different instrument). | Barnes et al. 2025, Nat. Astron. — variety and origin of accreted materials in Bennu. Protocol 2 of 2: NHM London, CAMECA SX100. Stated instrument: "CAMECA SX100 electron microprobe". Target minerals: olivine and pyroxene (anhydrous silicates). 20 kV, 1 µm focused beam. Beam current not stated for EPMA (3 nA in text refers to SEM-EDS on separate Zeiss EVO instrument). Detection limits ~250 ppm for transition metals. Standards, matrix correction, WDS spectrometer details not stated. Analyte list not explicitly given; implied Si, Mg, Fe, Ca, Mn, Cr, Ni, Al, Ti from context. SEM-EDS at NHM is a separate instrument (Zeiss EVO 15LS + Oxford X-Max80) calibrated at 20 kV, 3 nA. Carbon coat: initial coat for SEM/EPMA (thickness N); additional coat to ~30 nm total was for subsequent SIMS, not EPMA. | Multi-pass WDS mapping: two passes per stage map, five elements each; 18 hr per map; 20 x 10^6 fully quantitative analyses across all slides. Recorded in the acquisition-pass proposal (2026-09-08) as the evidence that EPMA partitions the target-species domain across passes.

## Legends
Procedure-Level Tier | Description | Analysis-Level Tier | Description
Basic | Mandatory for procedure registration. Must be provided to register a valid procedure. | Read-Only | Directly imported from the registered procedure; cannot be changed by the analyst. Changing this value means running a different procedure. Also used for fields relevant only at procedure level — the value is inherited from the procedure record and shown read-only in the analysis form.
Advanced | Optional for procedure registration. Strongly recommended but not required. | Editable | Imported from the registered procedure but may be adjusted within procedure-defined bounds (e.g., daily tuning, minor software updates). The procedure registers the target or typical value; the analyst confirms or adjusts it. Cannot be left void if the procedure-level tier is Basic.
N/A | Not applicable at procedure level. This field captures analysis-level information only. | Basic | Mandatory user input at analysis time. Value comes from the analysis itself and cannot be pre-specified in the procedure.
|  | Advanced | Optional user input at analysis time. Recommended for complete and reproducible documentation.
Mode Column | Definition
Purpose
EDS Point Analysis
EDS Mapping
WDS Point Analysis
WDS Mapping
Y | This field applies to this analytical mode and should be reported.
N | This field does not apply to this analytical mode.
Data Type | What to enter
Controlled list | Choose one of the values listed in "Example / Allowed Content". The list is CLOSED — it is meant to cover every case. If your value genuinely is not there, that is a gap in the list worth reporting, not a reason to write your own.
Controlled list / Text | Choose a listed value AND qualify it in the same cell — the detail is expected, not optional. Typically what was corrected, which target species or phases it applied to, the method or equation used, or the source it came from. An answer the list cannot express is also valid here.
Numeric (unit) | A number in the unit named in brackets. Enter the number only — the unit is fixed by the field.
Numeric + unit | A number AND its unit, because the unit varies between procedures. Write both, e.g. '50 us' or '0.5 s'.
Text (free) | Free text. No controlled vocabulary applies.
N/A ¦ None | Offered by every controlled list. "N/A" means the field does not apply to this procedure; "None" means it applies but nothing was used. Prefer either over leaving the cell blank.
Keyed By (Rule 7) | What the field's value repeats over
(none) | Scalar — one value per procedure or analysis. The default.
defines: monitored property per target species | ENUMERATES the monitored property domain — the header of the child table, not a column in it — and does so once per target species. One value per chemical species determined, at whatever granularity the procedure determines it.
defines: reported property | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
defines: sample | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it.
defines: sampling unit | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per subdivision of the physical sample that carries its own row — grain, spot, aliquot, phase, sub-volume.
defines: standard | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per reference material or reference database entry.
defines: target species | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per chemical species determined, at whatever granularity the procedure determines it.
monitored property
reported property | One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
sample
sample > sampling unit | Nested — sampling unit exists only within sample. One value per subdivision of the physical sample that carries its own row — grain, spot, aliquot, phase, sub-volume.
sample > sampling unit x reported property | Cross-product, read as "for each sample, one value per reported property". One value per subdivision of the physical sample that carries its own row — grain, spot, aliquot, phase, sub-volume. + One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
standard x reported property | Cross-product, read as "for each standard, one value per reported property". One value per reference material or reference database entry. + One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
target species | One value per chemical species determined, at whatever granularity the procedure determines it.