## Instructions
EPMA TAPP review workbook — EDS Mapping
This workbook asks you to review the EPMA TAPP: the list of metadata items that should be reported when an EPMA procedure is registered, and when a measurement session run under that procedure is documented. It has four tabs: Instructions (this tab), Reviewer information, EPMA TAPP, and Registration Metadata. Please review the 61 metadata items on the EPMA TAPP tab. The Registration Metadata tab lists 14 items that are set by registration requirements; they need no review, though you may comment on them in "Other comments".
This copy covers EDS Mapping only: the items that apply to that mode, 75 of the 92 in the full TAPP. "#" keeps the TAPP's own numbering, so gaps in the numbers are items that belong to other modes.
Every item is described for two records. The procedure record is the standing recipe, registered once and cited by DOI. The session record describes one measurement session run under that procedure; a session may cover several samples. The procedure record is filled in once per analytical setup, typically from a representative session. A session record is filled in when data from that session are submitted, not for every session run on the instrument.
What to do
Step 1 | Fill in the "Reviewer information" tab. Answers go in the "Your answer" column; the questions are locked.
Step 2 | On the "EPMA TAPP" tab, put all your feedback in the six shaded comment columns: "Needed?", "Structure OK?" (twice), "Tier OK?" (twice) and "Other comments". Each one judges the information column on its left; "Other comments" takes everything else. Leave a cell blank if you agree or have no view. All other columns are locked.
Step 3 | Follow the "Review criteria" at the bottom of this tab. Provide review that clearly presents your suggested change and reasons.
Step 4 | When you have finished, save the workbook as EPMA_TAPP_v89_Review_EDS_Mapping_<your initials>.xlsx (e.g., EPMA_TAPP_v89_Review_EDS_Mapping_RD.xlsx) and email it to info@astromat.org. The deadline for this round of review is December 1, 2026.
The EPMA TAPP tab: column explanations
The tab has 8 information columns and 6 comment columns. The information columns are: "#", "Metadata item", "Description", "Example (NOT recommendations)", "Structure — procedure record", "Procedure-Level Tier", "Structure — session record", "Session-Level Tier". The comment columns are: "Needed?", "Structure OK?" (twice), "Tier OK?" (twice) and "Other comments". Each comment column sits immediately to the right of the information column it judges and has the same colour: purple for the metadata item, blue for the procedure record, yellow for the session record.
Column | Meaning
# | The item's number in the TAPP: group number, then position in the group.
Metadata item | The name of the metadata item.
Description | What the item records.
Example (NOT recommendations) | An illustration of the item filled in, from a fictional worked example (an invented EPMA session written up as a methods section). It shows the shape of an answer. It is not recommended content, and the laboratory, samples and results are invented. Samples are CC-1 and CC-2; analysis points are named by mineral and number (ol3 = the third olivine point, ap2 = the second apatite point); maps are Map 1 and Map 2.
Structure — procedure record | How the procedure states the item: one entry, one entry for each entry of a list, or a list of its own. The session record follows it, as shown in the session structure column.
Procedure-Level Tier | Basic = mandatory for procedure registration. Advanced = recommended but optional. N/A = not part of the procedure: it can only be known once a session has run.
Structure — session record | How one measurement session reports the item.
Session-Level Tier | Read-Only = must follow the entry given in the procedure. Editable = may vary within the range the procedure allows; the session records what was used. Basic = the entry comes from the session itself and must be reported. Advanced = the entry comes from the session itself, and reporting it is recommended.
The EPMA TAPP tab: "Tiers"
Basic | Required
Advanced | Recommended
Read-Only | Fixed by the procedure. Session-level only.
Editable | Adjustable within the bounds the procedure sets. Session-level only.
N/A | Not applicable at that level. Procedure-level only.
The EPMA TAPP tab: "Structure"
Each structure cell is an instruction, written in one of these patterns. An entry is one complete answer to the item, in whatever form its Description asks for: a number with its unit, a name, a choice, or a short statement with several parts (for example a reference material with its source and a citation). The parts of one entry are given together, not as separate entries. A list is made of entries too, one per member, and other items refer to its members as entries.
Give one entry for each X. | One entry for every entry in the list X. "For the procedure" or "for the session" means nothing repeats: give a single entry. "One or more entries" means several may be given, but nothing else repeats over them.
For each A, give one entry for each B. | Two independent lists crossed: one entry for every pairing, with A as the outer list. Analytical Precision reads: for each "Secondary Reference Material", give one entry for each "Reported Variable".
For each A, give one entry for each of its B entries. | Nesting: B exists only inside A. For each "Sample Name", give one entry for each of its "Sampling Unit Name" entries — analysis points exist only inside a sample.
Give one entry; each "Sample Name" entry in a session inherits it. | Procedure column only. Samples, analysis points exist only in a session, so the procedure cannot list them: it gives one entry, and every sample (or point) in a session takes it.
List the … | The row is the list itself. Other rows are given for each of its entries, so it has to be filled in first.
Names in quotes are lists, each defined by one row of the TAPP; the number after a name is that row's "#":
Sample Name | the samples analysed in the session
Sampling Unit Name | the analysis points or map areas, per sample
Target Material | the material types the procedure analyses (row "Target Material")
Target Species | the elements the procedure determines (row "Target Species")
Monitored Element | what is measured to determine those elements (row "Monitored Elements")
Reported Variable | the reported quantities (row "Reported Variables and Units")
Secondary Reference Material | the secondary standards (row "Secondary Reference Materials")
Review criteria per column
Needed? | Is this item relevant to EPMA as you practise it? Is it one thing, and distinct from other items? Is it correctly named?
Structure OK? (procedure) | Does the procedure need one entry, one entry for each entry of a list (for example, per "Target Material"), or a list of its own?
Tier OK? (procedure) | Can it be known before any session? If not: "N/A". If so, is it essential for judging data quality or reproducing the work? Then: "Basic"; otherwise: "Advanced".
Structure OK? (session) | Does a session report it once, or once per sample, per analysis point, per element…?
Tier OK? (session) | Must the session copy the procedure's entry unchanged (Should be fixed by the procedure, i.e. "Read-Only")? May it vary within the procedure's bounds (Should be adjustable, i.e. "Editable")? Or is it known only from the session ("Basic"/"Advanced")?
Other comments | Anything else: the wording of Description or Example, missing items, wrong or missing modes, suggestions on the TAPP's structure.

## Reviewer information
About you
Please fill in the shaded cells. All answers are free text. Your details are used only to follow up on your review and to help us evaluate coverage of this review group; they are not published.
Question | Your answer
Reviewer's name
Reviewer's email
Reviewer's institute
Frequently used instrument model
Frequently used data acquisition software
Frequently used data processing software
Other contributors to this review (names and institutes)

## EPMA TAPP
# | Metadata item | Needed? | Description | Example (NOT recommendations) | Structure — procedure record | Structure OK? | Procedure-Level Tier | Tier OK? | Structure — session record | Structure OK? | Session-Level Tier | Tier OK? | Other comments
1 | Group 1: Procedure Identification
1.1–1.14 | Items 1.1–1.14 are on the "Registration Metadata" tab. They are set by registration requirements and need no review.
1.15 | Coupled Technique(s) |  | Other techniques applied to the same samples whose results are interpreted together with these, with the instrument, laboratory or purpose of each; 'None' if not coupled. | SEM; NanoSIMS (oxygen isotopes of carbonates, Example Isotope Laboratory) | Give one or more entries for the procedure. |  | Advanced |  | Give one or more entries for the session. |  | Editable
1.16 | Coupling Description |  | What data or context passes between the techniques, and which technique is performed first and why. Required whenever a coupled technique is named. | EPMA carbonate compositions select the matrix-matched reference material for the NanoSIMS mass-fractionation correction; SEM first, EPMA next (non-destructive), NanoSIMS last (sputtering destroys the analysed volume) | Give one entry for the procedure. |  | Advanced |  | Give one entry for the session. |  | Editable
1.17 | Coupled Procedure DOI |  | The registered procedure DOI of the coupled technique, a publication DOI, 'pending', or 'None'. | pending | — |  | N/A |  | Give one entry for the session. |  | Advanced
1.18 | Coupled Dataset or Publication Reference |  | The identifier of the dataset or publication reporting both datasets together, 'same submission', or 'pending'. | same submission | — |  | N/A |  | Give one or more entries for the session. |  | Advanced
2 | Group 2: Samples
2.1 | Target Material |  | The material types the procedure is designed to analyse. | Silicate mineral (olivine, pyroxene); oxide (magnetite); sulfide (pyrrhotite); phosphate (apatite); carbonate (dolomite, calcite) | List the material types the procedure is designed to analyse. |  | Basic |  | List the material types the procedure is designed to analyse. |  | Read-Only
2.2 | Sample Preparation Method |  | The form in which samples are presented to the instrument and the preparation that brought them to it; 'None' where material is analysed as received. | Epoxy mount, dry-polished with diamond to 0.25 µm, carbon coated to 20 nm, for CC-1; the same, plus Ar-ion polishing after the final diamond step, for CC-2 | Give one entry; each "Sample Name" (2.3) entry in a session inherits it. |  | Basic |  | Give one entry for each "Sample Name" (2.3). |  | Editable
2.3 | Sample Name |  | The name of each sample analysed in the session, as the laboratory records it. | CC-1; CC-2 | — |  | N/A |  | List every sample analysed in the session. |  | Basic
2.4 | Sampling Unit Type |  | The kind of subdivision of a sample that one row of reported results corresponds to, giving both levels where such units nest. | Map area (one map = one result; pixels are not reported individually) | Give one entry for the procedure. |  | Basic |  | Give one entry for the session. |  | Read-Only
2.5 | Sampling Unit Name |  | The label of each unit analysed and the sample it belongs to; where units are too numerous to name individually, the acquisition area they belong to. | Map 1 in CC-1; Map 2 in CC-2 | — |  | N/A |  | For each "Sample Name" (2.3), list the analysis points analysed in it. |  | Basic
2.7 | Sample Persistent Identifier |  | The persistent identifier of the sample, at the level actually analysed; IGSN recommended. | IGSN:EX0001 for CC-1; IGSN:EX0002 for CC-2 | — |  | N/A |  | Give one entry for each "Sample Name" (2.3). |  | Advanced
2.8 | Sampling Unit Selection Criteria |  | The criteria by which the units analysed are selected within a sample. | Map areas chosen to include every target phase present in a particle | Give one entry for the procedure. |  | Basic |  | Give one entry for the session. |  | Editable
2.9 | Pre-Analysis Imaging and Screening |  | Imaging or screening performed before analysis to locate and select the units analysed, with the technique, instrument and settings, and how individual analyses are linked back to the images. | BSE imaging and EDS maps (Mg, Si, Fe, Ca, S, P) on a field-emission SEM at 15 kV and 1 nA, registered to stage coordinates, each point labelled on its BSE image — for both CC-1 and CC-2 | Give one entry; each "Sample Name" (2.3) entry in a session inherits it. |  | Advanced |  | Give one entry for each "Sample Name" (2.3). |  | Editable
3 | Group 3: Instrument & Software
3.1 | Instrument Manufacturer |  | The manufacturer of the instrument. | JEOL | Give one entry for the procedure. |  | Basic |  | Give one entry for the session. |  | Read-Only
3.2 | Instrument Model |  | The model designation, including any generation or configuration suffix. | JXA-8530F Plus | Give one or more entries for the procedure. |  | Basic |  | Give one or more entries for the session. |  | Read-Only
3.3 | Electron Source |  | The type of electron source. | Schottky field-emission | Give one entry for the procedure. |  | Basic |  | Give one entry for the session. |  | Read-Only
3.4 | Acquisition Software |  | The software that controls the instrument and acquires the data, with version number. | Probe for EPMA v13 (Probe Software) | Give one entry for the procedure. |  | Basic |  | Give one entry for the session. |  | Editable
3.5 | Data Processing Software(s) |  | Every software package applied to the data after acquisition, with version numbers. | CalcImage v13 (quantitative maps); XMapTools 4 (phase classification) | Give one or more entries for the procedure. |  | Basic |  | Give one or more entries for the session. |  | Editable
3.7 | EDS Detector Configuration |  | The detector type, manufacturer, number of detector elements, active area and solid angle, window type and geometry, with multiple detectors listed separately. | Integrated silicon drift detector: 30 mm² active area, ultra-thin polymer window, 40° take-off angle | Give one or more entries for the procedure. |  | Advanced |  | Give one or more entries for the session. |  | Read-Only
4 | Group 4: Measurement Information
4.1 | Analytical Mode |  | The analytical modes the procedure covers. | EDS point analysis; EDS mapping; WDS point analysis; WDS mapping | Give one or more entries for the procedure. |  | Basic |  | Give one or more entries for the session. |  | Read-Only
4.3 | Accelerating Voltage |  | The accelerating voltage in kV, with justification for any departure from the standard operating voltage. | 15 kV (the laboratory's standard operating voltage) | Give one entry for the procedure. |  | Basic |  | Give one entry for the session. |  | Editable
4.8 | Mapping Beam Mode |  | Whether the beam was focused or defocused during X-ray mapping. | Focused, for Map 1 and Map 2 | Give one entry; each "Sampling Unit Name" (2.5) entry in a session inherits it. |  | Basic |  | For each "Sample Name" (2.3), give one entry for each of its "Sampling Unit Name" (2.5) entries. |  | Read-Only
4.9 | Mapping Beam Current |  | The probe current in nA used during X-ray mapping. | 50 nA, for Map 1 and Map 2 | Give one entry; each "Sampling Unit Name" (2.5) entry in a session inherits it. |  | Basic |  | For each "Sample Name" (2.3), give one entry for each of its "Sampling Unit Name" (2.5) entries. |  | Editable
4.10 | Mapping Beam Diameter |  | The beam diameter in µm during X-ray mapping. | 1 µm (focused), for Map 1 and Map 2 | Give one entry; each "Sampling Unit Name" (2.5) entry in a session inherits it. |  | Basic |  | For each "Sample Name" (2.3), give one entry for each of its "Sampling Unit Name" (2.5) entries. |  | Editable
4.11 | Drift Correction |  | How drift in beam current and spectrometer position is monitored and corrected during a session. | Beam current measured with the Faraday cup before each point and at the start of each map line, intensities normalised to it; primary standards re-measured at the start and end of the session, drift interpolated linearly with time | Give one entry for the procedure. |  | Advanced |  | Give one entry for the session. |  | Editable
4.12 | Target Species |  | The elements the procedure determines. Isotopes are never separate entries. | Si, Ti, Al, Cr, Fe, Mn, Mg, Ca, Na, K, P, S, F, Cl, Ni, Co, O, C (18 elements) | List the elements the procedure determines. |  | Basic |  | List the elements the procedure determines. |  | Editable
4.13 | Monitored Elements |  | The elements whose X-rays are measured, each listed under the target species it is used for. Usually the same elements as Target Species, but not always: an element determined without being measured (O by stoichiometry, C as CO₂) has no entry here, and an element measured only to correct or check another (Zn, to correct its overlap on Na Kα) is listed here but not under any target species. | Si, Mg, Fe, Ca, S | List each monitored element under the "Target Species" (4.12) it is used for; list one measured only to correct or check another on its own. |  | Basic |  | List each monitored element under the "Target Species" (4.12) it is used for; list one measured only to correct or check another on its own. |  | Read-Only
4.14 | Reported Variables and Units |  | The quantities the procedure finally reports, with their units, including intermediate quantities reported alongside final ones. | Element concentrations in wt% per pixel; phase map; modal abundances in area % | List every quantity reported, with its unit. |  | Basic |  | List every quantity reported, with its unit. |  | Read-Only
4.15 | X-ray Detection Method per Monitored Element |  | Whether the monitored element is measured by WDS or by EDS. | EDS for Si, Mg, Fe, Ca, S | Give one entry for each "Monitored Element" (4.13). |  | Basic |  | Give one entry for each "Monitored Element" (4.13). |  | Read-Only
4.16 | X-ray Line |  | The X-ray emission line measured. | Kα for Si, Mg, Fe, Ca, S | Give one entry for each "Monitored Element" (4.13). |  | Basic |  | Give one entry for each "Monitored Element" (4.13). |  | Read-Only
4.26 | EDS Acquisition Mode |  | How the beam is positioned during EDS acquisition, listing all approaches used. | Point spectra; spectrum images (full spectrum at every pixel); no line scans | Give one or more entries for the procedure. |  | Basic |  | Give one or more entries for the session. |  | Read-Only
4.27 | Dwell Time per Pixel |  | The dwell time per pixel: one value per spectrometer assignment for WDS, and a single live time per pixel for EDS. | 30 ms per pixel (EDS spectrum image, one live time for Si, Mg, Fe, Ca, S) | Give one entry for each "Monitored Element" (4.13). |  | Basic |  | Give one entry for each "Monitored Element" (4.13). |  | Editable
4.28 | Step Size / Pixel Size |  | The step size in micrometres, with X and Y given separately where they differ. | 2 µm in both X and Y | Give one entry for the procedure. |  | Basic |  | Give one entry for the session. |  | Editable
4.29 | Map Dimensions |  | The map dimensions in pixels. | 512 × 512 pixels for Map 1; 384 × 256 pixels for Map 2 | — |  | N/A |  | Give one entry for the session. |  | Basic
4.30 | Map Area |  | The physical extent of the mapped area. | 1.024 × 1.024 mm for Map 1; 0.768 × 0.512 mm for Map 2 | — |  | N/A |  | Give one entry for the session. |  | Basic
4.31 | Stage Scan vs. Beam Scan |  | Whether maps are acquired by stage scan or by beam scan. | Stage scan with the beam held fixed | Give one entry for the procedure. |  | Advanced |  | Give one entry for the session. |  | Read-Only
5 | Group 5: Data Processing
5.1 | Matrix Correction Method |  | The matrix correction algorithm, which also applies when count maps are converted to concentration maps. | Armstrong/Love–Scott φ(ρz) matrix correction (Armstrong 1995), applied pixel by pixel in CalcImage | Give one entry for the procedure. |  | Basic |  | Give one entry for the session. |  | Read-Only
5.2 | Mass Absorption Coefficients (MACs) |  | The database of mass absorption coefficients the correction uses. | FFAST (Chantler et al. 2005) | Give one entry for the procedure. |  | Basic |  | Give one entry for the session. |  | Read-Only
5.3 | X-ray Background Correction Method |  | How the background beneath the peak is estimated and subtracted. | Removed by filter fitting for Si, Mg, Fe, Ca, S | Give one entry for each "Monitored Element" (4.13). |  | Basic |  | Give one entry for each "Monitored Element" (4.13). |  | Editable
5.4 | Time-Dependent Intensity Correction |  | The time-dependent intensity correction applied. | Linear for Na; exponential for F (in apatite); none for other elements | Give one entry for each "Target Species" (4.12). |  | Advanced |  | Give one entry for each "Target Species" (4.12). |  | Editable
5.5 | Target Species Estimation Method |  | How the concentration of the element is obtained from the measured data. | Directly from measured intensities for every measured element; by stoichiometry from cation valences for O; as CO₂ by stoichiometry for C (carbonates only); OH in apatite by difference, F + Cl + OH = 1 apfu (Ketcham 2015) | Give one entry for each "Target Species" (4.12). |  | Advanced |  | Give one entry for each "Target Species" (4.12). |  | Read-Only
5.6 | Halogen Correction on Oxygen |  | Whether oxygen is adjusted for halogen substitution where oxygen is calculated by stoichiometry. | Oxygen equivalent of F and Cl subtracted from the total (apatite) | Give one entry for the procedure. |  | Advanced |  | Give one entry for the session. |  | Editable
5.8 | EDS Spectral Processing Type |  | How net peak intensities are extracted from the spectra before quantification. | Filter fitting: a top-hat filter followed by least-squares fitting to measured standard spectra | Give one entry for the procedure. |  | Advanced |  | Give one entry for the session. |  | Read-Only
5.9 | Blank Correction |  | The method and reference material used to determine and subtract a blank. | Apparent concentration measured on Ni- and Co-free synthetic forsterite subtracted, for Ni and Co; none for other elements | Give one entry for each "Target Species" (4.12). |  | Advanced |  | Give one entry for each "Target Species" (4.12). |  | Editable
5.10 | Normalization / Standards-Based Correction |  | Any normalisation applied beyond the primary calibration, or 'None'. | None | Give one entry for each "Reported Variable" (4.14). |  | Advanced |  | Give one entry for each "Reported Variable" (4.14). |  | Editable
5.11 | Calibration Factor and Determination Method |  | Any externally calibrated factor that converts the measured quantity into the reported quantity, how that factor was determined, and its uncertainty. | None, for every reported variable (concentrations from k-ratios against the primary standards) | Give one entry for each "Reported Variable" (4.14). |  | Basic |  | Give one entry for each "Reported Variable" (4.14). |  | Editable
5.12 | Procedural Blank Level |  | The blank level measured in the session, together with the blank's composition where the reported quantity is a ratio. | 35 ± 9 µg/g for Ni; 22 ± 8 µg/g for Co (n = 5) | — |  | N/A |  | Give one entry for each "Target Species" (4.12). |  | Basic
5.13 | Analysis Inclusion and Rejection Criteria |  | The rules deciding which individual results contribute to a combined value: the limits a result must meet and the grounds on which one is rejected. In the dataset: flag each excluded result, with the rule that excluded it. | None: maps are not combined | Give one entry for the procedure. |  | Basic |  | Give one entry for the session. |  | Editable
5.14 | Combination Method |  | Whether the reported variable is an individual result or combines several; if combined, the statistical model used and what the results are combined over. In the dataset: report each combined value with its uncertainty and number of results, and, for each result, the combined value it contributes to. | Not combined: maps are reported pixel by pixel | Give one entry for each "Reported Variable" (4.14). |  | Basic |  | Give one entry for each "Reported Variable" (4.14). |  | Editable
5.15 | Constants and Reference Values Used |  | The physical constants and citable reference values used in data reduction, with their sources, or 'None'. | IUPAC 2021 standard atomic weights (Prohaska et al. 2022) | Give one entry for the procedure. |  | Basic |  | Give one entry for the session. |  | Editable
6 | Group 6: QC & Uncertainty
6.1 | Primary Calibration Standard Name |  | The primary reference material against which the instrument is calibrated, its source, and a citation for the accepted values used; where quantification uses a stored library or response model, that instead. 'None' means no calibration was performed. | Springwater olivine for Si, Mg; Rockport fayalite for Fe; diopside for Ca; dolomite for Ca, Mg in carbonates; synthetic FeS for S (Smithsonian values: Jarosewich et al. 1980) | For each "Target Material" (2.1), give one entry for each "Target Species" (4.12). |  | Basic |  | For each "Target Material" (2.1), give one entry for each "Target Species" (4.12). |  | Editable
6.2 | Secondary Reference Materials |  | The reference materials measured as unknowns alongside the samples, with the source of each and a citation for its accepted values. | San Carlos olivine; Kakanui augite; Durango apatite; chromite; calcite (accepted values: Jarosewich et al. 1980) | List the secondary reference materials run as unknowns. |  | Basic |  | List the secondary reference materials run as unknowns. |  | Editable
6.3 | X-ray Line Overlap Corrections Applied |  | Whether an X-ray line overlap correction is applied and, if so, the overlapping lines and the correction method. | No for Si, Mg, Fe, Ca, S | Give one entry for each "Target Species" (4.12). |  | Basic |  | Give one entry for each "Target Species" (4.12). |  | Read-Only
6.4 | Interfering Elements |  | The elements whose lines overlap the measured peak. | None | Give one entry for each "Target Species" (4.12). |  | Advanced |  | Give one entry for each "Target Species" (4.12). |  | Read-Only
6.5 | Interference Correction Standard |  | The reference material used to calibrate the interference correction. | None required | Give one entry for each "Target Species" (4.12). |  | Advanced |  | Give one entry for each "Target Species" (4.12). |  | Read-Only
6.6 | Detection Limit |  | The detection limit, with units, stating whether the value is a typical estimate or a session-specific measurement. Must be consistent with the stated method. | About 0.7 wt% per pixel for the EDS elements | Give one entry for each "Reported Variable" (4.14). |  | Advanced |  | Give one entry for each "Reported Variable" (4.14). |  | Basic
6.7 | Detection Limit Method |  | The formula or approach used to calculate the detection limit, with a citation where one exists. | 3σ of the background counts, for every reported variable | Give one entry for each "Reported Variable" (4.14). |  | Basic |  | Give one entry for each "Reported Variable" (4.14). |  | Read-Only
6.8 | Analytical Precision |  | The reproducibility of repeated measurements, as a one-sigma relative standard deviation in percent, with the number of analyses and the measured value. | 0.4–0.7% for oxides above 10 wt%; 0.9–1.6% for FeO and Al₂O₃; 2–4% for TiO₂, Na₂O, F and Cl; 4% for NiO in San Carlos olivine (1σ RSD; San Carlos olivine n = 9, Kakanui augite n = 6, Durango apatite n = 6, chromite n = 5, calcite n = 6) | For each "Secondary Reference Material" (6.2), give one entry for each "Reported Variable" (4.14). |  | Advanced |  | For each "Secondary Reference Material" (6.2), give one entry for each "Reported Variable" (4.14). |  | Basic
6.9 | Analytical Accuracy |  | The offset between measured and accepted values, as a percent relative bias, with the source of the accepted value. | Within ±1.2% relative for all major oxides; within ±2.6% for minor oxides and halogens | For each "Secondary Reference Material" (6.2), give one entry for each "Reported Variable" (4.14). |  | Advanced |  | For each "Secondary Reference Material" (6.2), give one entry for each "Reported Variable" (4.14). |  | Basic
6.10 | Counting Statistics Error |  | The uncertainty predicted from counting statistics, including the counts on any background or blank subtracted, with the sigma level stated. | 1σ from counting statistics on the peak and on any subtracted background or blank, for every analysis and every reported variable; in olivine about 0.5% relative for SiO₂ and MgO, about 10% for NiO at 0.03 wt% | Give one entry for each "Reported Variable" (4.14). |  | Advanced |  | For each "Sample Name" (2.3), give one entry for every "Reported Variable" (4.14) at every "Sampling Unit Name" (2.5) entry. |  | Basic
6.11 | EDS Dead Time |  | The percent dead time recorded by the detector. | 24–31% during spectrum imaging | — |  | N/A |  | Give one entry for the session. |  | Basic
6.12 | Goodness-of-Fit or Dispersion Statistic |  | The statistic used to describe the scatter among the results contributing to a combined value, or to test whether it exceeds what analytical uncertainty predicts (for example one standard deviation of the mean, or the MSWD). In the dataset: report its value for each combined value, and any further statistic beside it, named. | MSWD of the included points about the mean, for every oxide and element concentration; none for structural formulas, which are not combined | Give one entry for each "Reported Variable" (4.14). |  | Basic |  | Give one entry for each "Reported Variable" (4.14). |  | Editable

## Registration Metadata
# | Metadata item | Needed? | Description | Example (NOT recommendations) | Structure — procedure record | Structure OK? | Procedure-Level Tier | Tier OK? | Structure — session record | Structure OK? | Session-Level Tier | Tier OK? | Other comments
1 | Group 1: Procedure Identification
1.1 | Procedure Name |  | A short descriptive name with a version number, identifying the instrument, the technique and the scope of what is measured. | EPMA-CC v2: major and minor elements in carbonaceous-chondrite minerals by WDS and EDS | Give one entry for the procedure. |  | Basic |  | Give one entry for the session. |  | Read-Only
1.2 | Technique |  | Whether the procedure uses WDS, EDS or both. | EPMA-WDS+EDS (combines wavelength-dispersive and energy-dispersive spectrometry) | Give one entry for the procedure. |  | Basic |  | Give one entry for the session. |  | Read-Only
1.3 | Procedure Author |  | The person or laboratory responsible for the procedure; ORCID recommended for a person. | Example Microanalysis Laboratory | Give one entry for the procedure. |  | Basic |  | Give one entry for the session. |  | Read-Only
1.4 | Laboratory |  | The laboratory or institution hosting the instrument. | Example Microanalysis Laboratory | Give one entry for the procedure. |  | Basic |  | Give one entry for the session. |  | Editable
1.5 | Laboratory ID |  | The laboratory's persistent identifier. | https://ror.org/0example00 | Give one entry for the procedure. |  | Advanced |  | Give one entry for the session. |  | Editable
1.6 | Procedure Start Date |  | The date this procedure configuration first went into production use. | January 2025 | Give one entry for the procedure. |  | Basic |  | Give one entry for the session. |  | Read-Only
1.7 | Funding Source for Procedure Development |  | The funding that supported the instrument, major upgrades and development of the procedure, with agencies and grant numbers. | Example Agency grant A (instrument); grant B (procedure development) | Give one entry for the procedure. |  | Advanced |  | Give one entry for the session. |  | Read-Only
1.8 | Procedure Reference(s) |  | Publications or technical reports describing, validating or benchmarking the procedure. | Example et al. (2025) | Give one entry for the procedure. |  | Advanced |  | Give one entry for the session. |  | Read-Only
1.9 | Procedure DOI |  | The DOI of the registered procedure followed, or 'pending'. | doi:10.0000/epma-cc.v2 | — |  | N/A |  | Give one entry for the session. |  | Basic
1.10 | Session Identifier |  | The laboratory's own run, sequence or batch identifier for the session, as generated by the instrument or acquisition software. | Run 2026-04-12 | — |  | N/A |  | Give one entry for the session. |  | Basic
1.11 | Analyst |  | The analyst or analysts who performed the session; ORCID recommended. | Analyst A (ORCID 0000-0000-0000-0000) | — |  | N/A |  | Give one entry for the session. |  | Basic
1.12 | Analysis Start Date |  | The date the session began. | 2026-04-12 | — |  | N/A |  | Give one entry for the session. |  | Basic
1.13 | Analysis End Date |  | The date the session ended. | 2026-04-14 | — |  | N/A |  | Give one entry for the session. |  | Basic
1.14 | Funding Source for Analysis |  | The funding that supported these analyses, with agencies and grant numbers. | Example Agency grant C | — |  | N/A |  | Give one entry for the session. |  | Basic

## Structure key (traceability) (hidden)
# | Metadata item | TAPP "Keyed By" (Column I) | Structure — procedure record | Structure — session record
1.1 | Procedure Name | (none) | Give one entry for the procedure. | Give one entry for the session.
1.2 | Technique | (none) | Give one entry for the procedure. | Give one entry for the session.
1.3 | Procedure Author | (none) | Give one entry for the procedure. | Give one entry for the session.
1.4 | Laboratory | (none) | Give one entry for the procedure. | Give one entry for the session.
1.5 | Laboratory ID | (none) | Give one entry for the procedure. | Give one entry for the session.
1.6 | Procedure Start Date | (none) | Give one entry for the procedure. | Give one entry for the session.
1.7 | Funding Source for Procedure Development | (none) | Give one entry for the procedure. | Give one entry for the session.
1.8 | Procedure Reference(s) | (none) | Give one entry for the procedure. | Give one entry for the session.
1.9 | Procedure DOI | (none) | — | Give one entry for the session.
1.10 | Session Identifier | (none) | — | Give one entry for the session.
1.11 | Analyst | (none) | — | Give one entry for the session.
1.12 | Analysis Start Date | (none) | — | Give one entry for the session.
1.13 | Analysis End Date | (none) | — | Give one entry for the session.
1.14 | Funding Source for Analysis | (none) | — | Give one entry for the session.
1.15 | Coupled Technique(s) | (none) | Give one or more entries for the procedure. | Give one or more entries for the session.
1.16 | Coupling Description | (none) | Give one entry for the procedure. | Give one entry for the session.
1.17 | Coupled Procedure DOI | (none) | — | Give one entry for the session.
1.18 | Coupled Dataset or Publication Reference | (none) | — | Give one or more entries for the session.
2.1 | Target Material | defines: target material | List the material types the procedure is designed to analyse. | List the material types the procedure is designed to analyse.
2.2 | Sample Preparation Method | sample | Give one entry; each "Sample Name" (2.3) entry in a session inherits it. | Give one entry for each "Sample Name" (2.3).
2.3 | Sample Name | defines: sample | — | List every sample analysed in the session.
2.4 | Sampling Unit Type | (none) | Give one entry for the procedure. | Give one entry for the session.
2.5 | Sampling Unit Name | defines: sample > sampling unit | — | For each "Sample Name" (2.3), list the analysis points analysed in it.
2.7 | Sample Persistent Identifier | sample | — | Give one entry for each "Sample Name" (2.3).
2.8 | Sampling Unit Selection Criteria | (none) | Give one entry for the procedure. | Give one entry for the session.
2.9 | Pre-Analysis Imaging and Screening | sample | Give one entry; each "Sample Name" (2.3) entry in a session inherits it. | Give one entry for each "Sample Name" (2.3).
3.1 | Instrument Manufacturer | (none) | Give one entry for the procedure. | Give one entry for the session.
3.2 | Instrument Model | (none) | Give one or more entries for the procedure. | Give one or more entries for the session.
3.3 | Electron Source | (none) | Give one entry for the procedure. | Give one entry for the session.
3.4 | Acquisition Software | (none) | Give one entry for the procedure. | Give one entry for the session.
3.5 | Data Processing Software(s) | (none) | Give one or more entries for the procedure. | Give one or more entries for the session.
3.7 | EDS Detector Configuration | (none) | Give one or more entries for the procedure. | Give one or more entries for the session.
4.1 | Analytical Mode | (none) | Give one or more entries for the procedure. | Give one or more entries for the session.
4.3 | Accelerating Voltage | (none) | Give one entry for the procedure. | Give one entry for the session.
4.8 | Mapping Beam Mode | sample > sampling unit | Give one entry; each "Sampling Unit Name" (2.5) entry in a session inherits it. | For each "Sample Name" (2.3), give one entry for each of its "Sampling Unit Name" (2.5) entries.
4.9 | Mapping Beam Current | sample > sampling unit | Give one entry; each "Sampling Unit Name" (2.5) entry in a session inherits it. | For each "Sample Name" (2.3), give one entry for each of its "Sampling Unit Name" (2.5) entries.
4.10 | Mapping Beam Diameter | sample > sampling unit | Give one entry; each "Sampling Unit Name" (2.5) entry in a session inherits it. | For each "Sample Name" (2.3), give one entry for each of its "Sampling Unit Name" (2.5) entries.
4.11 | Drift Correction | (none) | Give one entry for the procedure. | Give one entry for the session.
4.12 | Target Species | defines: target species | List the elements the procedure determines. | List the elements the procedure determines.
4.13 | Monitored Elements | defines: monitored property per target species | List each monitored element under the "Target Species" (4.12) it is used for; list one measured only to correct or check another on its own. | List each monitored element under the "Target Species" (4.12) it is used for; list one measured only to correct or check another on its own.
4.14 | Reported Variables and Units | defines: reported property | List every quantity reported, with its unit. | List every quantity reported, with its unit.
4.15 | X-ray Detection Method per Monitored Element | monitored property | Give one entry for each "Monitored Element" (4.13). | Give one entry for each "Monitored Element" (4.13).
4.16 | X-ray Line | monitored property | Give one entry for each "Monitored Element" (4.13). | Give one entry for each "Monitored Element" (4.13).
4.26 | EDS Acquisition Mode | (none) | Give one or more entries for the procedure. | Give one or more entries for the session.
4.27 | Dwell Time per Pixel | monitored property | Give one entry for each "Monitored Element" (4.13). | Give one entry for each "Monitored Element" (4.13).
4.28 | Step Size / Pixel Size | (none) | Give one entry for the procedure. | Give one entry for the session.
4.29 | Map Dimensions | (none) | — | Give one entry for the session.
4.30 | Map Area | (none) | — | Give one entry for the session.
4.31 | Stage Scan vs. Beam Scan | (none) | Give one entry for the procedure. | Give one entry for the session.
5.1 | Matrix Correction Method | (none) | Give one entry for the procedure. | Give one entry for the session.
5.2 | Mass Absorption Coefficients (MACs) | (none) | Give one entry for the procedure. | Give one entry for the session.
5.3 | X-ray Background Correction Method | monitored property | Give one entry for each "Monitored Element" (4.13). | Give one entry for each "Monitored Element" (4.13).
5.4 | Time-Dependent Intensity Correction | target species | Give one entry for each "Target Species" (4.12). | Give one entry for each "Target Species" (4.12).
5.5 | Target Species Estimation Method | target species | Give one entry for each "Target Species" (4.12). | Give one entry for each "Target Species" (4.12).
5.6 | Halogen Correction on Oxygen | (none) | Give one entry for the procedure. | Give one entry for the session.
5.8 | EDS Spectral Processing Type | (none) | Give one entry for the procedure. | Give one entry for the session.
5.9 | Blank Correction | target species | Give one entry for each "Target Species" (4.12). | Give one entry for each "Target Species" (4.12).
5.10 | Normalization / Standards-Based Correction | reported property | Give one entry for each "Reported Variable" (4.14). | Give one entry for each "Reported Variable" (4.14).
5.11 | Calibration Factor and Determination Method | reported property | Give one entry for each "Reported Variable" (4.14). | Give one entry for each "Reported Variable" (4.14).
5.12 | Procedural Blank Level | target species | — | Give one entry for each "Target Species" (4.12).
5.13 | Analysis Inclusion and Rejection Criteria | (none) | Give one entry for the procedure. | Give one entry for the session.
5.14 | Combination Method | reported property | Give one entry for each "Reported Variable" (4.14). | Give one entry for each "Reported Variable" (4.14).
5.15 | Constants and Reference Values Used | (none) | Give one entry for the procedure. | Give one entry for the session.
6.1 | Primary Calibration Standard Name | target material x target species | For each "Target Material" (2.1), give one entry for each "Target Species" (4.12). | For each "Target Material" (2.1), give one entry for each "Target Species" (4.12).
6.2 | Secondary Reference Materials | defines: standard | List the secondary reference materials run as unknowns. | List the secondary reference materials run as unknowns.
6.3 | X-ray Line Overlap Corrections Applied | target species | Give one entry for each "Target Species" (4.12). | Give one entry for each "Target Species" (4.12).
6.4 | Interfering Elements | target species | Give one entry for each "Target Species" (4.12). | Give one entry for each "Target Species" (4.12).
6.5 | Interference Correction Standard | target species | Give one entry for each "Target Species" (4.12). | Give one entry for each "Target Species" (4.12).
6.6 | Detection Limit | reported property | Give one entry for each "Reported Variable" (4.14). | Give one entry for each "Reported Variable" (4.14).
6.7 | Detection Limit Method | reported property | Give one entry for each "Reported Variable" (4.14). | Give one entry for each "Reported Variable" (4.14).
6.8 | Analytical Precision | standard x reported property | For each "Secondary Reference Material" (6.2), give one entry for each "Reported Variable" (4.14). | For each "Secondary Reference Material" (6.2), give one entry for each "Reported Variable" (4.14).
6.9 | Analytical Accuracy | standard x reported property | For each "Secondary Reference Material" (6.2), give one entry for each "Reported Variable" (4.14). | For each "Secondary Reference Material" (6.2), give one entry for each "Reported Variable" (4.14).
6.10 | Counting Statistics Error | sample > sampling unit x reported property | Give one entry for each "Reported Variable" (4.14). | For each "Sample Name" (2.3), give one entry for every "Reported Variable" (4.14) at every "Sampling Unit Name" (2.5) entry.
6.11 | EDS Dead Time | (none) | — | Give one entry for the session.
6.12 | Goodness-of-Fit or Dispersion Statistic | reported property | Give one entry for each "Reported Variable" (4.14). | Give one entry for each "Reported Variable" (4.14).