## TAPP
Metadata Item | Description / Purpose | Procedure-Level Tier | Analysis-Level Tier | Data Type | Example / Allowed Content | Comments | Last Update | Keyed By | TEM Sample Preparation | 3D Tomography | Literature Assessment | Garvie et al. 2008 ¦ Tagish Lake (C2) nanoglobules ¦ TEM Sample Preparation (FIB, FEI Nova 200 NanoLab) | Liu et al. 2017 ¦ High-rank coal (Qinshui basin) ¦ 3D Tomography (Carl Zeiss Crossbeam 540) | Zhou et al. 2017 ¦ Coal (SC + HBC, Junggar Basin) ¦ 3D Tomography (FEI Helios NanoLab 650) | Zega et al. 2025 ¦ Bennu asteroid particles (OSIRIS-REx) ¦ TEM Sample Preparation (Helios G3, U Arizona) | Zega et al. 2025 ¦ Bennu asteroid particles (OSIRIS-REx) ¦ TEM Sample Preparation (Helios G4 UX, UC Berkeley) | Zega et al. 2025 ¦ Bennu asteroid particles (OSIRIS-REx) ¦ TEM Sample Preparation (Quanta3D600, NASA JSC) | Barnes et al. 2025 ¦ Bennu asteroid particles (OSIRIS-REx) ¦ TEM Sample Preparation (FEI Helios G4 DualBeam, NASA JSC) | Barnes et al. 2025 ¦ Bennu asteroid particles (OSIRIS-REx) ¦ TEM Sample Preparation (FEI Helios 660 G3, NASA JSC)
1. Procedure Identification | [8 empty] | N | N | N | N | N | N | N | N | N | N | N
Procedure Name | A short descriptive name for this analytical procedure, including a version number. Should identify the instrument, the technique, and the scope of what is measured — analyte, material, or feature. | Basic | Read-Only | Text (free) | e.g., 'SEM-BSE/EDS Mineral Phase Mapping, JEOL 7600F, Bennu Astromaterials' ¦ 'FIB-SEM 3D Tomography, Helios NanoLab 660, Organic-Rich Chondrite Matrix' | Source: Core module | 2026-07-24 | (none) | Y | Y |  | N | N | N | N | N | N | N | N
Technique | Top-level analytical technique identifier. | Basic | Read-Only | Controlled list | SEM ¦ SEM/FIB-SEM ¦ Other: specify ¦ N/A ¦ None | Source: Core module | 2026-05-21 | (none) | Y | Y |  | N | N | N | N | N | N | N | N
Procedure Author | Person(s) or laboratory responsible for developing and registering this procedure. ORCID recommended for individuals. | Basic | Read-Only | Text (free) | e.g., 'T.J. McCoy, Smithsonian NMNH' ¦ 'ARES Electron Beam Analysis Laboratories, NASA JSC' | Source: Core module | 2026-07-24 | (none) | Y | Y |  | N | N | N | N | N | N | N | N
Laboratory | Name of the laboratory or institution hosting the instrument. | Basic | Editable | Text (free) | e.g., 'Electron Beam Analysis Laboratories, ARES, NASA Johnson Space Center' ¦ 'K-ALFAA, University of Arizona' | Source: Core module | 2026-05-21 | (none) | Y | Y |  | School of Earth and Space Exploration / School of Materials, Arizona State University | China University of Mining and Technology, Xuzhou, China | China University of Geosciences, Beijing, China | K-ALFAA (Kuiper-Arizona Laboratory for Astromaterials Analysis), University of Arizona | Molecular Foundry, Lawrence Berkeley National Laboratory (UC Berkeley) | NASA Johnson Space Center (JSC), Houston, TX, USA | N | N
Laboratory ID | Persistent identifier for the laboratory (e.g., ROR ID). | Advanced | Editable | URI / DOI | e.g., 'https://ror.org/027ka1x80' | Source: Core module | 2026-05-21 | (none) | Y | Y |  | N | N | N | N | N | N | N | N
Procedure Start Date | First date this procedure configuration was used in production. | Basic | Read-Only | Date | YYYY-MM-DD | Source: Core module | 2026-07-24 | (none) | Y | Y |  | N | N | N | N | N | N | N | N
Funding Source for Procedure Development | Grants and other funding sources that supported instrument acquisition, major upgrades, procedure development, and associated personnel time. Include grant numbers and funding agencies where applicable. | Advanced | Read-Only | Text (free) | e.g., 'NSF EAR-1234567' ¦ 'NASA 80NSSC20K0371' ¦ 'None' | Source: Core module | 2026-07-24 | (none) | Y | Y |  | N | N | N | N | N | N | N | N
Procedure Reference(s) | DOI or URL for peer-reviewed publications or technical reports describing, validating, or benchmarking this procedure. | Advanced | Read-Only | URI / DOI | e.g., 'https://doi.org/10.1234/example' | Source: Core module | 2026-07-24 | (none) | Y | Y |  | N | N | N | N | N | N | N | N
Procedure DOI | DOI or unique persistent identifier for the registered procedure used in this analysis, generated upon procedure submission to a data repository. This field is mandatory when submitting analytical data, to encourage formal procedure registration and ensure traceability. If a DOI has been applied for but not yet minted, enter "pending". | N/A | Basic | URI / DOI | e.g., 'https://doi.org/10.xxxxx/procedure-id' ¦ 'pending' | Source: Core module | 2026-07-24 | (none) | Y | Y |  | N | N | N | N | N | N | N | N
Session Identifier | Identifier for the analytical session this record describes — the laboratory's own run, sequence or batch identifier as generated by the instrument or acquisition software. The analysis record corresponds to one session, which may cover several samples, and this is the link back to the raw instrument files. Distinct from any persistent identifier a repository mints on submission. | N/A | Basic | Text (free) |  | Source: Core module | 2026-08-13 | (none) | Y | Y
Analyst | Name(s) of the analyst(s) who performed the analysis session. ORCID is recommended for persistent identification. | N/A | Basic | Text (free) | e.g., 'T.J. McCoy' ¦ 'Sara Singerling' | Source: Core module | 2026-05-21 | (none) | Y | Y |  | N | N | N | N | N | N | N | N
Analysis Start Date | Date on which the analytical session began. For sessions spanning multiple days, use the date of the first session. Format: YYYY-MM-DD. | N/A | Basic | Date | YYYY-MM-DD | Source: Core module | 2026-05-21 | (none) | Y | Y |  | N | N | N | N | N | N | N | N
Analysis End Date | Date on which the analytical session ended. May equal Analysis Start Date for single-day sessions. Format: YYYY-MM-DD. | N/A | Basic | Date | YYYY-MM-DD | Source: Core module | 2026-05-21 | (none) | Y | Y |  | N | N | N | N | N | N | N | N
Funding Source for Analysis | Grants and other funding sources that supported the specific sample analyses and data collection reported here. Include grant numbers and funding agencies. Distinct from procedure development funding, which may have been awarded years earlier. | N/A | Basic | Text (free) | e.g., 'NSF EAR-2345678' ¦ 'NASA 80NSSC23K0415' ¦ 'None' | Source: Core module | 2026-05-21 | (none) | Y | Y |  | NASA NNG06GE37G (LAJG); NASA NNG06GF08G (PRB) | N | N | NASA PSEF 80NSSC23K0327; NASA Planetary Major Equipment NNX12AL47G; NSF MRI 0619599 | US DOE contract DE-AC02-05CH11231 (Advanced Light Source / Molecular Foundry) | NASA award NNH09ZDA007O; contract NNM10AA11C (OSIRIS-REx New Frontiers) | N | N
Coupled Technique(s) | Other analytical techniques applied to the same sample(s) whose results are intended to be interpreted together with data from this procedure. Document coupling with any technique whose results are functionally linked to this dataset — providing calibration inputs, complementary spatial context, or required companion measurements. Use the same controlled vocabulary as the Technique field. Enter "None" if no coupling is intended. | Advanced | Editable | Controlled list / Text | e.g., 'TEM/STEM (FIB sections)' ¦ 'EPMA; NanoSIMS' ¦ 'XCT (pre-SEM overview)' ¦ 'None' ¦ N/A | Source: Core module | 2026-08-11 | (none) | Y | Y |  | SE Imaging (same instrument, pre- and post-FIB milling) | X-ray CT (Xradia 520 Versa, Carl Zeiss) | N | TEM analysis (HF5000 STEM, U Arizona); BSE/SE Imaging (Hitachi S-4800, U Arizona) | Synchrotron XANES (ALS, Berkeley); TEM analysis | TEM analysis (NASA JSC); BSE/SE Imaging (JEOL 7600F, JSC) | N | N
Coupling Description | Description of how this procedure is coupled with the technique(s) listed above. Include: (1) the functional relationship — what data or context flows between techniques, or how results are combined (e.g. which output from the coupled technique serves as input to data reduction for this technique); and (2) the analytical sequence — which technique is performed first and why (e.g. non-destructive before destructive). Required when Coupled Technique(s) is not "None". | Advanced | Editable | Text (free) | e.g., 'SEM-BSE/EDS used for phase identification and target selection before FIB lift-out; FIB performed after SEM survey because lift-out is destructive. TEM-EDS provides nanoscale chemistry of FIB-prepared sections.' | Source: Core module | 2026-05-21 | (none) | Y | Y |  | N | N | N | N | N | N | N | N
Coupled Procedure DOI | Registered procedure DOI for the coupled technique named above. Provides a stable, citable link to the companion method independent of whether a dataset has been deposited. If the coupled procedure has not yet been registered, enter the DOI of a publication describing the coupled method, or "pending". Enter "None" if no coupling is planned. | N/A | Advanced | URI / DOI | e.g., 'https://doi.org/10.xxxxx/coupled-procedure' ¦ 'pending' ¦ 'None' | Source: Core module | 2026-07-24 | (none) | Y | Y |  | N | N | N | N | N | N | N | N
Coupled Dataset or Publication Reference | DOI or other persistent identifier for the co-registered dataset or publication where both datasets are reported together. Accepts: a dedicated dataset DOI (if separately deposited), a shared dataset DOI (if co-submitted in the same package), or a publication DOI. Use "same submission" if the coupled dataset is included in this data package, or "pending" if not yet assigned. If coupling is documented through a shared sample identifier only, that information is already captured in Sample Persistent Identifier (Group 2). | N/A | Advanced | URI / DOI / Text | e.g., 'https://doi.org/10.xxxxx/dataset' ¦ 'same submission' ¦ 'pending' ¦ 'None' | Source: Core module | 2026-05-21 | (none) | Y | Y |  | N | N | N | N | N | N | N | N
2. Samples | [8 empty] | N | N | N | N | N | N | N | N | N | N | N
Target Material | General description of the material type(s) this procedure is designed to analyse. Used for discoverability and procedure matching, and because the material type constrains sample preparation, calibration and matrix-matching requirements. | Basic | Read-Only | Controlled list / Text | Silicate mineral ¦ Oxide ¦ Sulfide ¦ Carbonate ¦ Phosphate ¦ Metal alloy ¦ Organic matter ¦ Glass ¦ Regolith ¦ Whole rock / polished section ¦ Porous material ¦ Other: specify ¦ N/A ¦ None | Source: Core module | 2026-08-14 | (none) | Y | Y |  | Carbonaceous nanoglobules, Tagish Lake (C2) meteorite (HCl/HF acid residue concentrate) | High-rank coal (anthracite from Bofang Mine; lean coal from Yuwu Mine), southern Qinshui basin, China | Subbituminous coal (SC) and high-volatile bituminous coal (HBC), Xishanyao Formation, southern Junggar Basin, NW China | Asteroid (101955) Bennu particles (OSIRIS-REx sample return); fine-grained matrix areas | Asteroid (101955) Bennu particles (OSIRIS-REx sample return) | Asteroid (101955) Bennu particles (OSIRIS-REx sample return) | N | N
Sample Preparation Method | Method by which samples were prepared for SEM analysis prior to loading in the instrument. Includes mounting medium (epoxy, carbon tape, stub), polishing steps (alumina, colloidal silica, argon ion mill), and conductive coating type and thickness. For VP-SEM/ESEM analyses, note whether an uncoated sample was used and the gas type used. FIB-specific in-session operations (protective coating deposition, milling conditions, lamella preparation) are documented separately in Group 4. | Basic | Editable | Text (free) | e.g., 'Polished thin section, 10 nm C coat' ¦ 'Fragments on C tape, uncoated (VP-SEM, 30 Pa H2O)' ¦ '25 mm epoxy mount, 0.25 µm diamond polish, 5 nm Ir coat' |  | 2026-05-22 | (none) | Y | Y |  | HCl and HF acid dissolution residue from pristine Tagish Lake pieces; deposited on lacey C TEM grid attached to Al-SEM stub; uncoated | Small coal pillars (~2 mm diameter, 2 mm height) drilled orthogonal to bedding; polished with cross section polisher to remove ~1-2 µm oxide layer; no coating applied | Cuboidal samples (~0.5×1×1 cm) polished with dry emery paper and argon ion polishing; high-pressure freezing and freeze-drying; glued onto sample holder | Sections extracted from fine-grained matrix areas in polished sections; BSE and SE images acquired before and after sectioning | Particles placed on PELCO carbon conductive tabs on Al SEM round; no protective coating stated | Particles dispersed on conductive carbon dots on Al SEM pin mounts | N | N
Sample Name | Name or identifier of each sample analysed in this session, as used in the laboratory — a sample mount, section or aliquot counts as one entry where that is what the laboratory tracks. The analysis record corresponds to one session and may cover several samples; fields whose Keyed By column declares 'sample' take one value per entry. Should match the identifier used in associated publications or data tables. | N/A | Basic | Text (free) | e.g., 'Bennu OREX-800045-103' ¦ 'NWA 8657 polished thin section 1' | Source: Core module | 2026-08-12 | defines: sample | Y | Y |  | N | N | SC; HBC | N | N | N | N | N
Sampling Unit | The physical subdivision of the sample to which one row of reported values corresponds — the unit that is analysed and reported, as distinct from the sample as a whole. State the unit type at procedure level and the units actually analysed at analysis level. Where units nest (e.g. confined tracks within grains), state both levels. | Basic | Basic | Controlled list / Text | Whole sample ¦ Aliquot ¦ Grain ¦ Spot ¦ Analysis point ¦ Phase ¦ Sub-volume ¦ Region of interest ¦ N/A ¦ None ¦ Other: specify | Source: Core module | 2026-08-11 | defines: sampling unit | Y | Y
Sample Persistent Identifier | Globally unique, persistent identifier for each sample listed in Sample Name. IGSN (International Geo Sample Number) is the recommended standard for geological and cosmochemical samples, as used by Astromat, EarthChem and SESAR. Where a sample and its sub-samples are separately registered, record the identifier at the level actually analysed. | Advanced | Advanced | URI / IGSN | e.g., 'IGSN:AU1234567' | Source: Core module | 2026-08-12 | sample | Y | Y |  | N | N | N | N | N | N | N | N
Target Selection Criteria | The rules governing which part of the sample is analysed, and why. Covers the criteria applied when choosing grains, aliquots, spots, or a region of interest — size, morphology, clarity, freedom from inclusions or alteration, phase identity, or spatial position. Distinct from Target Material, which states the material type the procedure is designed for: this field states how, within such a sample, the analysed portion is picked out. | Basic | Editable | Text (free) |  | Source: Target Selection module | 2026-08-08 | (none) | Y | Y
Pre-Analysis Imaging and Screening | Imaging or other characterisation performed before the measurement in order to select or locate the analysed target, including the technique, instrument and settings used, and how individual analyses are linked back to the images. Distinct from any imaging the procedure performs as its own measurement. Where the imaging is performed on a separate instrument, it should also be recorded in the Group 1 coupling fields. | Advanced | Editable | Text (free) |  | Source: Target Selection module | 2026-08-08 | sample | Y | Y
3. Instrument & Software | [8 empty] | N | N | N | N | N | N | N | N | N | N | N
Instrument Manufacturer | Manufacturer of the instrument that performs the measurement, recorded as a controlled value so that procedures can be found by vendor. Where a procedure couples a sample-introduction system to an analysing instrument, this records the analysing instrument. Instrument Model gives the specific designation. | Basic | Read-Only | Controlled list | JEOL ¦ Zeiss ¦ FEI / Thermo Fisher Scientific ¦ Hitachi ¦ Tescan ¦ Phenom ¦ Other: specify ¦ Unknown ¦ N/A ¦ None | Source: Core module | 2026-08-12 | (none) | Y | Y |  | FEI | Carl Zeiss | FEI (now Thermo Fisher Scientific) | Thermo Fisher Scientific (ThermoScientific) | FEI (now Thermo Fisher Scientific) | FEI (now Thermo Fisher Scientific) | N | N
Instrument Model | Model designation of the instrument that performs the measurement, including any generation or configuration suffix. Conventionally written with the manufacturer name included; Instrument Manufacturer records the vendor separately, as a controlled value, so that procedures remain findable by vendor. | Basic | Read-Only | Text (free) | e.g., 'JEOL JSM-7600F' ¦ 'Zeiss Supra 55 FEG-SEM' ¦ 'FEI Quanta 650 FEG' ¦ 'Thermo Fisher Quattro S' ¦ 'FEI Helios NanoLab 660' ¦ 'Zeiss Crossbeam 540' | Source: Core module | 2026-05-21 | (none) | Y | Y |  | Nova 200 NanoLab DualBeam | Crossbeam 540 | Helios NanoLab 650 | Helios G3 | Helios G4 UX | Quanta 3D 600 | N | N
Electron Source | Type of electron gun used in the instrument. | Basic | Read-Only | Controlled list | Cold-FEG ¦ Schottky FEG ¦ LaB6 / CeB6 ¦ Tungsten (W) ¦ Other: specify ¦ Unknown ¦ N/A ¦ None |  | 2026-08-12 | (none) | Y | Y |  | Other: FEG (type not specified in paper) | N | N | Field emission gun (FEG) | Field emission gun (FEG) | Field emission gun (FEG) | N | N
Instrument Variant | Broad platform type of the instrument. 'Standard SEM': dedicated electron-only SEM column. 'FIB-SEM dual-beam': combined focused ion beam and SEM columns (enables TEM specimen preparation, 3D serial sectioning, ion-beam milling). 'VP-SEM / ESEM': variable-pressure or environmental SEM for imaging uncoated, hydrated, or charging specimens. An instrument may combine categories (e.g., FIB-SEM with VP capability). | Basic | Read-Only | Controlled list | Standard SEM ¦ FIB-SEM dual-beam ¦ VP-SEM / ESEM ¦ FIB-SEM dual-beam + VP ¦ Other: specify ¦ N/A ¦ None |  | 2026-05-21 | (none) | Y | Y |  | FIB-SEM dual-beam | N | Dual Beam (FIB-SEM) | Dual Beam (FIB-SEM) | Dual Beam (FIB-SEM) | Dual Beam (FIB-SEM) | N | N
Ion Beam Source | Ion beam source type in a FIB-SEM dual-beam system. Gallium LMIS (Ga+) is the most common ion source; xenon plasma FIB (PFIB) provides higher material removal rates for large-volume milling; helium and neon ion microscopes (GFIS) provide nanometre-resolution imaging and low-damage milling. | Basic | Read-Only | Controlled list | Gallium LMIS (Ga+) ¦ Xenon plasma FIB (PFIB) ¦ Helium GFIS ¦ Neon GFIS ¦ Other: specify ¦ N/A ¦ None |  | 2026-08-11 | (none) | Y | Y |  | Gallium LMIS (Ga+) | N | N | N | Gallium (Ga+) ion beam | Gallium (Ga+) ion beam | N | N
SE Detector Type | Type of secondary electron detector used. Everhart-Thornley (ET) detector is the standard off-axis collector sensitive to SE2 and some BSE; in-lens (TLD) detectors collect high-resolution SE1 signal at short working distances; GSED/ESED detectors operate in VP/ESEM mode by using the chamber gas as the signal amplification medium. | Basic | Read-Only | Controlled list | Everhart-Thornley (ET) ¦ In-lens / TLD (through-the-lens) ¦ GSED (VP/ESEM) ¦ ESED (VP/ESEM) ¦ Other: specify ¦ N/A ¦ None |  | 2026-05-22 | (none) | Y | Y |  | N | N | N | N | N | N | N | N
BSE Detector Type | Type of backscattered electron detector. Solid-state diode detectors (single or segmented) are standard; YAG scintillator detectors offer high sensitivity at low voltage. Segmented detectors can operate in composition mode (segments summed) or topography mode (differential signal between segments). In-lens BSE detectors provide improved BSE collection at short working distances. | Basic | Read-Only | Controlled list | Solid-state diode (single) ¦ Solid-state diode (segmented, composition mode) ¦ Solid-state diode (segmented, topography mode) ¦ In-lens BSE ¦ YAG scintillator ¦ Other: specify ¦ N/A ¦ None |  | 2026-05-22 | (none) | Y | Y |  | N | N | N | N | N | N | N | N
Acquisition Software | Instrument control, automation and data acquisition software used to collect the raw data, including version number. Distinct from Data Processing Software(s), which covers everything applied to the data after acquisition. | Basic | Editable | Text (free) | e.g., 'JEOL Analysis Station v3.5' ¦ 'Oxford AZtec 6.0' ¦ 'Zeiss SmartSEM 6.06' ¦ 'Thermo Fisher Auto Slice and View G4 (3D tomography)' | Source: Core module | 2026-08-14 | (none) | Y | Y |  | N | N | N | N | N | N | N | N
Data Processing Software(s) | All software applied to the data after acquisition in order to produce the reported quantities, including version numbers. List every package used. Distinct from Acquisition Software, which controls the instrument and collects the raw data. | Basic | Editable | Text (free) | e.g., 'Oxford AZtec 6.0 (EDS quantification); Oxford AZtec Crystal 2.1 (EBSD); Gatan DigitalMicrograph 3.5 (CL); FIJI 2.15 (image processing); Avizo 2023 (3D segmentation)' | Source: Core module | 2026-08-14 | (none) | N | Y |  | N | Avizo 7 (3D digital core software); Multiple-point geostatistics for pore network model | Fiji/ImageJ (StackReg/TurboReg for slice alignment; VolumeJ for volume rendering); Adobe Photoshop CS6 (image enhancement); FEI Avizo Fire 8.1.1 (pore volume reconstruction and segmentation) | N/A | N/A | N/A | N/A | N/A
4. Measurement Information | [8 empty] | N | N | N | N | N | N | N | N | N | N | N
Analytical Mode | Primary analytical mode(s) executed under this procedure. Records one value for a single-mode procedure and all applicable modes for a multi-mode one. Serves as the procedure-level declaration of measurement type, distinct from the mode flag columns, which state per-field applicability. Required even where a procedure defines no mode flag columns, so that every procedure record is self-describing and comparable across the library. | Basic | Read-Only | Controlled list | 'TEM Sample Preparation' ¦ '3D Tomography' ¦ 'TEM Sample Preparation; 3D Tomography' | Source: Core module | 2026-08-24 | (none) | Y | Y |  | TEM Sample Preparation | 3D Tomography | 3D Tomography | TEM Sample Preparation | TEM Sample Preparation | TEM Sample Preparation | TEM Sample Preparation | TEM Sample Preparation
Reported Variables and Units | The final variable(s) this procedure reports and their units — distinct from the fields recording what was *acquired* rather than what is reported. A procedure may acquire many channels and report a small number of derived quantities; without this field a data consumer cannot tell which. Record every reported variable, including intermediate quantities reported alongside final ones (e.g. both the 206Pb/238U ratio and the 206Pb/238U date). Where a reported variable is a nominal property with no magnitude (e.g. a mineral identification), record the variable and give the unit as 'N/A — nominal property'. | Basic | Read-Only | Text (free) | e.g., 'Element concentration (ppm); oxide (wt%)' ¦ 'd56Fe (permil vs IRMM-014); d57Fe (permil vs IRMM-014)' ¦ '206Pb/238U date (Ma); 207Pb/206Pb date (Ma)' ¦ 'Mineral identity (N/A — nominal property); match score (0-1)' | Source: Core module | 2026-08-12 | defines: reported property | Y | Y
Accelerating Voltage | Electron beam accelerating voltage in kilovolts. Affects X-ray generation depth (EDS/WDS), EBSD pattern quality, imaging resolution, and beam penetration. Low voltages (1-5 kV) improve surface sensitivity and reduce beam damage; high voltages (15-20 kV) improve X-ray generation for quantitative analysis. | Basic | Editable | Numeric (kV) | e.g., 2 ¦ 5 ¦ 10 ¦ 15 ¦ 20 |  | 2026-05-21 | (none) | Y | Y |  | N | N | 2 kV (SEM imaging) | 30 keV (FIB milling and thinning) | 16 to 30 keV (coarse milling); down to 1 keV (polishing) | 30 kV (initial milling); 16 kV (intermediate); 5 kV (final thinning) | N | N
Beam Current | Electron beam probe current. Higher current improves signal-to-noise for X-ray analysis (EDS/WDS, EBSD) and CL but may increase beam damage and reduce spatial resolution. Express in nA; for sub-nA values use decimal notation (e.g., 0.4 nA). | Basic | Editable | Numeric (nA) | e.g., 0.4 ¦ 0.9 ¦ 2 ¦ 5 ¦ 10 ¦ 20 |  | 2026-08-11 | (none) | Y | Y |  | N | N | N | 0.8 to 2.5 nA (thinning) | N | N | N | N
Working Distance | Distance between the objective lens pole piece and the specimen surface in millimetres. Affects spatial resolution, depth of focus, EDS X-ray take-off angle, and EBSD geometry. | Basic | Editable | Numeric (mm) | e.g., 4 ¦ 8 ¦ 10 ¦ 15 |  | 2026-05-21 | (none) | Y | Y |  | 5.4 mm (eucentric height for electron and ion columns) | N | 4 mm | N | N | N | N | N
Dwell Time per Pixel | Time the electron beam dwells on each pixel during raster scanning (imaging modes) or on each step position during compositional mapping (EDS and WDS mapping modes), in microseconds or milliseconds. Longer dwell time improves signal-to-noise and counting statistics but increases total dose and can cause beam damage or contamination on sensitive materials. For WDS mapping, the dwell time is per spectrometer per pixel. | Basic | Editable | Numeric + unit | e.g., '400 ns' ¦ '10 µs' ¦ '1 ms' |  | 2026-06-01 | (none) | N | Y |  | N/A | N | N | N/A | N/A | N/A | N/A | N/A
Protective Coating Deposition | Type and deposition conditions of the protective coating applied to the sample surface before FIB milling. E-beam deposition causes less surface damage than ion-beam deposition and should be applied as the initial layer. Typical coatings: platinum (Pt) or carbon (C). State material, deposition method, beam conditions, and approximate thickness. | Basic | Editable | Text (free) | e.g., 'E-beam Pt (5 kV, 1.5 nA, ~500 nm) then Ga+ ion-beam Pt (30 kV, 0.3 nA, ~2 µm)' ¦ 'E-beam C (5 kV) + ion-beam C (2-3 µm total)' |  | 2026-08-11 | (none) | Y | Y |  | N | No coating applied; not sputtered with gold or other materials | N | 12-µm wide × 4-µm tall carbon capping layer deposited on matrix areas | N | Electron-beam deposited carbon (~0.5–1 µm); followed by ion beam-deposited carbon (~2–3 µm capping layer) | N | N
Coarse Milling Conditions | Ion beam voltage and current used for bulk material removal during FIB milling. For TEM specimen preparation: bulk trenching to isolate the lamella and intermediate thinning. For 3D tomography: face preparation and initial slice removal. | Basic | Editable | Text (free) | e.g., '30 kV, 2.5 nA (bulk trenching); 16 kV, 0.8 nA (intermediate thinning)' |  | 2026-08-11 | (none) | Y | Y |  | 30 kV, 10 pA Ga beam | N | N | Standard stair step; 30 keV, currents 2.5 to 0.8 nA | Ga+ ion beam at 16–30 keV (coarse milling) | Ga+ ion beam at 30 kV (initial milling); 16 kV (intermediate) | N | N
Fine Polishing Conditions | Ion beam voltage and current for final thinning and surface polishing of the TEM lamella. Low-voltage polishing (2 kV or below) minimises Ga implantation depth, surface amorphisation, and curtaining artifacts. | Basic | Editable | Text (free) | e.g., '5 kV, 0.3 nA; final polish: 2 kV, 72 pA' ¦ '1 kV, 0.1 nA final step' |  | 2026-08-11 | (none) | Y | N |  | N | N/A | N/A | N | Various voltages down to 1 keV (polishing) | Ga+ ion beam at 5 kV (final thinning) until ~100 nm thick | N | N
Foil Thickness | Target thickness of the electron-transparent TEM lamella after final FIB polishing, in nanometres. Actual thickness may differ from target. Typical range: 50-150 nm for standard TEM/STEM; 200-600 nm for XANES or tomography sections. | Basic | Editable | Numeric + unit | e.g., '~100 nm' ¦ '50-80 nm' ¦ '100-200 nm' |  | 2026-08-11 | (none) | Y | N |  | N | N/A | N/A | N | <100 to 600 nm (variable, depending on targeted experiment) | ~100 nm | N | N
Lift-out Method | Method used to transfer the FIB-prepared lamella to the TEM support grid. In-situ lift-out uses a micromanipulator inside the FIB-SEM chamber and is the standard method for small or precious specimens. | Basic | Read-Only | Controlled list | In-situ (micromanipulator, Cu or Mo TEM half-grid) ¦ Ex-situ ¦ Other: specify ¦ N/A ¦ None |  | 2026-08-11 | (none) | Y | N |  | N | N/A | N/A | In-situ lift-out (standard stair step) | N | In-situ lift-out; ion beam-deposited Pt weld to Cu half grids | N | N
Ion Milling Conditions (3D Tomography) | Ion beam voltage and current used to mill each slice during FIB-SEM serial sectioning. These parameters determine material removal rate per slice and exposed surface quality. | Basic | Editable | Text (free) | e.g., 'Ga+ 30 kV, 2.5 nA (slice milling)' |  | 2026-08-11 | (none) | N | Y |  | N/A | N | N | N/A | N/A | N/A | N/A | N/A
Slice Thickness | Thickness of each FIB-milled slice during serial sectioning in nanometres. Controls the Z-axis resolution of the 3D reconstruction. | Basic | Editable | Numeric (nm) | e.g., 10 ¦ 20 ¦ 50 ¦ 100 |  | 2026-08-11 | (none) | N | Y |  | N/A | 15 nm (single layer scanning thickness = 9.0 µm total / 600 slices) | N | N/A | N/A | N/A | N/A | N/A
Voxel Size and Image Stack Dimensions | X, Y, Z dimensions of the reconstructed 3D voxel in nanometres (X-Y pixel size from SEM image calibration; Z from slice thickness), and the total number of slices in the stack. Determined at analysis time. | N/A | Basic | Text (free) | e.g., '14.6 x 8.3 x 10.0 nm, 600 slices' ¦ '14.8 x 14.8 x 20 nm, 800 slices' |  | 2026-08-11 | (none) | N | Y |  | N/A | 9.8 × 9.8 × 15 nm voxel size; 600 slices; 7.8 × 7.8 µm scanning area; 9.0 µm total thickness | 14.8×14.8 nm pixel size (XY); ~800 total slices; sub-volumes: SC=5.609×3.08×5.446 µm; HBC=4.679×3.2×4.24 µm; SEM image resolution 2.5 nm | N/A | N/A | N/A | N/A | N/A
5. Data Processing | [8 empty] | N | N | N | N | N | N | N | N | N | N | N
3D Image Registration | Method used to align consecutive SEM image slices in the 3D stack to correct for drift, vibration, and curtaining artifacts. Include software used. | Basic | Editable | Text (free) | e.g., 'Cross-correlation (FIJI/ImageJ)' ¦ 'Fiducial marker alignment' ¦ 'Auto Slice and View G3 built-in drift correction' ¦ 'Avizo drift correction' |  | 2026-08-11 | (none) | N | Y |  | N/A | N | Fiji/ImageJ StackReg and TurboReg plugins used for slice realignment | N/A | N/A | N/A | N/A | N/A
3D Segmentation Method | Method and software used to segment phases and features in the aligned 3D image stack, transforming the grayscale stack into labelled 3D regions (pores, mineral phases, grain boundaries, organic matter). | Basic | Editable | Text (free) | e.g., 'Greyscale thresholding (Avizo Fire 8.1)' ¦ 'Machine learning (Dragonfly 2022)' ¦ 'Manual tracing' ¦ 'Region growing' ¦ 'None' |  | 2026-08-11 | (none) | N | Y |  | N/A | Image denoising, binarization, and segmentation; 3D model established using Avizo 7 and Multiple-point geostatistics | Semi-automatic porosity segmentation by grayscale thresholding; pore volume reconstruction using FEI Avizo Fire 8.1.1; connected component analysis for pore network extraction (PNE) | N/A | N/A | N/A | N/A | N/A
Constants and Reference Values Used | Physical constants and reference values used in data reduction to calculate the final reported quantity (e.g., decay constants for age calculation, standard isotope ratios, or other citable reference values used in a correction or calculation), together with their source. Distinct from the Group 6 reference-material fields, which document accepted values for specific calibration/validation materials rather than universal physical constants. Record "None" if no citable, revisable physical constants feed into this procedure's data reduction. | Basic | Editable | Text (free) | e.g., "λ238U = 1.55125×10⁻¹⁰ yr⁻¹, λ235U = 9.8485×10⁻¹⁰ yr⁻¹ (Jaffey et al. 1971); 238U/235U = 137.818 (Hiess et al. 2012)" ¦ "None" | Source: Core module | 2026-08-12 | (none) | Y | Y
6. Quality Control & Uncertainty | [8 empty] | N | N | N | N | N | N | N | N | N | N | N
Additional Notes | Any procedure- or analysis-specific information not captured by a structured field anywhere in this TAPP — including anomalies, deviations from the registered procedure, instrument modifications, and supplementary context. Scope is the whole document, not Group 6: this is the last field of the TAPP and covers all six groups. Use sparingly; a structured field is preferred for anything that can be formally categorised. | Advanced | Advanced | Text (free) | Free text | Source: Core module | 2026-08-11 | (none) | Y | Y |  | N | N | Stage tilt 52° between electron and ion columns; SEM range 20V–30kV and FIB range 500V–30kV (system specs); destriping filter xStripes.jar applied; deconvolves y-axis by y/sin(52°) for pixel scale correction | Sections thinned to electron transparency; BSE/SE images acquired before and after sectioning; methods follow refs. 72-75 | Thicker sections (<100 nm) for TEM; sections up to 600 nm for Fe-L XANES and tomography | Multi-step milling: e-beam C deposition then FIB C capping, 30 kV → 16 kV → 5 kV; Pt weld to Cu half grids | No FIB-based TEM sample preparation at NASA JSC described in this paper (Barnes et al. 2025). This paper does not include TEM analysis or FIB-SEM-based sample preparation. FIB-based TEM foil preparation using FEI Helios G4 and related instruments at NASA JSC is described in the companion paper Zega et al. 2025 (Nat. Geosci., ref. 7 therein). | No FIB-based TEM sample preparation at NASA JSC described in this paper (Barnes et al. 2025). This paper does not include TEM analysis or FIB-SEM-based sample preparation. TEM foil preparation using FEI Helios 660 G3 and related instruments at NASA JSC is described in the companion paper Zega et al. 2025 (Nat. Geosci., ref. 7 therein).

## Legends
Procedure-Level Tier | Description | Analysis-Level Tier | Description
Basic | Mandatory for procedure registration. Must be provided to register a valid procedure. | Read-Only | Directly imported from the registered procedure; cannot be changed by the analyst. Changing this value means running a different procedure. Also used for fields relevant only at procedure level — the value is inherited from the procedure record and shown read-only in the analysis form.
Advanced | Optional for procedure registration. Strongly recommended but not required. | Editable | Imported from the registered procedure but may be adjusted within procedure-defined bounds (e.g., daily tuning, minor software updates). The procedure registers the target or typical value; the analyst confirms or adjusts it. Cannot be left void if the procedure-level tier is Basic.
N/A | Not applicable at procedure level. This field captures analysis-level information only. | Basic | Mandatory user input at analysis time. Value comes from the analysis itself and cannot be pre-specified in the procedure.
|  | Advanced | Optional user input at analysis time. Recommended for complete and reproducible documentation.
Mode Column | Definition
TEM Sample Preparation
3D Tomography
Y | This field applies to this analytical mode and should be reported.
N | This field does not apply to this analytical mode.
Keyed By (Rule 7) | What the field's value repeats over
(none) | Scalar — one value per procedure or analysis. The default.
defines: reported property | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
defines: sample | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it.
defines: sampling unit | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per subdivision of the physical sample that carries its own row — grain, spot, aliquot, phase, sub-volume.
sample