## TAPP
Metadata Item | Description | Procedure-Level Tier | Analysis-Level Tier | Data Type | Example / Allowed Content | Comments | Last Update | Keyed By | Purpose | Literature Assessment | Desem+etal2022 ¦ Nu Attom SC-SF-ICP-MS ¦ Univ Melbourne | Li+etal2016 ¦ Thermo Element I ¦ IGGCAS Beijing | Lu+etal2007 ¦ Finnigan ELEMENT ¦ PML Okayama | Milne+etal2010 ¦ Thermo Finnigan Element I ¦ FSU NHMFL | Misra+etal2014 ¦ Thermo Element XR ¦ Univ Cambridge | Willbold2005 ¦ ThermoFinnigan ELEMENT2 ¦ MPI Mainz
1. Procedure Identification | [9 empty] | N | N | N | N | N | N | N
Procedure Name | A short descriptive name for this analytical procedure, including a version number. Should identify the instrument, the technique, and the scope of what is measured — target species, material, or feature. | Basic | Read-Only | Text (free) |  | Source: Core module | 2026-09-01 | (none) |  |  | N | N | N | N | N | N
Technique | Top-level analytical technique identifier. | Basic | Read-Only | Controlled list | Solution SF-ICP-MS | Source: Core module | 2026-06-16 | (none) |  |  | Solution SF-ICP-MS | Solution SF-ICP-MS | Solution SF-ICP-MS | Solution SF-ICP-MS | Solution SF-ICP-MS | Solution SF-ICP-MS
Procedure Author | Person(s) or laboratory responsible for developing and registering this procedure. ORCID recommended for individuals. | Basic | Read-Only | Text (free) |  | Source: Core module | 2026-07-24 | (none) |  |  | N | N | N | N | N | N
Laboratory | Name of the laboratory or institution hosting the instrument. | Basic | Editable | Text (free) |  | Source: Core module | 2026-06-16 | (none) |  |  | University of Melbourne, Australia (affiliation) | Institute of Geology and Geophysics, Chinese Academy of Sciences (IGGCAS), Beijing (affiliation) | Pheasant Memorial Laboratory (PML), Okayama University (section 2.1.2) | National High Magnetic Field Laboratory (NHMFL), Florida State University (stated section 2.4) | Godwin Laboratory for Palaeoclimate Research, University of Cambridge (affiliation) | Max-Planck-Institut fuer Chemie (MPIC), Mainz, Germany (affiliation)
Laboratory ID | Persistent identifier for the laboratory (e.g., ROR ID). | Advanced | Editable | URI / DOI |  | Source: Core module | 2026-06-16 | (none) |  |  | N | N | N | N | N | N
Procedure Start Date | First date this procedure configuration was used in production. | Basic | Read-Only | Date |  | Source: Core module | 2026-07-24 | (none) |  |  | N | N | N | N | N | N
Funding Source for Procedure Development | Grants and other funding sources that supported instrument acquisition, major upgrades, procedure development, and associated personnel time. Include grant numbers and funding agencies where applicable. | Advanced | Read-Only | Text (free) |  | Source: Core module | 2026-07-24 | (none) |  |  | N | N | N | N | N | N
Procedure Reference(s) | DOI or URL for peer-reviewed publications or technical reports describing, validating, or benchmarking this procedure. | Advanced | Read-Only | URI / DOI |  | Source: Core module | 2026-07-24 | (none) |  |  | N | N | N | N | N | N
Procedure DOI | DOI or unique persistent identifier for the registered procedure used in this analysis, generated upon procedure submission to a data repository. If a DOI has been applied for but not yet minted, enter "pending". | N/A | Basic | URI / DOI |  | Source: Core module | 2026-08-25 | (none) | Encourages formal procedure registration and ensures traceability. |  | N/A | N/A | N/A | N/A | N/A | N/A
Session Identifier | Identifier for the analytical session this record describes — the laboratory's own run, sequence or batch identifier as generated by the instrument or acquisition software. Distinct from any persistent identifier a repository mints on submission. | N/A | Basic | Text (free) |  | Source: Core module | 2026-08-25 | (none) | The analysis record corresponds to one session, which may cover several samples, and this is the link back to the raw instrument files. |  | N -- "A typical session comprised analyses of up to 50 unknowns and 15 standards"; no session identifier stated | N | N | N -- "Each analytical session would begin and end with the analysis of a series of Mo standards (1-100 nM)"; "an analysis sequence"; "1 day's analysis"; "three separate days of analyses". No session or sequence identifier stated | N -- "a single instrument session" referenced; no session identifier stated | N
Analyst | Name(s) of the analyst(s) who performed the analysis session. ORCID is recommended for persistent identification. | N/A | Basic | Text (free) |  | Source: Core module | 2026-08-25 | (none) |  |  | N/A | N/A | N/A | N/A | N/A | N/A
Analysis Start Date | Date on which the analytical session began. For sessions spanning multiple days, use the date of the first session. Format: YYYY-MM-DD. | N/A | Basic | Date |  | Source: Core module | 2026-08-25 | (none) |  |  | N/A | N/A | N/A | N/A | N/A | N/A
Analysis End Date | Date on which the analytical session ended. May equal Analysis Start Date for single-day sessions. Format: YYYY-MM-DD. | N/A | Basic | Date |  | Source: Core module | 2026-08-25 | (none) |  |  | N/A | N/A | N/A | N/A | N/A | N/A
Funding Source for Analysis | Grants and other funding sources that supported the specific sample analyses and data collection reported here. Include grant numbers and funding agencies. Distinct from procedure development funding, which may have been awarded years earlier. | N/A | Basic | Text (free) |  | Source: Core module | 2026-08-25 | (none) |  |  | N/A | N/A | N/A | N/A | N/A | N/A
Coupled Technique(s) | Other analytical techniques applied to the same sample(s) whose results are intended to be interpreted together with data from this procedure. Document coupling with any technique whose results are functionally linked to this dataset — providing calibration inputs, complementary spatial context, or required companion measurements. Name the instrument, laboratory or purpose alongside each technique where known, and join multiple techniques with '; '. Enter "None" if no coupling is intended. | Advanced | Editable | Text (free) | e.g., 'EPMA (Brown University CAMECA SX-100)' ¦ 'FIB-SEM (Crossbeam 540) for 3D tomography; EDS for mineral analysis' ¦ 'NanoSIMS; STXM; Raman' ¦ 'None' ¦ 'N/A' | Source: Core module | 2026-08-31 | (none) |  |  | Agilent 7700x Q-ICP-MS (used for Pb isotope comparison; stated section 2.4) | N | Agilent 7500cs Q-ICP-MS at PML (B, Zr, Nb, Mo, Sn, Sb, Hf, Ta on same samples; stated section 2.1.1) | N | N | N
Coupling Description | Description of how this procedure is coupled with the technique(s) listed above. Include: (1) the functional relationship — what data or context flows between techniques, or how results are combined (e.g. which output from the coupled technique serves as input to data reduction for this technique); and (2) the analytical sequence — which technique is performed first and why (e.g. non-destructive before destructive). Required when Coupled Technique(s) is not "None". | Advanced | Editable | Text (free) |  | Source: Core module | 2026-08-25 | (none) |  |  | Q-ICP-MS Pb isotope ratios compared with SF-ICP-MS values for validation; Q-ICP-MS showed higher uncertainties (stated section 2.4) | N | The SF-ICP-MS measures Ti as the (47Ti + 49Ti)/93Nb ratio; the Q-ICP-MS measures B, Zr, Nb, Mo, Sn, Sb, Hf and Ta on the same solutions and gives the Nb concentration used for Ti — abstract; §2.7 | N | N | N
Coupled Procedure DOI | Registered procedure DOI for the coupled technique named above. If the coupled procedure has not yet been registered, enter the DOI of a publication describing the coupled method, or "pending". Enter "None" if no coupling is planned. | N/A | Advanced | URI / DOI |  | Source: Core module | 2026-08-25 | (none) | Provides a stable, citable link to the companion method independent of whether a dataset has been deposited. |  | N/A | N/A | N/A | N/A | N/A | N/A
Coupled Dataset or Publication Reference | DOI or other persistent identifier for the co-registered dataset or publication where both datasets are reported together. Accepts: a dedicated dataset DOI (if separately deposited), a shared dataset DOI (if co-submitted in the same package), or a publication DOI. Use "same submission" if the coupled dataset is included in this data package, or "pending" if not yet assigned. If coupling is documented through a shared sample identifier only, that information is already captured in Sample Persistent Identifier (Group 2). | N/A | Advanced | URI / DOI / Text |  | Source: Core module | 2026-08-25 | (none) |  |  | N/A | N/A | N/A | N/A | N/A | N/A
2. Samples | [9 empty] | N | N | N | N | N | N | N
Target Material | General description of the material type(s) this procedure is designed to analyse. | Basic | Read-Only | Controlled list / Text | Basalt ¦ Chondrite ¦ Peridotite ¦ Synthetic solution ¦ N/A ¦ None | Source: Core module | 2026-09-28 | defines: target material | Used for discoverability and procedure matching, and because the material type constrains sample preparation, calibration and matrix-matching requirements. |  | Soil samples (sequential acid leach fractions; stated abstract) | magnetite; pyrite — mineral separates from alternating magnetite-rich and pyrite-rich mesobands, Wutai greenstone belt; FER-2 is the reference material | basalt; andesite; peridotite; carbonaceous chondrite — GSJ and USGS silicate reference materials, and four carbonaceous chondrites (Ivuna, Orgueil, Cold Bokkeveld, Allende) | Open-ocean seawater (stated abstract) | Foraminifera calcite (benthic and planktonic species; stated abstract) | basalt; andesite; granite; rhyolite; shale; peridotite; synthetic glass — the rock types of the 17 reference materials of Table 2, including the USGS and MPI-DING glasses
Sample Name | Name or identifier of each sample analysed in this session, as used in the laboratory — a sample mount, section or aliquot counts as one entry where that is what the laboratory tracks. Should match the identifier used in associated publications or data tables. | N/A | Basic | Text (free) | e.g., 'Allende MS-1 digest A' ¦ 'NWA 8657 aliquot 2' ¦ 'BHVO-2 replicate 3' | Source: Core module | 2026-08-25 | defines: sample | The analysis record corresponds to one session and may cover several samples. |  | Soil and rock samples from boreholes BH1, BH2 (Sunbury), BH3, BH4 (Kalkallo), BH5 (Greenvale), BH6, BH (Wallan), incl. BH3a; reference materials BCR-2, BR, AGV-2, JB-2, JB-3, NIST SRM981, and Broken Hill Main Lode galena | mag_1, mag_3, mag_5, py_2, py_4; iron-formation reference material FER-2 (CCRMP, CANMET MMSL, Canada) | JB-1, JB-2, JB-3, JA-1, JA-2, JA-3, JP-1 (GSJ); BHVO-1, AGV-1, PCC-1, DTS-1 (USGS); Ivuna (CI1), Orgueil (CI1), Cold Bokkeveld (CM2), Allende (USNM 3529, Split 1, Pos. 23) | Open-ocean seawater reference materials SAFe S1, SAFe D2 and NASS-5; GEOTRACES inter-calibration samples GS (surface) and GD (deep); depth-profile samples from the BATS station, 31 deg 45' N, 64 deg 05' W, 23 June 2008 | In-house consistency standards CAM-wuellerstorfi, CAM-Uvig-1, CAM-Uvig-2 and CAM-Mix; Globigerinoides sacculifer specimens of the 300-355 um size fraction | AGV-1, AGV-2, BCR-1, BCR-2, BCR-2G, BIR-1, BIR-1G, BHVO-1, BHVO-2, BHVO-2G, G-2, JR-1, KL2-G, ML3B-G, NIST SRM 612, OU-6, PCC-1 -- tabulated with issuing organisation and split/position numbers (e.g. BHVO-1 Split 15 Pos 26; G-2 Split 58 Pos 23)
Sampling Unit Type | The kind of physical subdivision of the sample to which one row of reported values corresponds — the unit that is analysed and reported, as distinct from the sample as a whole. Where units nest (e.g. confined tracks within grains), state both levels. The units themselves are listed in Sampling Unit Name. | Basic | Read-Only | Controlled list / Text | Whole sample ¦ Aliquot ¦ Grain ¦ Spot ¦ Analysis point ¦ Phase ¦ Sub-volume ¦ Region of interest ¦ N/A ¦ None | Source: Core module | 2026-09-15 | (none) |  |  | Weighed split of a digest or leachate -- rock chips 0.05-0.24 g, soils 1-2.3 g; "weighed splits taken for trace element and high-precision Pb isotope analysis by MC-ICPMS. At least 50% of each solution was retained for Pb isotope analysis by SC-SF-ICP-MS and Q-ICP-MS"; "Small splits of the soil samples (TD, AR) were used for Pb isotope analysis on a Nu Instruments Attom" | Aliquot of the digest solution -- 50 mg FER-2 and "approximately 100 mg of the studied mineral samples" digested; "a small aliquot sample solution was taken for column separation", "7.2 mg Fe in 10% aliquot of magnetite solution"; "A 1.8 g sample solution (in 2 g of 10 M HCl) was weighed and loaded" | Weighed test portion -- "Approximately 20 mg of basalt and andesite samples were weighed"; "Approximately 50 mg for peridotites and approximately 10 mg for meteorites"; 9-18 mg for carbonaceous chondrites | 12 mL sub-sample (aliquot) of an acidified seawater sample -- "Acidified seawater samples ... were sub-sampled (12 mL) into clean 30 mL FEP Teflon bottles. The 12 mL aliquots were spiked"; "standard additions ... were added to individual 12 mL sub-samples of the same sample"; "Standard additions of Co and Mn were performed on a further four aliquots (1 mL) of the elution acid" | Dissolved foraminiferal test aliquot -- "capable of analyzing small masses of calcite (5-10 mg), including single foraminifera specimens"; "Leached samples were dissolved in a minimum volume of 1 M HNO3 (40-60 uL) ... centrifuged for 2 min at 10,000 rpm and the supernatant was used for Me/Ca analysis. A 5 uL aliquot ..." | Digestion, with determinations nested inside it -- "Five independent analyses (different spikings/digestions) of BHVO-1 were carried out over a time period of 4 months. Triplicate determinations were performed for each digestion"; "Only one digestion was prepared for the USGS reference glasses ... and were measured in triplicate"
Sampling Unit Name | The name or label of each sampling unit analysed in this session, as the laboratory records it, together with the sample it belongs to — e.g. a spot number, a grain label, a map or region-of-interest name, or an aliquot identifier. Where units are too numerous to name individually, such as map pixels or reconstructed voxels, name the acquisition area they belong to instead. | N/A | Basic | Text (free) | e.g., NWA 8657-1: spots 1–24 ¦ Zircon mount ZM-3: grains Z1–Z30 ¦ Core 73001: map areas A1–A5 ¦ Bennu aggregate OREX-803015-100: aliquot for Ti | Source: Core module | 2026-09-15 | defines: sample > sampling unit | Gives every field keyed by sampling unit a row to attach its values to, and ties each unit to the sample it belongs to; without it those fields collapse to one value per session. |  | Labelled by digestion: "rock TD, soil TD, soil AR splits" of each borehole sample (p.3), e.g. sample "BH3a" (p.2); splits are not numbered | Sample name only — "mag_1", "mag_3", "mag_5", "py_2", "py_4", each reported as "Mean ± s (n = 3)" with the replicates unlabelled (table, p.8) | Sample name only, except the Allende powder: "the Smithsonian reference Allende powder (USNM 3529, Split 1, Pos. 23)" (p.5). Solutions "#1"–"#8" (p.7) are synthetic yield-test solutions, not samples | Sample name only — seawater samples by name (SAFe S1, SAFe D2, NASS-5; GEOTRACES "(GS)" and "(GD)", p.4); the 12 mL sub-samples are not labelled | Sample name only — consistency standards "CAM-Uvig-1", "CAM-Uvig-2", "CAM-wuellerstorfi", "CAM-Mix" (p.6); in the cleaning test each core-top sample was "split into three fractions" identified only by cleaning sequence (p.4) | Labelled for BHVO-1: digestions "BHVO-1 (1)" to "BHVO-1 (5)", each with determinations "1 2 3" (Table 4, pp.8–9); the other reference materials by name with their issuing split and position (e.g. AGV-1 Split 35 Pos 13; Table 2, p.6)
Sample Persistent Identifier | Globally unique, persistent identifier for each sample listed in Sample Name. IGSN (International Geo Sample Number) is the recommended standard for geological and cosmochemical samples, as used by Astromat, EarthChem and SESAR. Where a sample and its sub-samples are separately registered, record the identifier at the level actually analysed. | N/A | Advanced | URI / IGSN |  | Source: Core module | 2026-09-15 | sample |  |  | N/A | N/A | N/A | N/A | N/A | N/A
Sample Description | Brief description of sample provenance, form, or preparation state relevant to this analysis. | N/A | Basic | Text (free) |  |  | 2026-06-16 | (none) |  |  | N/A | N/A | N/A | N/A | N/A | N/A
Sample Preparation Method | The form in which the sample is presented to the instrument, and the preparation that brought it to that form — for example mounting, sectioning, polishing, coating, crushing, fusion, or extraction of an electron-transparent section. Record 'None' where the material is analysed as received. | Basic | Editable | Controlled list / Text | Whole-rock powder ¦ Mineral separate ¦ Biogenic material (cleaned) ¦ Natural water ¦ Leachate or partial-dissolution fraction ¦ Dissolved residue ¦ Pre-digested liquid ¦ N/A ¦ None | Source: Core module | 2026-09-28 | sample |  |  | Soils oven-dried (50 °C, 7 days), not sieved; total-dissolution (TD) and aqua-regia (AR) splits; the SC-SF-ICP-MS splits 'redissolved and diluted to 2 ml in high purity 2% HNO3 doped with 1 ppb of high-purity thallium' — §2.1, §2.2, §2.4; the rock chips were analysed by MC-ICP-MS only | Mineral separates (~100 mg powder; stated section 2.3) | Whole-rock powder (decomposed in TFM bomb; same as Q-ICP-MS portion; stated section 2.1.1) | Open-ocean seawater; off-line pre-concentration using Toyopearl AF-Chelate-650M chelating resin (stated section 2.2) | Handpicked 1–2 mg of species- and size-specific shells, cracked, clay removed (MQ water, methanol), reductively and/or oxidatively cleaned, leached in 0.001 M HNO3, dissolved in 1 M HNO3 and centrifuged; 5 µL of supernatant diluted with 200 µL 0.1 M HNO3 as the Me/Ca stock, Ca measured by ICP-AES, then diluted to [Ca] 10 ppm with 0.1 M HNO3 + 0.3 M HF — §2.4; the HF-bearing matrix 'was used only in the final dilution step' | About 100 mg whole-rock powder spiked with the three multi-element spikes (MES 1–3), digested in HF-HNO3, converted to chlorides, taken up in 7 mol/l HNO3 and diluted for LR and HR, with a Ru-Re solution added to each dilution — cleanroom, twice sub-boiled acids (Samples and sample preparation)
Digestion Acid(s) | Acid mixture used for sample dissolution. Record acid type(s) and concentration(s) where known (e.g., concentrated HF–HNO3 in Parr bombs is standard for silicate dissolution). | Basic | Read-Only | Text (free) | e.g., HF–HNO3 ¦ HF–HNO3–HClO4 ¦ HNO3–HCl ¦ HNO3 only ¦ Aqua regia ¦ N/A ¦ None | Source: Solution Introduction module | 2026-08-31 | preparation step | For isotope-ratio work, complete matrix dissolution is required to avoid isotopic fractionation during partial dissolution. |  | TD HNO3 leach: conc. HNO3; TD HF: conc. HF; TD HNO3: conc. HNO3; TD HCl: 6 M HCl; AR leach: aqua regia (3:1 HCl:HNO3) — §2.2 | HCl-HNO3 attack: 6 M HCl + 8 M HNO3; re-dissolution: 10 M HCl — §2.3.1 | ultrasonic HF decomposition: HF; bomb HF decomposition: 30 mol/l HF; re-dissolution: 0.5 mol/l HF — §2.5 | N/A — seawater; pre-concentration by chelation; no acid digestion | dissolution: 1 M HNO3 — §2.4 | hotplate HF-HNO3: 1–2 ml HF (24 mol/l) + 0.2 ml HNO3 (14 mol/l); bomb HF-HNO3: 1–2 ml HF (24 mol/l) + 0.2 ml HNO3 (14 mol/l), and 0.5 ml HF after 3 days; fluoride removal: 14 mol/l HNO3; chloride conversion: 6 mol/l HCl; final uptake: 7 mol/l HNO3
Digestion Vessel Type | Type of vessel used for acid digestion. | Basic | Read-Only | Text (free) | e.g., Savillex beaker ¦ Parr bomb ¦ TFE/TFM bomb ¦ Carius tube ¦ HPA-S ¦ Open beaker ¦ N/A ¦ None | Source: Solution Introduction module | 2026-08-31 | (none) | Vessel type constrains maximum temperature and pressure and determines suitability for refractory mineral dissolution (e.g., Parr bombs required for complete silicate digestion at high temperatures). |  | N | 15 mL Teflon capsule (stated section 2.3) | TFM bomb (TFM-981; stated section 2.1.1) | N/A — no digestion vessel; seawater pre-concentration | N/A — simple dissolution; stated section 2.4 | Closed 15 ml Savillex PFA vials, placed in Parr bombs for samples with refractory minerals
Digestion Step | Each distinct acid digestion step applied to dissolve the sample, listed in the order performed and named by its attack. A step is distinct when its acid mixture, vessel, temperature or duration differs from the one before; an evaporation or dry-down that carries no attack of its own is part of the step it follows, and an identical attack repeated on the residue is a repeat of that step rather than a new one. Enumerating the steps is what allows Digestion Acid(s), Digestion Temperature and Digestion Duration to be recorded per step. Record N/A where the sample is introduced without acid digestion. | Basic | Read-Only | Text (free) | e.g., '1: HF-HNO3, Parr bomb ¦ 2: aqua regia reflux of the residue' ¦ '1: HF-HNO3-HClO4 ¦ 2: HNO3-HCl ¦ 3: HNO3-H2O2' ¦ 'Single step' ¦ 'N/A (no acid digestion)' | Source: Solution Introduction module | 2026-09-08 | defines: preparation step | Naming each step is what lets acid, temperature and duration be read per step; a bare count cannot say which conditions belong to which attack, and cannot settle whether a repeated attack or an intervening evaporation counts. |  | TD HNO3 leach (3 ml conc. HNO3, 80 °C, 48 h, leachate discarded); TD HF (4 ml conc. HF, 100 °C, 48 h, evaporated); TD HNO3 (2 × 1 ml conc. HNO3, 100 °C); TD HCl (5 ml 6 M HCl, 80 °C, 15 h, then centrifuged); AR leach (3 ml aqua regia, 3:1 HCl:HNO3, 20 °C, 15 h, then centrifuged) — §2.2 | HCl-HNO3 attack (1.5 ml 6 M HCl + 0.5 ml 8 M HNO3, 130 °C, 48 h); re-dissolution (evaporated at 130 °C close to dryness, then 1.5 ml 10 M HCl) — §2.3.1, for magnetite and pyrite; FER-2 took 1 ml 28 M HF + 1 ml 8 M HNO3 at 130 °C for at least 48 h, then 1.5 ml 10 M HCl for 24 h | ultrasonic HF decomposition (basalts and andesites, <70 °C); bomb HF decomposition (peridotites and meteorites, 30 mol/l HF, 245 °C, mannitol added); re-dissolution (dried, then 5 ml 0.5 mol/l HF in an ultrasonic bath, fluorides removed by centrifuging) — abstract; §2.5 | N/A | dissolution (minimum volume of 1 M HNO3, 40–60 µL, then centrifuged 2 min at 10,000 rpm) — §2.4 | hotplate HF-HNO3 (non-refractory samples of basaltic composition, 12 h at 130 °C in closed 15 ml Savillex PFA vials); bomb HF-HNO3 (samples with refractory minerals such as granites, stirred 7 days at 180 °C in Parr bombs, reopened after 3 days and refilled with 0.5 ml HF); fluoride removal (re-dissolved in a few drops of 14 mol/l HNO3 and evaporated to incipient dryness, repeated twice); chloride conversion (2 ml 6 mol/l HCl heated at 80 °C, then evaporated at 50 °C); final uptake (5 ml 7 mol/l HNO3) — Samples and sample preparation
Digestion Temperature | Temperature at which acid digestion was carried out (°C). | Basic | Editable | Numeric (°C) | e.g., '90–120 °C (hotplate, Savillex beaker)' ¦ '150–190 °C (Parr bomb, refractory phases)' ¦ '220 °C (HPA-S)' | Source: Solution Introduction module | 2026-08-25 | preparation step |  |  | N | HCl-HNO3 attack: 130 °C; re-dissolution: 130 °C (evaporation) — §2.3.1 | ultrasonic HF decomposition: <70 °C; bomb HF decomposition: 245 °C; re-dissolution: N — abstract | N/A | N | hotplate HF-HNO3: 130 °C; bomb HF-HNO3: 180 °C; fluoride removal: ca. 80 °C (evaporation); chloride conversion: 80 °C, then 50 °C (evaporation); final uptake: N
Digestion Duration | Duration of the primary acid digestion step (hours or days). | Basic | Editable | Text (free) | e.g., '12 h' ¦ '24 h' ¦ '48 h (refractory accessory phases)' ¦ '3 days (multi-step with intermediate dry-down)' | Source: Solution Introduction module | 2026-08-25 | preparation step |  |  | N | HCl-HNO3 attack: 48 h; re-dissolution: N — §2.3.1 | N | N/A | N/A | hotplate HF-HNO3: 12 h; bomb HF-HNO3: 7 days; other: N
Final Solution Matrix | Acid type and concentration of the final dissolved sample solution introduced to the instrument. For isotope-ratio procedures, match acid molarity precisely between samples and bracketing standards. Procedures targeting HFSE may add trace HF to the final solution to keep those elements in solution. | Basic | Read-Only | Text (free) | e.g., 2% HNO3 ¦ 3% HNO3 ¦ 2% HNO3 + trace HF ¦ 6M HCl (intermediate) ¦ N/A ¦ None | Source: Solution Introduction module | 2026-09-08 | acquisition pass | Directly affects instrument response and ion transmission. Mismatched molarity introduces matrix-induced mass bias offsets. |  | all: 2% HNO3 doped with 1 ppb Tl — §2.4 | all: 2% HNO3 containing 5 ng/mL Rh — §2.3.2 | all: 0.5 mol/l HF with mannitol, diluted — 'The mannitol and HF concentrations in all samples and standard solutions were diluted to be similar to each other' (§2.5) | all: 1.0 M HNO3 — 'extracted trace elements were eluted with 1 mL of 1.0 M Q-HNO3' (§2.2) | all: 0.1 M HNO3 + 0.3 M HF — 'The acid matrix containing HF was used only in the final dilution step' | LR: 0.4 mol/l HNO3, about 120 µl of the uptake solution diluted to 50 ml (dilution factor ~21000); HR: 0.4 mol/l HNO3, 2.5 ml of the uptake solution diluted with H2O to 50 ml (dilution factor ~1000) — ca. 6 µl of a Ru-Re solution (60 µg/g Ru, 20 µg/g Re) added to each dilution
Sample Aliquot Mass or Volume | Mass (mg) of solid material digested or volume (mL) of liquid taken for dissolution. | Advanced | Editable | Numeric (mg or mL) | e.g., '50 mg (whole-rock silicate)' ¦ '25 mg (mass-limited meteorite)' ¦ '100 mg (low-abundance trace elements)' ¦ '1–5 mL (aqueous sample)' | Source: Solution Introduction module | 2026-08-25 | sample | Used for yield calculations and concentration back-calculation. |  | N | ~100 mg mineral powder (stated section 2.3) | N | 12 mL seawater (stated section 2.2) | 1-2 mg foraminifera shells (stated section 2.4) | About 100 mg of whole-rock powder
Chromatographic Separation Applied | Whether chromatographic separation of element groups (e.g., using ion exchange resin) was performed prior to analysis. Record resin type and elution matrix conditions if applied. Record 'None' if the direct digestion solution is analyzed. | Basic | Read-Only | Controlled list / Text | Yes ¦ N/A ¦ None | Source: Solution Introduction module | 2026-08-31 | (none) | Separation may be required to remove matrix elements that cause variable mass bias, to eliminate isobaric interferences, or to isolate an element group before measurement (e.g., Fe and Cu–Zn for isotope dilution, Hf separation from REE, removal of Cr and Ni from Fe fractions). It is usually mandatory for isotope-ratio procedures and optional for concentration measurement. |  | None — the SC-SF-ICP-MS splits were analysed unseparated; the anion-exchange separation (§2.3) was for MC-ICP-MS | AG1-X8 anion exchange resin + TRUspec resin (stated section 2.3.2) | None (direct analysis of 0.5 mol/l HF solution; stated section 2.1.2) | Toyopearl AF-Chelate-650M chelating resin (pre-concentration from seawater; stated section 2.2) | None (direct dissolution analysis; stated section 2.4) | None — 'dissolved rock samples are analysed directly by ID without previous separation'
Isotope Dilution Spike | Isotopic spike material added to the sample before or during digestion for isotope dilution (ID) analysis or double-spike mass bias correction. For double-spike procedures, add and equilibrate the spike before digestion. For concentration ID, the spike must be added before digestion for complete equilibration. Record spike name or certificate reference, spike isotopes, and whether added pre- or post-digestion. Record 'None' if ID or double-spike technique is not used. | Basic | Read-Only | Text (free) | e.g., Mixed multi-element spike ¦ Single-element spike ¦ None ¦ N/A | Source: Solution Introduction module | 2026-08-31 | (none) | Spiking before digestion corrects for isotopic fractionation during dissolution and column chemistry. |  | None (Tl added for mass fractionation correction only, not ID) | None (external calibration used; stated section 2.3) | N — Ti is not spiked; the Zr–Hf and Mo–Sn–Sb spikes serve the ICP-QMS elements | 57Fe, 62Ni, 65Cu, 68Zn, 111Cd, 207Pb enriched isotope spikes (stated Table 1) | None | Three multi-element spikes: MES 1 (86Sr, 135Ba, 145Nd, 149Sm, 235U), MES 2 (155Gd, 161Dy, 167Er, 171Yb, 207Pb), MES 3 (91Zr, 179Hf) — Table 1; calibrated by reverse ID against Alfa Aesar specpure solutions, spike-concentration uncertainty 0.5–1% RSD
3. Instrument & Software | [9 empty] | N | N | N | N | N | N | N
Instrument Manufacturer | Manufacturer of the instrument that performs the measurement, recorded as a controlled value. Where a procedure couples a sample-introduction system to an analysing instrument, this records the analysing instrument. Instrument Model gives the specific designation. | Basic | Read-Only | Controlled list | Thermo Fisher Scientific ¦ Agilent ¦ PerkinElmer ¦ Nu Instruments ¦ Analytik Jena ¦ Shimadzu ¦ Unknown ¦ N/A ¦ None | Source: Core module | 2026-08-30 | (none) | Recording it as a controlled value lets procedures be found by vendor. |  | Nu Instruments -- "a Nu Instruments Attom SC-SF-ICP-MS" | Thermo Fisher Scientific -- "An Inductively Coupled Sector Plasma Mass Spectrometer (Element I, ThermoFisher, USA)" | Thermo Fisher Scientific -- "(b) ICP-SFMS, Finnigan ELEMENT"; stated as "Finnigan ELEMENT", the pre-Thermo brand name | Thermo Fisher Scientific -- "a Thermo-Finnigan Element I (E1) HR-ICP-MS"; "Instrument ThermoFinnigan Element1" | Thermo Fisher Scientific -- "a Thermo Element XR, a single collector sector field high resolution inductively coupled plasma mass spectrometer" | Thermo Fisher Scientific -- "a ThermoFinnigan ELEMENT2 mass spectrometer"
Instrument Model | Model designation of the instrument that performs the measurement, including any generation or configuration suffix. Conventionally written with the manufacturer name included; Instrument Manufacturer records the vendor separately, as a controlled value, so that procedures remain findable by vendor. | Basic | Read-Only | Text (free) | e.g., Thermo Element 2 ¦ Thermo Element XR ¦ Nu Instruments Attom HR ¦ Nu Instruments AttoM ES ¦ N/A ¦ None | Source: Core module | 2026-08-31 | (none) |  |  | Nu Instruments Attom SC-SF-ICP-MS (stated section 2.3) | Thermo Fisher Element I HR-ICP-MS (stated Table 1) | Finnigan ELEMENT sector-field ICP-MS (stated section 2.1.2) | Thermo Finnigan Element I (E1) HR-ICP-MS (stated section 2.4) | Thermo Element XR single-collector SF-ICP-MS (stated section 2.3) | ThermoFinnigan ELEMENT2 (stated section on instrumentation)
ICP-MS Type | Type of ICP-MS system - the mass analyser and, where the instrument has one, the collector configuration. | Basic | Read-Only | Controlled list | Single-collector sector-field (SF-ICP-MS) ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-30 | (none) | Single-collector sector-field instruments use a magnetic sector combined with an electrostatic analyser to achieve high mass resolution through physical separation of ion beams. |  | Single-collector sector-field (SF-ICP-MS) | Single-collector sector-field (SF-ICP-MS) | Single-collector sector-field (SF-ICP-MS) | Single-collector sector-field (SF-ICP-MS) | Single-collector sector-field (SF-ICP-MS) | Single-collector sector-field (SF-ICP-MS)
Instrument Serial Number or Lab Identifier | Serial number or laboratory-internal identifier for the specific instrument unit. | Advanced | Editable | Text (free) |  | Source: ICP-MS module | 2026-08-25 | (none) | Supports traceability to instrument service records. |  | N | N | N | N | N | N
Torch Type | Type of plasma torch installed (e.g. standard quartz, high-matrix, low-flow). | Advanced | Read-Only | Text (free) | Standard quartz ¦ High-matrix ¦ Low-flow ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-30 | (none) |  |  | N | N | Quartz glass torch with sapphire injector (stated Table in section 2.1.2) | N | N | N
Interface Cone Configuration | Geometry and designation of the sampler and skimmer cones installed during analysis. State both cones. | Basic | Read-Only | Text (free) | e.g., Ni sample cone + standard Ni skimmer ¦ Ni sample cone + Ni H-skimmer ¦ Al sample cone + Al skimmer ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-31 | (none) | Some SF-ICP-MS instruments offer multiple cone geometries that differ in aperture size and affect ion transmission efficiency and matrix tolerance. |  | N | 1.1 mm Ni sampler + 0.8 mm Ni skimmer (Table 1) | 1 mm Ni sampler + 0.8 mm Ni skimmer (stated Table in section 2.1.2) | Ni/Cu sampler (Spectron Inc.) + Ni skimmer (Spectron Inc.; Table 2) | Pt-normal sampler + Pt-H skimmer (Table 1) | 1.0 mm Ni sampler + 0.5 mm Ni skimmer (Table 3)
Sampler and Skimmer Cone Material | Material composition of the sampler and skimmer cones. State both sampler and skimmer materials; a single statement is acceptable where they are identical. | Advanced | Read-Only | Controlled list / Text | Ni sampler and Ni skimmer ¦ Pt sampler and Pt skimmer ¦ Ni sampler and Pt skimmer ¦ Al sampler and Al skimmer ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-31 | (none) | Nickel is standard for most aqueous and nitric matrices; platinum is used for HCl-rich, high-TDS or organic matrices for its greater corrosion resistance; aluminium is used in some laboratories for high-purity work or enhanced sensitivity. |  | N | Ni sampler and Ni skimmer (Table 1) | Ni sampler and Ni skimmer (stated Table in section 2.1.2) | Ni/Cu sampler and Ni skimmer (Table 2) | Pt sampler and Pt skimmer (Table 1) | Ni sampler and Ni skimmer (Table 3)
Torch Depth | Distance between the load coil and the sampling cone tip (mm), also called injector depth or torch position depending on the instrument manufacturer. | Advanced | Editable | Text (free) | 5 mm ¦ Standard position (manufacturer default) ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-25 | (none) | Affects ion transmission efficiency, oxide formation, and doubly-charged species production. |  | N | N | N | N | N | N
Guard Electrode | Whether a guard electrode - also called a grounded shield electrode, extraction lens, or Fassel torch shield - is installed and active on the torch assembly. | Basic | Read-Only | Controlled list | On ¦ Off ¦ Not installed ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-30 | (none) | Capacitively decouples the plasma from the load coil, reducing secondary discharge and improving ion extraction efficiency. Actively used as a design feature on Thermo Element 2 and Element XR instruments. |  | N | N | N | N | N | N
Mass Resolution Setting | Operating mass resolution of the mass analyser. Where the analyser permits selection, state the mode; where individual target species are assigned to different modes, state each. | Basic | Editable | Controlled list / Text | Unit resolution (quadrupole, fixed) ¦ Low resolution (LR) ¦ Medium resolution (MR) ¦ High resolution (HR) ¦ Multiple modes assigned per target species ¦ N/A ¦ None | Source: ICP-MS module | 2026-09-01 | (none) | Sector-field instruments allow selection of low resolution (LR; m/Δm ≈ 300–400), medium resolution (MR; m/Δm ≈ 2500–5000), or high resolution (HR; m/Δm ≈ 10,000). |  | N — Nu Attom uses deflector peak jump; no LR/MR/HR designation stated; stated section 2.3 | N | MR (M/Delta-m = 3000; stated section 2.1.2) | LR (R ~300) and MR (R ~4000; stated section 2.4) | LR and MR (stated Table 1 and section 2.3) | LR (M/Delta-m = 300) and HR (M/Delta-m = 11000; stated section on instrumentation)
Detector Configuration | Type(s) of detector(s) installed in the mass spectrometer and the detection mode(s) used: pulse counting, analog and, on some sector-field instruments, Faraday. The cross-calibration between detector modes is recorded in Pulse/Analog Detector Nonlinearity Correction (Group 5). | Basic | Read-Only | Controlled list / Text | Single SEM, pulse counting only ¦ Single SEM, dual mode (pulse counting + analog) ¦ Single SEM, triple mode (pulse counting + analog + Faraday) ¦ N/A ¦ None | Source: single-collector ICP-MS module | 2026-09-11 | (none) | SF-ICP-MS instruments typically use a single SEM. Some instruments (e.g., Thermo Element XR) support triple mode: pulse counting for low signals, analog for mid-range signals, and Faraday cup for high signals. |  | N — the Attom's 'deflector peak jump' mode is stated (§2.4) | Counting — Table 1 'Detection mode' | Pulse counting mode — §2.1.2 | N | Dual mode, fixed for each isotope: counting for 7Li, 11B, 111Cd, 137Ba, 238U, 55Mn, 56Fe, 66Zn and analog for 25Mg, 27Al, 43Ca, 87Sr, 23Na — Tables 1 and 3; 'measured at a fixed detection mode to avoid detection mode switch (pulse to analog) during analysis' (§2.3.1) | N
Nebulizer Type | Type and material of the pneumatic nebulizer used for sample introduction. | Basic | Read-Only | Text (free) | e.g., PFA MicroFlow ¦ Meinhard ¦ Micromist ¦ PFA-ST ¦ N/A ¦ None | Source: Solution Introduction module | 2026-08-31 | (none) | Affects droplet size distribution and sample introduction efficiency. |  | Glass Expansion glass nebulizer (stated section 2.3) | N | Micro-flow PFA nebulizer PFA-20 (ESI, USA); self-aspiration (stated Table in section 2.1.2) | PFA microflow PFA-100 (Elemental Scientific; Table 2) | ESI 50 uL microconcentric nebulizer; self-aspirating (Table 1) | ESI microconcentric Teflon nebulizer (stated section on instrumentation)
Spray Chamber Type and Cooling Temperature | Type of spray chamber and cooling temperature if thermostatted (°C). | Basic | Read-Only | Text (free) | e.g., Scott double-pass (glass) ¦ Scott double-pass (PFA) ¦ Cyclonic at 7°C ¦ Cinnabar cyclonic at 7°C ¦ N/A ¦ None | Source: Solution Introduction module | 2026-08-31 | (none) | Chamber type and temperature control aerosol droplet size and the solvent vapor load reaching the plasma. In some procedures a spray chamber is placed downstream of a desolvating nebulizer to improve signal stability. |  | Glass Expansion cyclonic spray chamber (stated section 2.3) | N | Scott double-pass, uncooled, Teflon (stated Table in section 2.1.2) | PFA Teflon Savillex 100 mL with internal baffle (Table 2) | Teflon Scott-type single-pass (Savillex PFA; Table 1 footnote) | ESI Teflon spray chamber (stated section on instrumentation)
Desolvation System | Desolvating nebulizer or membrane desolvator used upstream of the plasma to reduce solvent load. Record 'None' if not used. | Advanced | Read-Only | Controlled list / Text | Aridus I ¦ Aridus II ¦ Apex Q ¦ Apex IR ¦ Apex HF ¦ None ¦ N/A | Source: Solution Introduction module | 2026-09-08 | acquisition pass | Lowers oxide production rates and increases sensitivity, and is commonly used for low-abundance target species or small sample sizes. Note the trade-off: desolvation can introduce additional instrumental mass bias instability relative to wet plasma introduction. |  | N | N | N | N | N | N
Sample Uptake Rate | Peristaltic pump speed or self-aspiration flow rate for sample introduction (µL/min or mL/min). | Basic | Editable | Numeric (µL/min or mL/min) | e.g., '50–100 µL/min (self-aspirating PFA nebulizer)' ¦ '200–400 µL/min (micro-flow, peristaltic)' ¦ '1 mL/min (conventional concentric)' | Source: Solution Introduction module | 2026-08-26 | (none) |  |  | 0.33 ml/min (stated section 2.3) | 200 uL/min (Table 1) | Self-aspiration; uptake time 60 s (stated section 2.1.2); volumetric flow rate N | 150 uL/min (Table 2) | 50 uL ESI nebulizer; uptake time 70 s (Table 1); volumetric flow rate N | ~100 uL/min (Table 3)
Acquisition Software | Instrument control, automation and data acquisition software used to collect the raw data, including version number. Distinct from Data Processing Software(s), which covers everything applied to the data after acquisition. | Basic | Editable | Text (free) |  | Source: Core module | 2026-08-26 | (none) |  |  | N | N | N | N | N | N
Data Processing Software(s) | All software applied to the data after acquisition in order to produce the reported quantities, including version numbers. List every package used. Distinct from Acquisition Software, which controls the instrument and collects the raw data. | Basic | Editable | Text (free) |  | Source: Core module | 2026-08-26 | (none) |  |  | N | N | N | N | N | N
4. Measurement Information | [6 empty] | 2026-08-26 |  |  | N | N | N | N | N | N | N
Analytical Mode | Primary analytical mode(s) executed under this procedure. Records one value for a single-mode procedure and all applicable modes for a multi-mode one. Applies even where a procedure defines no mode flag columns. | Basic | Read-Only | Controlled list | Solution nebulisation (continuous) ¦ Flow injection | Source: Core module | 2026-08-26 | (none) | Serves as the procedure-level declaration of measurement type, distinct from the mode flag columns, which state per-field applicability. Ensures every procedure record is self-describing and comparable across the library. |  | Solution nebulisation (continuous) -- "The Attom was operated with a Glass Expansion cyclonic spray chamber and glass nebulizer (uptake rate 0.33 ml/min)"; "both operated in wet plasma mode"; ESI SC-2 DX autosampler | Solution nebulisation (continuous) -- "Sample uptake rate 200 uL min-1"; "The components of the sample introduction system: nebulizer, spray chamber, torch, and the cones" [sec 2.1] | Solution nebulisation (continuous) -- stated in the acquisition parameters: "Middle resolution 50 s with 30 scans (continuous nebulization)" | Solution nebulisation (continuous) -- "Nebuliser PFA microflow (PFA-100), Elemental Scientific"; "Nebuliser sample uptake rate 150 uL min-1"; "Autosampler CETAC ASX-100". The flow-injection manifold of sec 2.2 is the offline pre-concentration step, not the ICP-MS introduction: "prevented the online coupling of the flow injection system directly to an ICP-MS" [sec 2.3] | Solution nebulisation (continuous) -- "a Teflon Scott type (single pass) spray chamber was constructed"; "we used a platinum injector (1.8 mm I.D.)"; ESI nebulizer | Solution nebulisation (continuous) -- "The ELEMENT2 was equipped with an ESI microconcentric Teflon nebuliser (flow rate ca. 100 ul min-1) and an ESI Teflon spray chamber"; "Sample uptake rate ca. 100 ul min-1"
RF Power | Radiofrequency forward power applied to the ICP, in watts. | Basic | Editable | Numeric (W) | e.g., 1200 ¦ 1300 ¦ 1350 ¦ 1550 | Source: ICP-MS module | 2026-08-31 | (none) | Controls ionization efficiency and oxide production rates. |  | N | 1300 W (Table 1) | 1.1 kW (stated Table in section 2.1.2) | 1300 W (Table 2) | 1250 W (Table 1) | 1235 W (Table 3)
Coolant (Plasma) Gas Flow Rate | Flow rate of the outer (coolant/plasma) argon gas stream that sustains the ICP plasma, in L/min. | Basic | Editable | Numeric (L/min) | 13–16 L/min | Source: ICP-MS module | 2026-08-26 | (none) | Influences plasma temperature and oxide ion formation. |  | N | 14.6 L/min (Table 1) | 14 L/min (stated Table in section 2.1.2) | 13 L/min (Table 2) | N | 15 L/min (Table 3)
Auxiliary Gas Flow Rate | Flow rate of the intermediate (auxiliary) argon gas stream, between the torch body and the injector tube, which positions the plasma relative to the load coil, in L/min. Distinct from the outer coolant (plasma) gas and from the gas stream that carries sample into the plasma. | Basic | Editable | Numeric (L/min) | 0.8–1.1 L/min | Source: ICP-MS module | 2026-08-26 | (none) |  |  | N | 0.96 L/min (Table 1) | 1.2 L/min (stated Table in section 2.1.2) | 1.05 L/min (varied and optimized daily; Table 2) | N | 0.9 L/min (Table 3)
Nebulizer Gas Flow Rate | Flow rate of the carrier argon gas delivered through the nebulizer (L/min). | Basic | Editable | Numeric (L/min) | 0.8–1.1 L/min | Source: Solution Introduction module | 2026-08-26 | (none) | Controls aerosol transport and strongly influences signal sensitivity and stability. |  | N | 0.90 L/min (sampling gas; Table 1) | 0.90 L/min (stated Table in section 2.1.2) | 1.2 L/min (varied and optimized daily; Table 2) | N | 1.0 L/min (sample gas; Table 3)
Make-up Gas and Flow Rate | Supplementary gas added to the sample-carrying stream between the sample introduction system and the plasma, with its identity and the procedure-registered target flow rate. Record any small nitrogen or hydrogen addition with its own flow, whose unit commonly differs from the make-up flow. Record 'None' explicitly where no supplementary gas is added, to distinguish it from not reported. | Advanced | Editable | Numeric (L/min) / Text | e.g., '0.25 (Ar, no desolvator fitted)' ¦ '0.85 Ar + 4 ml/min N2 (sensitivity enhancement)' ¦ 'None' | Source: ICP-MS module | 2026-08-26 | (none) | Argon make-up is standard and maintains total gas delivery where the carrier flow alone is insufficient — downstream of an ablation cell, or of a desolvation system that has removed solvent load. Small nitrogen or hydrogen additions enhance sensitivity for some elements. |  | N/A | N/A | N/A | N/A | N/A | N/A
Plasma Thermal Mode | Whether the ICP plasma is operated under cool plasma (<=900 W RF) or normal (hot) plasma (>1000 W RF) conditions, with the RF power that defines it. This field records the RF-POWER axis only. Solvent loading also affects plasma thermal state, but dry versus wet aerosol is recorded by Desolvation System — do not answer this field with 'dry plasma' or 'wet plasma'. Document independently of the RF Power field. | Basic | Read-Only | Controlled list / Text | Normal plasma (>1000 W RF) ¦ Cool plasma (≤900 W RF) ¦ N/A ¦ None | Source: ICP-MS module | 2026-09-29 | acquisition pass | Normal plasma (>1000 W RF) is standard for most solution SF-ICP-MS analyses. Cool plasma (≤900 W RF) reduces argide-based interferences (e.g., 40Ar12C+ on 52Cr) at the cost of reduced sensitivity and ionization efficiency for most elements. |  | N | N — RF power 1300 W (Table 1) | N — plasma power 1.1 kW (Table 1b) | N — incident RF power 1300 W (Table 2) | N — RF power 1250 W (Table 1) | N — RF power 1235 W (Table 3)
ICP Tuning | Description of the approach used to optimise ICP plasma conditions prior to analysis, including the reference material used for tuning and the acceptance criteria (e.g., oxide production threshold, sensitivity targets, mass calibration). | Advanced | Editable | Text (free) | e.g., 'Daily tune on 5 ng/ml Be-Rh-U to maximise sensitivity across the mass range; CeO+/Ce+ <2 permil' ¦ 'Tuned to ~1000 kcps/ppb Pb while maintaining flat-topped peaks' ¦ 'Optimised on 11B, 115In and 175Lu, with 2.5x10^6 cps/ppb on 115In set as an operational criterion; medium-resolution tune on 56Fe for separation from 40Ar16O' | Source: ICP-MS module | 2026-08-26 | (none) |  |  | "The instrument was tuned to provide ~1000 kcps/ppb Pbtotal while maintaining flat-topped peaks" | "The instrument was tuned using a 5 ng mL-1 solution containing Be, Rh and U in order to maximize the sensitivity covering the low-mid-high mass range. MO+/M+ ratios measured for Ce under the routine experiment conditions maintained at less than 2 permil" | Partially -- an oxide acceptance criterion is registered in the operating conditions, "Oxide forming rate <1% (CeO+/Ce+)"; no tuning solution or tuning procedure stated | "At the start of each day, prior to any sample analysis, the instrument was first tuned to produce maximum sensitivity and stability while also maintaining low oxide formation ... using a 5ppb solution of In"; "Further tuning using a 5ppb solution of Fe in the medium resolution mode ensured maximum separation of the 56Fe peak from the 40Ar16O peak"; oxide formation "typically below 5%", above 10% causing problems | "Instrumental sensitivity was optimized on 11B, 115In, and 175Lu", with sensitivities "set as operational criteria"; "in medium resolution tuning, the instrument was optimized on 56Fe to achieve a mass resolution" | N
Instrument Warm-up / Session Duration Limit | Minimum warm-up time required after plasma ignition before analyses begin, and any maximum session duration enforced to maintain stable operating conditions. | Advanced | Read-Only | Text (free) | e.g., '>=1 h warm-up after plasma ignition' ¦ 'Instrument conditioned with a matrix-matched solution before the first sample; no session duration limit' | Source: ICP-MS module | 2026-08-26 | (none) |  |  | N | N | N | N | Partially -- "The instrument was conditioned with a 10 [ppb solution]" before analysis, and detector mode was "kept fixed through the entire run to avoid detection mode switch induced changes in sensitivity". No warm-up time or session duration limit stated | N
Instrument Sensitivity | Instrument sensitivity achieved in the session, with the isotope or monitored property it was measured on and the conditions it applies to. May be expressed either as detected signal per unit concentration or per unit mass of target species delivered — counts per second per ppb, volts per ppm, counts per picogram — or as useful yield, the percentage of sampled atoms ultimately detected as ions, with the method used to derive it cited. A sensitivity the procedure requires before analyses may begin belongs with the tuning acceptance criteria. | N/A | Advanced | Numeric + unit / Text | e.g., '572 V/ppm' ¦ '~2.5 x 10^6 cps/ppb' ¦ '0.04 count pg-1 ml' ¦ '0.1% useful yield' ¦ 'N/A' | Source: ICP-MS module | 2026-09-29 | monitored property | Useful yield is the more comparable of the two wherever the amount of material consumed varies between procedures, as it does with spot size, fluence and repetition rate. |  | all: ~1000 kcps/ppb Pb total — 'tuned to provide ~1000 kcps/ppb Pbtotal while maintaining flat-topped peaks' (§2.4) | N | N | 59Co: ~5000 cps/nM; 55Mn: ~3100 cps/nM; other: N — slopes of the standard additions (Fig. 2, Table 4); the highest Mn sensitivity, 3560 cps/nM, at pH 8.1; the 5 ppb In tuning solution gave ~1,000,000 cps | N — 'a typical sensitivity of 2.5 × 10^6 cps/ppb on 115In' (§3.1), the tuning isotope; 1 ppb 11B and 115In sensitivities by spray chamber, injector and acid matrix in Table 2 | N
Target Species | The chemical species this procedure is designed to determine, recorded at whatever resolution the chemistry is resolved — element(s) for this technique; valence species where a procedure resolves oxidation state; compounds where it resolves molecules. Isotopes are not target species: isotopes of an element are the same chemical species. The subset actually measured may be narrower than the registered suite where species were excluded due to interferences or scope reduction. | Basic | Editable | Text (free) |  | Source: Target Species module | 2026-09-01 | defines: target species |  |  | Pb — Hg and Tl are monitored: ²⁰²Hg for the Hg interference on ²⁰⁴Pb, ²⁰³Tl and ²⁰⁵Tl for mass bias (§2.4) | Li, Be, Sc, Cr, Co, Ni, Cu, Zn, Ge, Rb, Sr, Y, Cs, Ba, La, Ce, Pr, Nd, Sm, Eu, Gd, Tb, Dy, Ho, Er, Tm, Yb, Lu — the 28 trace elements of Tables 2 and 3 | Ti — 'the analysis of Ti by middle resolution' (§2.1.2); ⁹³Nb is measured as the internal standard | Mn, Fe, Co, Ni, Cu, Zn, Cd, Pb — abstract; Table 1 | Li, B, Na, Mg, Al, Mn, Fe, Zn, Sr, Cd, Ba, U — 'The concentrations of foraminiferal trace elements of interest (Li, B, Na, Mg, Al, Mn, Fe, Zn, Sr, Cd, Ba, and U)' (§2.3.1); Ca is the ratio denominator | Rb, Sr, Y, Zr, Nb, Cs, Ba, La, Ce, Pr, Nd, Sm, Eu, Gd, Tb, Dy, Ho, Er, Tm, Yb, Lu, Hf, Ta, Pb, Th, U — twelve by isotope dilution and fourteen by relative sensitivity factors (Table 1)
Reported Variables and Units | The final variable(s) this procedure reports and their units — distinct from the fields recording what was *acquired* rather than what is reported. Record every reported variable, including intermediate quantities reported alongside final ones (e.g. both the 206Pb/238U ratio and the 206Pb/238U date). Where a reported variable is a nominal property with no magnitude (e.g. a mineral identification), record the variable and give the unit as 'N/A — nominal property'. | Basic | Read-Only | Text (free) | e.g., 'Element concentration (ppm); oxide (wt%)' ¦ 'd56Fe (permil vs IRMM-014); d57Fe (permil vs IRMM-014)' ¦ '206Pb/238U date (Ma); 207Pb/206Pb date (Ma)' ¦ 'Mineral identity (N/A — nominal property); match score (0-1)' | Source: Core module | 2026-08-26 | defines: reported property | A procedure may acquire many channels and report a small number of derived quantities; without this field a data consumer cannot tell which. |  | 206Pb/204Pb, 207Pb/204Pb, 208Pb/204Pb, 207Pb/206Pb, 208Pb/206Pb — dimensionless Pb isotope ratios (§2.4) | Li, Be, Sc, Cr, Co, Ni, Cu, Zn, Ge, Rb, Sr, Y, Cs, Ba, La, Ce, Pr, Nd, Sm, Eu, Gd, Tb, Dy, Ho, Er, Tm, Yb, Lu (µg/g) — Table 3 (FER-2) in µg g⁻¹; the unit of Table 4 (samples) is not stated in its header | Ti (µg/g); TiO2 — Ti by ICP-SFMS; Nb is from the ICP-QMS. Detection limits in solution (ng/g) and in rock (µg/g) in Table 2b | Mn, Fe, Co, Ni, Cu, Zn, Cd, Pb (nM, dissolved) — dissolved concentrations in seawater | B/Ca (µmol/mol); Li/Ca, Mg/Ca, Al/Ca, Sr/Ca, Cd/Ca, Ba/Ca, U/Ca, Na/Ca, Mn/Ca, Fe/Ca, Zn/Ca (µmol/mol or mmol/mol) — Me/Ca ratios; Li, Mg, Al, Sr, Cd, Ba and U in low resolution, Na, Mn, Fe and Zn in medium (Table 3) | Rb, Sr, Y, Zr, Nb, Cs, Ba, La, Ce, Pr, Nd, Sm, Eu, Gd, Tb, Dy, Ho, Er, Tm, Yb, Lu, Hf, Ta, Pb, Th, U (µg/g) — Equations 1 and 2
Monitored Masses | Specific masses monitored in this procedure, grouped by the target species element they serve where they serve one. Covers atomic isotopes and, where a reaction cell shifts an target species onto a different mass, the product mass actually measured. Includes interference-monitor and internal-standard masses, which serve no target species and so have no parent element. The target species list is given by the Target Species field and is never inferred from the element symbols appearing here. | Basic | Read-Only | Text (free) | e.g., 71Ga ¦ 109Ag ¦ 115In ¦ 205Tl ¦ 209Bi ¦ N/A ¦ None |  | 2026-09-10 | defines: monitored property per target species |  |  | 204Pb, 206Pb, 207Pb, 208Pb → Pb; 202Hg, 203Tl, 205Tl → none — 'masses 202Hg, 203Tl, 204Pb, 205Tl, 206Tl, 207Pb and 208Pb' (§2.4); '206Tl' is the paper's typo for 206Pb | 7Li → Li; 9Be → Be; 45Sc → Sc; 52Cr → Cr; 59Co → Co; 60Ni → Ni; 63Cu → Cu; 66Zn → Zn; 72Ge → Ge; 85Rb → Rb; 88Sr → Sr; 89Y → Y; 133Cs → Cs; 138Ba → Ba; 139La → La; 140Ce → Ce; 141Pr → Pr; 143Nd → Nd; 147Sm → Sm; 153Eu → Eu; 157Gd → Gd; 159Tb → Tb; 163Dy → Dy; 165Ho → Ho; 166Er → Er; 169Tm → Tm; 172Yb → Yb; 175Lu → Lu; 103Rh → none — Table 2 'Isotope measured'; Table 3 prints 74Ge where Table 2 has 72; ¹⁰³Rh is the internal standard (Table 1) | 47Ti, 49Ti → Ti; 93Nb → none — 'In one scan, 47Ti and 49Ti together with 93Nb were measured. Signals of 47Ti and 49Ti were summed up to determine a total Ti signal' (§2.1.2) | 56Fe, 57Fe → Fe; 60Ni, 62Ni → Ni; 63Cu, 65Cu → Cu; 66Zn, 68Zn → Zn; 110Cd, 111Cd → Cd; 207Pb, 208Pb → Pb; 55Mn → Mn; 59Co → Co; 95Mo → none — the spike and reference isotopes of Table 1, ⁵⁵Mn and ⁵⁶Fe named in §2.4, and ⁹⁵Mo counted for the MoO⁺ correction; Co is monoisotopic | 7Li → Li; 11B → B; 25Mg → Mg; 27Al → Al; 43Ca → none; 87Sr → Sr; 111Cd → Cd; 137Ba → Ba; 238U → U; 23Na → Na; 55Mn → Mn; 56Fe → Fe; 66Zn → Zn — Table 3; ⁴³Ca is measured in both resolutions | 85Rb → Rb; 86Sr, 88Sr → Sr; 89Y → Y; 90Zr, 91Zr → Zr; 93Nb → Nb; 133Cs → Cs; 135Ba, 137Ba → Ba; 139La → La; 140Ce → Ce; 141Pr → Pr; 143Nd, 145Nd → Nd; 147Sm, 149Sm → Sm; 151Eu, 153Eu → Eu; 155Gd, 158Gd → Gd; 159Tb → Tb; 161Dy, 163Dy → Dy; 165Ho → Ho; 166Er, 167Er → Er; 169Tm → Tm; 171Yb, 172Yb → Yb; 175Lu → Lu; 178Hf, 179Hf → Hf; 181Ta → Ta; 207Pb, 208Pb → Pb; 232Th → Th; 235U, 238U → U; 47Ti, 49Ti, 99Ru, 101Ru, 185Re, 187Re → none — the determined ratios of Table 1; Ti, Ru, Eu and Re ratios measured for the in-run mass-fractionation correction
Mass Resolution Assignment | Mass resolution mode used for acquisition. One target species may be acquired at more than one resolution, so the assignment is per acquired mass rather than per element. The overall mode(s) used in the procedure are recorded in Mass Resolution Setting (Group 3). | Basic | Read-Only | Text (free) | e.g., LR ¦ MR (m/Δm ~2500) ¦ HR ¦ N/A ¦ None |  | 2026-09-29 | acquisition pass | The selected resolution determines which polyatomic interferences are physically resolved by the magnetic sector. |  | N — Nu Attom uses deflector peak jump; resolution mode not stated in LR/MR/HR terms | N | all: middle resolution (M/ΔM = 3000) — 'At this resolution, major interferences on Ti such as 15N16O2+ or 31P16O+ can be separated' (§2.1.2) | all: medium resolution (R = 4000) for Mn, Fe, Co, Ni, Cu and Zn, low resolution (R = 300) for Cd and Pb — §2.4 | LR: low resolution (Δm/m = 300), for 7Li, 11B, 25Mg, 27Al, 43Ca, 87Sr, 111Cd, 137Ba, 238U; MR: medium resolution (Δm/m = 4000), for 23Na, 43Ca, 55Mn, 56Fe, 66Zn — Table 3 | LR: M/ΔM = 300, for Rb to Sm and Hf to U; HR: M/ΔM = 11000, for Eu to Lu — Table 1; HR resolves light REE oxides on the heavy REE
E-scan Range | Electric scan range used for peak acquisition, expressed as percentage of the centre mass (%). Record 'N/A' if E-scan acquisition mode is not used. | Advanced | Read-Only | Numeric (%) | e.g., 15% ¦ 10% |  | 2026-08-31 | (none) | Varies the accelerating voltage to cover masses in the vicinity of the set magnetic mass without re-scanning the magnet. |  | N/A — Nu Attom uses deflector peak jump, not E-scan | N | N | N | N | N
Triple Scanning Mode | Whether each mass peak is scanned three times per cycle and the results averaged (Y/N). Record 'N/A' if not applicable to the instrument. | Advanced | Read-Only | Controlled list | Yes ¦ No ¦ N/A ¦ None |  | 2026-08-30 | (none) | Used to reduce noise from short-term magnetic field instabilities on sector-field instruments. Triple scanning affects the effective integration time per cycle. |  | N/A — Nu Attom instrument; not applicable | N | N | N | N | N
Dwell Time per Mass | Count (dwell) time at the mass position, in milliseconds. Where the procedure defines it per sweep or per scan rather than per measurement, state that basis. | Basic | Editable | Numeric (ms) / Text |  |  | 2026-09-10 | monitored property |  |  | all: 500 µs — §2.4 | all: 100 ms per isotope — Table 1 | 47Ti: 0.18 s; 49Ti: 0.18 s; 93Nb: 0.18 s — integration time per 1 s (Table 2b) | N | 7Li: 0.005 s; 11B: 0.005 s; 25Mg: 0.01 s; 27Al: 0.01 s; 43Ca: 0.02 s (LR), 0.01 s (MR); 87Sr: 0.04 s; 111Cd: 0.02 s; 137Ba: 0.02 s; 238U: 0.05 s; 23Na: 0.02 s; 55Mn: 0.01 s; 56Fe: 0.05 s; 66Zn: 0.01 s — 'Sample Time' (Table 3); samples per peak 50 or 100 in LR, 25 in MR | all: 0.1 s, 15 samples per peak — 'Peaks were monitored in electrical scanning mode for 0.1 s' and Table 3
Number of Scans per Replicate | Number of complete mass scans accumulated per analytical replicate. A scan — also called a sweep or a pass — is one complete traversal of the monitored masses by a sequentially scanning analyser, so the scan count multiplied by the per-mass dwell time gives the total integration time per replicate. Distinct from a cycle in simultaneous multi-collection, which is one readout of all detectors at once rather than a traversal of masses. | Basic | Read-Only | Integer | e.g., 20 ¦ 40 ¦ 100 ¦ 250 ¦ 2000 |  | 2026-09-08 | acquisition pass |  |  | all: 30 sets of 2000 sweeps — total analysis time 4.5 min (§2.4) | N | all: 30 scans in 50 s — Table 1b 'Middle resolution 50 s with 30 scans (continuous nebulization)' | N | LR: 15 passes × 3 runs; MR: 5 passes × 3 runs — Table 3; Table 1 prints 3 passes for MR | all: 70 to 120 scans of the whole mass spectrum — for one analysis, a total acquisition time of 10 minutes per sample
Number of Replicates | Number of replicate measurements performed on the same sample, or on the same nominal location where the technique is spatially resolved. For spot analysis this is the number of individual spots per grain or location; for transects, the number of replicate lines; for mapping, the number of map acquisitions of the same area; for solution work, the number of discrete replicate measurements acquired per sample solution. | Advanced | Basic | Integer | e.g., 3 ¦ 5 ¦ 10 |  | 2026-08-31 | (none) |  |  | 1 — one acquisition of 30 sets of 2000 sweeps (§2.4) | N | N | N | 3 runs (LR and MR; Table 3) | 3 — 'Triplicate determinations were performed for each digestion'
Analysis Sequence | The repeating order in which calibration or bracketing standards, quality-control and secondary reference materials, blanks and unknowns are interleaved within a measurement session. Adjustments must maintain the bracketing strategy defined in the procedure. | Basic | Editable | Text (free) | e.g., Matrix-matched bracketing every 5 samples ¦ Standard after every 3 samples ¦ Dual-aliquot design (Barrat approach) ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-31 | (none) |  |  | N | N | Standard solution every two samples; each sample measurement ~6 min, including ~3 min washing — §2.1.2 | N | Blocks of seven samples, each bracketed by a pair of acid blanks and internal consistency standards (Standard 3 of the calibration) — §2.3.1 | N — a total of ca. 20 minutes per sample 'including the measurement of blank, standard solution, washout'; the order is not stated
Wash Time Between Samples | Duration of instrument rinse between successive sample solutions (seconds). | Basic | Read-Only | Numeric (s) | e.g., 120 ¦ 180 ¦ 210 ¦ 300 | Source: Solution Introduction module | 2026-08-31 | (none) | Controls carryover and memory effects. For isotope-ratio work, complete washout is critical to prevent cross-contamination between samples with differing isotopic compositions. |  | 10 s in each of two 2% HNO3 reservoirs, then a third reservoir for the blank — the blank solution 'was replaced every 20 samples' (§2.4) | 60 s — Table 1; '1 min with 3% v/v HNO3' (§2.1) | ~3 min with 0.5 mol/l HF — §2.1.2; Table 1b: background measured before each sample 'after 200 s wash' | N | 120 s (Table 1) | N
Internal Standard Element | Element(s) used as the internal standard, the monitored isotope, and how the internal-standard concentration was established for unknowns. Where the element serves isotope-ratio work it must have a known isotopic composition and a mass close to the target species's. Record 'None' together with the alternative used where no internal standard is applied. | Basic | Read-Only | Text (free) | e.g., Y + Ho ¦ Rh (10 ppb) ¦ Tm spike aliquot ¦ None ¦ N/A | Source: ICP-MS module | 2026-09-08 | acquisition pass |  |  | all: Tl (203Tl, 205Tl), for mass-bias correction — §2.4 | all: Rh (103Rh) — Table 1 | all: Nb (93Nb) — §2.1.2 | all: none — isotope dilution for Fe, Ni, Cu, Zn, Cd and Pb, standard additions for Co and Mn (§2.5) | all: none — Me/Ca ratios, with Ca measured in both resolutions | all: the ID-determined elements, for the RSF elements (Table 1 ratios, e.g. 93Nb/90Zr, 175Lu/172Yb) — Ru and Re are added for the mass-fractionation correction, not as internal standards
Internal Standard Concentration | Target concentration of internal standard element(s) in all measured solutions (µg/L or ppb). Record 'N/A' where Internal Standard Element is 'None'. | Advanced | Read-Only | Numeric (µg/L) | e.g., '5 µg/L In' ¦ '10 µg/L Rh' ¦ '1–10 µg/L (In, Rh, Re multi-element)' ¦ '20 µg/L (high-dilution runs)' | Source: Solution Introduction module | 2026-08-26 | (none) |  |  | 1 ppb Tl in sample matrix (stated section 2.3) | 5 ng/ml Rh (stated section 2.3) | N | N/A | N/A | N
Oxide Production Method and Threshold | Method used to quantify plasma oxide production and the acceptance threshold applied before commencing analysis. Record both the monitored mass ratio(s) and the maximum allowed threshold(s). Measured values are recorded in Oxide Production. The monitored ratio is procedure-dependent: ThO+/Th+ (m/z 248/232), UO+/U+ (254/238), CeO+/Ce+ (156/140) and BaO+/Ba+ are all in use. | Basic | Read-Only | Text (free) | e.g., CeO+/Ce+ (m/z 156/140) < 0.4%; BaO+/Ba+ (m/z 154/138) < 0.15% ¦ CeO+/Ce+ < 0.5% ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-31 | (none) | Stricter thresholds than Q-ICP-MS are commonly required. |  | N | N | CeO+/Ce+ < 1% (stated section 2.1.2) | In (5 ppb) used for LR tuning; oxide rate typically <5% (stated section 2.4) | Sensitivity criteria used (not oxide threshold): >=250000 cps/ppb for 11B; >=2500000 for 115In; >=2000000 for 175Lu (stated section 2.3.1) | N
Oxide Production | Measured oxide production ratio obtained during session-start tuning, for the proxy specified in Oxide Production Method and Threshold. Record the measured value and confirm whether the procedure threshold was met. | N/A | Basic | Text (free) | CeO+/Ce+ = 0.3% (threshold < 0.4% met) ¦ BaO+/Ba+ = 0.08% (threshold < 0.15% met) | Source: ICP-MS module | 2026-08-26 | (none) |  |  | N/A | N/A | N/A | N/A | N/A | N/A
Doubly-Charged Species Monitor | The mass ratio monitored to estimate doubly-charged ion (M²⁺) formation during instrument tuning. The monitor species and the mass positions monitored should be stated explicitly. Analogous to Oxide Production Method and Threshold for oxide monitoring. | Advanced | Editable | Text (free) | Ba2+/Ba+ (m/z 69/138) ¦ Ce2+/Ce+ (m/z 70/140) ¦ N/A ¦ None | Source: ICP-MS module | 2026-09-11 | (none) | Doubly-charged ions appear at half the nominal mass of the parent ion and can interfere with lighter target species masses. In SF-ICP-MS, doubly-charged species are not suppressed by collision cells and require monitoring through plasma tuning. Ba2+/Ba+ (m/z 69/138) is the most common proxy. |  | N | N | N | N | N | N
Doubly-Charged Species Production | Measured percentage of doubly-charged ion production for the monitored species at the time of instrument tuning. The acceptable threshold is typically <1% or <3%. Record both the threshold and the measured value. | Advanced | Editable | Text (free) | Ba2+/Ba+ < 5% (threshold); 2.1% (measured) ¦ Ce2+/Ce+ < 3% (threshold); 1.8% (measured) ¦ N/A ¦ None | Source: ICP-MS module | 2026-09-11 | (none) | Elevated doubly-charged production indicates incomplete ionization and potential interference on elements at approximately half the mass of abundant matrix components. |  | N | N | N | N | N | N
Ion Counter Dead Time | Dead time of the ion-counting detector(s), used in the dead-time correction applied to high count rates. Distinct from pulse/analog cross-calibration, which relates the two detector modes rather than correcting counting losses within the pulse-counting mode. | Basic | Editable | Numeric (ns) | e.g., '23' ¦ '20 ns, determined from a U calibration series' ¦ 'N/A (analog detection only)' |  | 2026-09-10 | (none) |  |  | N | N | N | N | N | N
Acquisition Pass | The acquisition passes the procedure is divided into, named and described so that fields keyed by acquisition pass can point at them. A pass is a sub-procedure: a distinct traversal of the measurement with its own configuration, run in sequence on the same material. State what distinguishes each pass — resolution mode, cup or cell configuration, plasma or introduction path, spot size — since that differs by technique. Identical repeats of one configuration are replicates, not passes, and belong in Number of Replicates. | Basic | Read-Only | Text (free) | e.g., 'Pass 1 (LR, M/dm 300); Pass 2 (MR, M/dm 4000)' ¦ 'Pass 1 (KED, He, on-mass); Pass 2 (MS/MS, O2 mass-shift)' ¦ 'Pre-ablation; Run 1 (major, MR, 30 um); Run 2 (trace, LR, 130 um)' ¦ 'Single pass' | Source: ICP-MS module | 2026-09-29 | defines: acquisition pass | [6 empty] | LR; MR — low resolution (Δm/m = 300) and medium resolution (Δm/m = 4000) methods, 'Calcium was measured in both low and medium resolution to maintain accuracy of Me/Ca ratios obtained from the two mass resolution modes' (Table 3) | LR; HR — two solutions of different dilution: dilution factor ~21000 for LR and ~1000 for HR; 'the transmission decreases by a factor of about 100 from the LR to the HR mode'
Number of Acquisition Passes | Number of acquisition passes the procedure runs. A count of the passes enumerated in Acquisition Pass, recorded separately so multi-pass procedures are findable without parsing that field. | Basic | Read-Only | Integer | e.g., 1 ¦ 2 ¦ 3 | Source: ICP-MS module | 2026-09-08 | (none) | [6 empty] | 2 — low and medium resolution (Tables 1 and 3) | 2 — the LR and HR solutions
5. Data Processing | [6 empty] | 2026-08-26 |  |  | N | N | N | N | N | N | N
Drift Correction Method | Method used to correct for instrumental signal drift across a session. | Basic | Read-Only | Controlled list | IS normalization ¦ Standard bracketing ¦ IS normalization + bracketing ¦ None ¦ N/A |  | 2026-08-30 | (none) |  |  | N | Rh internal standard — 'In order to correct the instrumental drift, the internal standard concentration of Rh was kept constant at 5 ng mL−1 in the sample, calibration, and blank solutions' (§2.2) | Standard solution every two samples — 'No time drift of fTi was observed during measurements over 2 h, thus all fTi were averaged and used' (§2.7) | Elution acid, an enriched-isotope standard and a natural-abundance commercial standard measured every 10–12 samples — 'to assess instrument drift and mass bias' (§2.4) | N — the blocks of seven are bracketed by acid blanks and a consistency standard; at low calcium concentrations there was 'minimal instrumental sensitivity drift' (§2.3.1) | N — 'Instrumental drift of SF-ICP-MS has only a negligible effect on the reproducibility of ID determined concentrations since isotope ratios are used'
Calibration Strategy per Target Species | Approach used to convert measured ion signals to reported concentrations, and specifically any case where different target species or target species groups within one procedure are calibrated differently - different primary standards for different mass ranges or phases, or one element serving as internal standard while others are externally calibrated. Where a single strategy applies to all target species, record that strategy. Where the procedure reports isotope ratios only and no concentrations, record 'Not applicable (isotope ratios only)'. | Basic | Read-Only | Text (free) | e.g., External calibration ¦ Isotope dilution ¦ Isotope dilution (67Zn spike) ¦ External calibration + standard addition ¦ N/A ¦ None | Source: ICP-MS module | 2026-09-29 | target species |  |  | Pb: internal normalisation to 205Tl/203Tl = 2.3871, with Pb mass-bias factors from master Pb vs Tl correlation lines of numerous SRM981 analyses — §2.3–2.4; SRM981 results 'were primarily used to update the long-term Pb vs Tl mass bias master correlations' | all: external calibration with 103Rh internal standardisation — multi-element certified solutions (VHG Labs) diluted to 0.1, 1, 5 and 10 ng/mL (§2.2); Table 1 'Calibration: External' | Ti: two-stage internal standardisation against Nb, X_Ti = X_Nb × f_Ti × S_Ti — S_Ti is the (47Ti + 49Ti)/93Nb intensity ratio and f_Ti the relative concentration factor from the standard solution; X_Nb from the ICP-QMS Nb/Mo and Nb/Zr ratios (§2.7) | Fe, Ni, Cu, Zn, Cd, Pb: isotope dilution; Mn, Co: standard addition — standard additions of 50, 100 and 150 µL of the mixed Co and Mn solution to 12 mL sub-samples (§2.3, §2.5) | all: external calibration with matrix-matched standards, samples and standards concentration-matched within ±5% — ±1% in the Fig. 2 dilution series; at [Ca] of 10 ppm with normal cones, 5 ppm with the Jet and X cones; samples first measured for Ca by ICP-AES (§2.3.1) | Sr, Zr, Ba, Nd, Sm, Gd, Dy, Er, Yb, Hf, Pb, U: isotope dilution (Eq. 1); Rb, Y, Nb, Cs, La, Ce, Pr, Eu, Tb, Ho, Tm, Lu, Ta, Th: relative sensitivity factor against an ID-determined element measured in the same resolution and run (Eq. 2) — Table 1
Signal Integration Interval Method | Method used to define which portion of the acquired signal is integrated for each reported measurement, and the rule for excluding unstable or anomalous portions. State whether the interval is the full acquisition, a stable sub-interval selected by inspection, or selected by an automated criterion, and name the criterion. | Basic | Read-Only | Controlled list / Text | Full acquisition ¦ Stable interval selected by inspection ¦ Automated threshold ¦ Automated outlier rejection ¦ Fixed time window ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-26 | (none) |  |  | N | N | N | N | N | N
Signal Integration Time | The integration time actually used for the reported signal, in seconds - the time retained after the signal-integration interval has been applied, and therefore equal to or shorter than the total acquisition. Where the procedure computes it from an acquisition structure, record that basis. | N/A | Basic | Numeric (s) |  | Source: ICP-MS module | 2026-08-26 | (none) |  |  | N/A | N/A | N/A | N/A | N/A | N/A
Blank / Background Correction Method | Method used to subtract background and, where relevant, procedural blank from the measured signal. State what is subtracted - instrument background acquired with no sample reaching the plasma, an on-peak zero acquired on the same matrix, or a separately measured procedural blank - over what interval it is measured, and at what granularity it is applied. | Basic | Read-Only | Controlled list / Text | Gas blank ¦ On-peak zero ¦ Solution blank ¦ Procedural blank ¦ On-peak zero + procedural blank ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-26 | (none) |  |  | Blank determination before each sample acquisition (average 900 cps on 208Pb, equivalent to 1.8 ppt Pb); on-line baseline correction — §2.4 | N — a blank solution in 2% HNO3 is part of the calibration set (§2.2); the blank correction is not described | Background measured before each sample after a 200 s wash — Table 1b | All sample concentrations corrected for the ammonium acetate buffer and the extraction procedure (flow manifold, chelating resin, elution acid), which includes the ICP-MS background — §2.5 | N | N — 'After correction for background and mass fractionation'; the method is not described
Spike / Outlier Filtering Approach | Criteria used to identify and exclude anomalous data - signal spikes, individual cycles, or whole replicate measurements - before the reported value is calculated. State where in the reduction sequence the filter is applied. | Advanced | Editable | Text (free) | 2σ exclusion ¦ Grubbs test ¦ Visual inspection ¦ None ¦ N/A | Source: ICP-MS module | 2026-08-30 | (none) |  |  | N | N | N | N | N | Dixon outlier test on each block of ten ratios — 'Generally, less than one ratio had to be excluded from the whole data set'
Pulse/Analog Detector Nonlinearity Correction | Whether a correction was applied for nonlinear detector response at the transition between pulse-counting and analog (and Faraday, for triple-mode instruments) detection modes. Cross-calibration factors between detector modes must be confirmed, typically measured each session. Record 'Applied' and describe the method, the detector modes involved and the target species affected; 'None' where a crossover exists on this instrument but no correction was made, giving the reason; and 'N/A' where the detector is pulse-counting only and no crossover exists. | Advanced | Editable | Controlled list / Text | Applied ¦ N/A ¦ None | Source: single-collector ICP-MS module | 2026-09-11 | acquisition pass | For SF-ICP-MS instruments, the signal transitions between pulse counting (low concentrations) and analog detection (high concentrations). For triple-mode instruments (e.g., Thermo Element XR), an additional transition to Faraday cup detection occurs at the highest signal intensities. |  | N | N | N | N | all: detector cross-calibration between pulse and analog modes performed daily — §2.3.1 | N
Spectral Interference Corrections Applied | Whether mathematical corrections for isobaric, polyatomic or residual interferences are applied in data reduction, supplementary to any suppression already achieved by chemical separation, mass resolution, or a collision/reaction cell. Detail for each affected mass is carried by Interfering Species and Interference Correction Method. | Basic | Read-Only | Controlled list / Text | e.g., 'Yes — ⁸⁷Rb on ⁸⁷Sr, corrected from the ⁸⁵Rb monitor' ¦ 'Yes — doubly-charged Er and Yb on the Sr and Rb masses' ¦ 'No explicit corrections applied; medium resolution resolves the polyatomic interferences' ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-27 | (none) |  |  | Y — 'on-line correction for baselines and Hg interference (202Hg/204Pb always <0.043)' (§2.4) | N | N | Y — MoO⁺ on ¹¹⁰Cd and ¹¹¹Cd corrected by counting ⁹⁵Mo (§2.4); Fe and Mn interferences resolved in medium resolution | N — Na, Mn, Fe and Zn are measured in medium resolution (Table 3); no correction is described | Y — the heavy REE are measured in HR to resolve light REE oxides; NIST SRM 612 data for Rb, Zr, Nb, Hf, Ta, Pb and U omitted because of oxide interferences and spiking
Interfering Species | The isobaric, polyatomic and doubly charged species that overlap the measured masses and are corrected in data reduction - direct isobars, oxides and argides, hydrides, and abundance-sensitivity tailing from an adjacent large beam. Name each species and the mass it affects. | Basic | Read-Only | Text (free) | e.g., BaO+ ¦ Zr2+ ¦ Oxides ¦ None ¦ N/A | Source: ICP-MS module | 2026-09-29 | monitored property | In SF-ICP-MS, mass resolution is the primary interference mitigation strategy; mathematical corrections address residual interferences not resolved at the operating resolution. |  | 204Pb: Hg; other: N — monitored at 202Hg (§2.4) | N | N | 110Cd, 111Cd: MoO+; 56Fe: 40Ar16O and 40Ca16O; 55Mn: 40Ar15N; other: N — §2.4 | N — MR resolution used to resolve polyatomic interferences; no additional mathematical correction reported; stated section 2.3 | 153Eu, 155Gd, 158Gd, 159Tb, 161Dy, 163Dy, 165Ho, 166Er, 167Er, 169Tm, 171Yb, 172Yb, 175Lu: light REE oxides; other: N — in NIST SRM 612 also 69Ga16O on 85Rb, 74Ge16O on 90Zr, 162Dy16O on 178Hf and 165Ho16O on 181Ta
Interference Correction Method | Equation or procedure used to calculate and remove each interference contribution, together with how its magnitude was established - a monitor mass measured simultaneously and scaled by natural abundance ratios, a production-rate factor measured on a reference material or interference standard solution, or a tailing factor measured on a pure standard. Name the reference material used. | Basic | Read-Only | Text (free) | e.g., Interference standard solutions (Burney–Neal approach) ¦ Mass-balance equations ¦ Empirical correction factors from pure element solutions ¦ None ¦ N/A | Source: ICP-MS module | 2026-09-10 | monitored property |  |  | 204Pb: on-line correction for Hg interference, monitored at 202Hg; other: N — §2.4 | N | N | 110Cd, 111Cd: 95Mo counted in all analyses, with slopes from linear regressions of 95Mo/110Cd and 95Mo/111Cd on Cd-free Mo standards (1–100 nM) run at the start and end of each session; 56Fe, 55Mn: resolved in medium resolution; other: N — §2.4 | N | 153Eu, 155Gd, 158Gd, 159Tb, 161Dy, 163Dy, 165Ho, 166Er, 167Er, 169Tm, 171Yb, 172Yb, 175Lu: resolved in HR (M/ΔM = 11000); other: N
Isotope Dilution Data Reduction Method | Mass balance approach used to calculate sample mass fractions, and where applicable isotope ratios, from spike-sample isotope ratio measurements. For double-spike procedures the ratios and concentrations are co-determined through the spike inversion; for conventional isotope dilution the IUPAC ID equations are used. Record 'None' where neither isotope dilution nor double-spike methods are applied. | Basic | Read-Only | Controlled list / Text | IUPAC isotope dilution equations ¦ Double-spike inversion ¦ ID-IS (isotope dilution internal standardisation) ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-31 | (none) |  |  | None | None | N — Ti is determined by internal standardisation, not isotope dilution | Standard isotope dilution equation (de Jong et al.), with a per-element mass-bias factor applied to the sample isotope ratios — §2.5 | None | Eq. 1 for the ID elements and Eq. 2 for the RSF elements, from mass-fractionation- and background-corrected mean ratios, with R_ik in Eq. 2 corrected for the spike contribution to isotope k
Memory Effect Mitigation | Procedure applied to identify and minimise carry-over of high-concentration or isotopically distinct material from a preceding measurement into the current one. Mitigation is applied primarily at measurement time, by allowing sufficient washout or rinse between successive introductions. At data processing level, record any flagging or exclusion of measurements where the required washout may not have been achieved. | Advanced | Editable | Text (free) | e.g., Extended rinse between high-concentration samples and unknowns; flagged in sequence log ¦ Replicate first measurement discarded if preceded by high-concentration standard ¦ None required for this target species suite ¦ N/A ¦ None | Source: ICP-MS module | 2026-09-01 | (none) |  |  | 10 s wash in two 2% HNO3 reservoirs between samples (stated section 2.3) | 60 s wash between samples (Table 1) | 0.5 mol/l HF carrier and wash solution; ~3 min wash per sample — §2.1.2 | N | 120 s washout (Table 1) | N
Uncertainty Level | The convention at which reported uncertainties are quoted — 1-sigma, 2-sigma, or a 95% confidence interval — and whether a measured spread is reported as a standard error or a standard deviation. State the convention applying to all values reported under this procedure, and state each separately where different reported quantities use different conventions. Distinct from Uncertainty Propagation Method, which addresses how component uncertainties are combined rather than how the result is quoted. | Basic | Read-Only | Controlled list / Text | 2σ (95% confidence) ¦ 1σ (68% confidence) ¦ 95% confidence interval ¦ 2 standard errors (2SE) ¦ 1 standard deviation (1SD) ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-26 | (none) | A reported uncertainty is not interpretable without it: the same numeric value means different things at 1-sigma and at 95% confidence. |  | 2 standard errors for within-run precision; 2sd for the averages of standards — 'Typical within-run precision (2 standards errors)' (§2.4); '(2sd, n = 22)' | 1 standard deviation -- "The mean values and respective standard deviations (s) for three analyses"; "Mean +/- s (n = 3)"; "RSD = standard deviation/mean x 100%" | RSD% with observed ranges in parentheses | Mixed and each stated: "Mean blank +/- 1 S.D. (pmoles)"; "The precision is calculated as the percent relative standard deviation (%RSD) (n = 3)" [Table 4 footnote]; "95% confidence limit" [Table 6 caption] | 2 sigma -- "with 2r analytical uncertainty" and "the gray area represents the 2r spread in the B/Ca measured at 10 ppm [Ca]Matrix" (r = sigma in the extracted text) | RSD for repeatability of triplicate determinations; "confidence intervals (1s)"; the method result is quoted as a "combined standard uncertainty"
Uncertainty Propagation Method | The approach used to propagate analytical uncertainty through the data reduction chain to the final reported value. State which sources are included in the propagation: counting statistics, calibration standard uncertainty, internal standard uncertainty, drift correction, and any systematic contributions. Distinct from Uncertainty Level, which states the convention at which the resulting uncertainty is quoted. | Advanced | Editable | Text (free) | Quadrature propagation ¦ Monte Carlo ¦ ISO GUM ¦ None ¦ N/A | Source: ICP-MS module | 2026-08-30 | reported property |  |  | N | N | N | N | N | all: combined standard uncertainty (Eurachem/CITAC 2000), 1–2% for ID and 2–3% for RSF — components: instrumental drift, matrix effects, ratio measurement, spectroscopic interferences, spike calibration, RSF calibration, mass fractionation and constants (Willbold et al. 2003)
Normalization / Standards-Based Correction | Post-acquisition normalization applied to the reported data beyond the primary calibration — for example correction to a reference value derived from secondary reference materials, or correction for a systematic bias those materials reveal. Record 'None' if no additional normalization is applied. | Advanced | Editable | Text (free) | CI chondrite normalization ¦ Normalization to BHVO-2 working values ¦ None ¦ N/A | Source: Composition QC module | 2026-08-30 | reported property |  |  | N | N | N | N | N | N
Calibration Factor and Determination Method | An externally-calibrated factor that converts the measured quantity into the reported quantity, how it was determined, and its uncertainty. Applies where the conversion depends on a factor calibrated against a reference of independently known value, rather than on the instrument response alone. Distinct from the fields that name the calibration material and that state which approach applies to which target species, where the technique has them: this field records the resulting factor itself. | Basic | Editable | Text (free) |  | Source: Calibration Factor module | 2026-09-01 | reported property |  |  | all: Pb mass-bias factors from Pb vs Tl mass-bias plots of numerous SRM981 analyses — 'Raw Pb isotope ratios were corrected for instrumental mass fractionation using the measured 205Tl/203Tl and Pb mass bias factors derived from Pb vs Tl isotope mass bias plots generated from numerous analyses of SRM981' (§2.4) | N — external calibration against multi-element solutions at 0.1, 1, 5 and 10 ng/mL with ¹⁰³Rh held at 5 ng/mL (§2.2); no factor is stated | Ti: relative concentration factor f_Ti against Nb, from the standard solution and averaged over the session; other: N — Eq. 2 (§2.7) | Fe, Ni, Cu, Zn, Cd, Pb: mass-bias correction factor, measured over true natural isotopic ratio, from replicate analyses of a commercial ICP standard (High Purity Standards); Mn, Co: slopes of the standard-addition curves; other: N — §2.5 | N — 'Before sample analysis, mass bias determination was carried out using a foraminiferal matrix matched solution' (§2.3.1); no factor is stated | all: in-run mass-fractionation factors from a power law MF = a × b^(m−c) + d fitted to 47Ti/49Ti, 99Ru/101Ru (LR), 151Eu/153Eu (HR) and 185Re/187Re, and RSF calibrated once per analytical run on a multi-element standard solution — 'Typical values are a: 0.08 to 0.23, b: 0.96 to 0.98, c: -10 to 10 and d: -0.005 to 0.005'
Procedural Blank Level | The level of the analytical blank and, where the reported quantity is a ratio, its composition, since a blank subtracted from a ratio biases the result unless its own composition is known. For the procedure: the typical or maximum blank expected. For the session: the blank measured in this session. Companion to the blank correction method. | Advanced | Basic | Text (free) |  | Source: Blank module | 2026-10-29 | target species | A blank that is large relative to the sample, or of unknown composition, biases every result it is subtracted from. |  | Pb: total procedural blank <100 pg — 'total procedural blanks (dissolution and/or leaching, including centrifuging) are estimated to be <100 pg'; sample/blank ratios ≥1500 (§2.3) | Li: 0.026; Be: 0.008; Sc: 0.020; Cr: 0.068; Co: 0.007; Ni: 0.035; Cu: 0.089; Zn: 0.216; Ge: 0.020; Rb: 0.012; Sr: 0.010; Cs: 0.004; Ba: 0.029; other: N — ng/mL (Table 2); 'the highest blank level in Zn would contribute less than 0.01% of the amount of analyte' | N | Mn: 0.433 ± 0.026; Fe: 2.791 ± 0.083; Co: 0.078 ± 0.006; Ni: 0.457 ± 0.104; Cu: 0.184 ± 0.027; Zn: 3.044 ± 0.018; Cd: 0.045 ± 0.003; Pb: 0.017 ± 0.001 — pmol, mean reagent blank ± 1 SD from one day's analysis, elution acid plus ammonium acetate buffer (Table 5) | B: 2.0 ± 1.0 µmol/mol (as B/Ca); other: N — abstract | N — the total procedural blanks enter only through the LOD
Analysis Inclusion and Rejection Criteria | The rules determining which individual results contribute to a combined result: the limits a result must meet (for example on analytical total or stoichiometry) and the grounds on which one is rejected (for example overlap onto a neighbouring phase, or an outlier test with its threshold), including any acceptance threshold on the dispersion statistic. An individual result is the value of the reported quantity obtained from one acquisition: a replicate measurement of the same solution or location, a spot or grain within a sample, or an independently prepared aliquot or digestion, whichever the procedure combines. Distinct from filtering the acquired signal during data reduction (removing spikes, cycles or scans, or discarding an acquisition whose signal is compromised): this field records the rules by which finished results enter a combined result. In the dataset: flag each individual result excluded from a combined value, with the rule that excluded it. | Basic | Editable | Text (free) | e.g., 'Totals 98.5–101.5 wt%; points overlapping a neighbouring phase rejected' ¦ 'All results included; no filter applied' | Source: Aggregation module | 2026-10-26 | (none) | The same measurements give different combined values under different acceptance rules, so the rules are needed to reproduce a combined value from its individual results. |  | N — no acceptance or rejection rule is stated. Re-read 2026-10-26 for the rules-only definition (Module_Aggregation v6); the earlier reading, kept for provenance: n stated per averaged result (BCR-2 n = 39, AGV-2 n = 13, BR n = 11, JB-2 n = 9, JB-3 n = 11, SRM981 n = 22 and n = 16). One documented exclusion, from the quality assessment rather than from a reported aggregate: "Results for the pure Pb standard NIST SRM981, analysed many times with the soil samples, are not included here, because it contains no matrix and may thus not a be a good indicator of data quality for the soil samples analysed here" [sec 3.1]. No acceptance or rejection rule, and no acquired-versus-included count, stated | N — no acceptance or rejection rule is stated. Re-read 2026-10-26 for the rules-only definition (Module_Aggregation v6); the earlier reading, kept for provenance: "The mean values and respective standard deviations (s) for three analyses were listed in Table 3"; n = 3 throughout. No acceptance or rejection rule stated | N — no acceptance or rejection rule is stated. Re-read 2026-10-26 for the rules-only definition (Module_Aggregation v6); the earlier reading, kept for provenance: "Orgueil and Allende were analyzed 4 times and twice from the sample digestion, respectively ... analytical results for each run are shown in the table" alongside the averages. No acceptance or rejection rule stated | N — no acceptance or rejection rule is stated. Re-read 2026-10-26 for the rules-only definition (Module_Aggregation v6); the earlier reading, kept for provenance: "The blank solutions were analysed at least three times on the ICP-MS"; "parallel triplicate samples"; n = 3 for reference materials and n = 5 for the GEOTRACES samples. No acceptance or rejection rule stated | N — no acceptance or rejection rule is stated. Re-read 2026-10-26 for the rules-only definition (Module_Aggregation v6); the earlier reading, kept for provenance: "Open symbols represent an average of 10 measurements acquired during a single instrument session. The solid symbols represent the average of the open symbols"; and for a second figure "which is a total of 15 measurements"; acquisition structured as 3 runs x 15 passes (low resolution) or 3 x 5 (medium). No acceptance or rejection rule stated | N — no acceptance or rejection rule is stated. Re-read 2026-10-26 for the rules-only definition (Module_Aggregation v6); the earlier reading, kept for provenance: Partially, and the most complete of the six -- "Five independent analyses (different spikings/digestions) of BHVO-1 were carried out over a time period of 4 months. Triplicate determinations were performed for each digestion"; "the results of three to four independent analyses of sixteen other RMs"; "Only one digestion was prepared for the USGS reference glasses BCR-2G, BHVO-2G and BIR-1G, and NIST SRM 612 respectively and were measured in triplicate". No acceptance or rejection rule stated
Combination Method | Whether the reported variable is an individual result or combines several, and if combined, how: the statistical model used, including any criteria governing which model is applied, and what the individual results are combined over (for example per grain, per phase within a sample, or per sample). Record 'Not combined' where each individual result is reported on its own. Record the model only; the software implementing it belongs in Data Processing Software. In the dataset: report each combined value with its uncertainty and the number of individual results it combines, and identify, for each individual result, the combined value it contributes to. | Basic | Editable | Text (free) | e.g., Arithmetic mean per phase within each sample ¦ Weighted mean with MSWD-based uncertainty expansion ¦ York (1969) isochron regression ¦ Not combined | Source: Aggregation module | 2026-10-26 | reported property | A methodological choice that changes the result: an arithmetic and a weighted mean of the same data, or a Model-1 and a Model-3 regression, give different values and different uncertainties. |  | all: average of repeat analyses of each reference material — 'Results for SRM981 analysed with the NVP samples yield average ...'; averages with '2sd' and n for BCR-2, AGV-2, BR, JB-2 and JB-3 | all: mean of three analyses — 'The mean values and respective standard deviations (s) for three analyses were listed in Table 3' | all: average of repeat runs per sample — 'Orgueil and Allende were analyzed 4 times and twice from the sample digestion, respectively ... analytical results for each run are shown in the table' alongside the averages | all: mean of replicate analyses, and average slope of the standard additions — 'Mean blank ± 1S.D.'; 'Average slopes resulting from the regression analysis of Co and Mn standard additions' | all: average of 10 measurements within a session, and average of the session averages — 'Open symbols represent an average of 10 measurements acquired during a single instrument session. The solid symbols represent the average of the open symbols' | all: mean of triplicate determinations per digestion, and mean over the independent analyses of each reference material — 'the mean values of the triplicate determinations and their RSDs' (Table 4); within an analysis, ratios are the mean of all scans (n = 70–120)
Constants and Reference Values Used | Physical constants and reference values used in data reduction to calculate the final reported quantity (e.g., decay constants for age calculation, standard isotope ratios, or other citable reference values used in a correction or calculation), together with their source. Distinct from the Group 6 reference-material fields, which document accepted values for specific calibration/validation materials rather than universal physical constants. Record "None" if no citable, revisable physical constants feed into this procedure's data reduction. | Basic | Editable | Text (free) | e.g., "λ238U = 1.55125×10⁻¹⁰ yr⁻¹, λ235U = 9.8485×10⁻¹⁰ yr⁻¹ (Jaffey et al. 1971); 238U/235U = 137.818 (Hiess et al. 2012)" ¦ "None" | Source: Core module | 2026-08-26 | (none) |  |  | 205Tl/203Tl = 2.3871 (Woodhead 2002), used to correct instrumental mass fractionation by internal normalisation with the exponential law | N | N | Partially -- "A mass bias correction factor for each of the six elements was calculated from the measured natural isotopic ratio divided by the true natural isotopic ratio". The true natural isotopic ratios used, and their source, are not stated | N | Relative atomic masses M_El and M_S "(Loss 2003)"; "the known natural isotopic abundances of the isotopes i and k in the sample (Rosman and Taylor 1998)", stated to be adequately known ("uncertainty < 0.2%"); in-run mass fractionation determined "by comparing determined 47Ti/49Ti, 99Ru/101Ru (in LR mode), 151Eu/153Eu (in HR mode) and 185Re/187Re ratios with known values (Rosman and Taylor 1998)". For Pb the paper compares two reference choices -- "average Pb isotope abundances (Rosman and Taylor 1998)" versus the BHVO-1 TIMS composition of "Woodhead and Hergt 2000" -- and quantifies the consequence: "The difference between both approaches is 0.4% (concentration of Pb: 2.13 ug g-1 versus 2.14 ug g-1)"
6. Quality Control & Uncertainty | [6 empty] | 2026-08-26 |  |  | N | N | N | N | N | N | N
Primary Calibration Standard Name | Name and reference material identifier of the primary reference material(s) against which the instrument is calibrated — converting raw signal intensities to concentrations, or anchoring an isotope ratio as the bracketing standard or zero-delta reference. Give the material name, its source or supplier, and a citation for the accepted values used. Where calibration instead uses the vendor's stored library or theoretical response factors rather than measured reference materials — 'standardless' or 'semi-quantitative' quantification — record that here, naming the library or model used. 'None' means no calibration was performed at all, which is a different answer. | Basic | Editable | Text (free) | e.g., BHVO-2 (working values) ¦ Orgueil CI chondrite (matrix-matched) ¦ NIST SRM 1640a ¦ Lab-prepared single-element standard mix ¦ N/A ¦ None | Source: Composition QC module | 2026-09-29 | target material x target species | Results calibrated against different published values for the same material are not directly comparable. |  | all [Pb: NIST SRM981] — for the Pb vs Tl mass-bias master correlations (§2.4) | all [all: multi-element certified solutions (VHG Labs)] — diluted to 0.1, 1, 5 and 10 ng/mL (§2.2) | N — 'The standard solution was measured every two samples' (§2.1.2); its Ti and Nb contents are not stated | all [Fe, Ni, Cu, Zn, Cd, Pb: enriched isotope spike (57Fe, 62Ni, 65Cu, 68Zn, 111Cd, 207Pb) calibrated against ICP High Purity Standards; Mn, Co: mixed Co and Mn standard-addition solution from Spectrosol AA standards] — §2.1 | all [all: matrix-matched Me/Ca calibration standards (Standards 0–8)] — §2.2, §2.3.1 | all [Sr, Zr, Ba, Nd, Sm, Gd, Dy, Er, Yb, Hf, Pb, U: multi-element spike solutions MES 1–3 calibrated by reverse ID against Alfa Aesar specpure solutions; Rb, Y, Nb, Cs, La, Ce, Pr, Eu, Tb, Ho, Tm, Lu, Ta, Th: multi-element standard solution (16–17 µg/g stock from Alfa Aesar specpure solutions, diluted to 1–5 ng/g)] — Table 1
Calibration Standard Measurement Frequency | How often the primary calibration or bracketing standard is measured relative to unknown samples within a session. Specify the bracketing scheme where a fixed one is used, such as one standard measurement before and after each sample. | Basic | Read-Only | Text (free) | e.g., Before and after every 3–5 samples ¦ Matrix-matched bracketing ¦ Start and end of session only ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-31 | (none) |  |  | N | N | Every two samples (stated section 2.1.2) | N | Standard 3 brackets each block of seven samples — §2.3.1 | Standard solution once per analytical run (e.g. once per day) — RSF values constant within 1s over at least ten hours
Secondary Reference Materials | Quality-control reference material(s) measured as unknowns alongside samples to assess accuracy independently and to monitor drift. Give the material name, its source, and a citation for the accepted or reference values used for comparison. | Basic | Editable | Text (free) | e.g., BHVO-2 ¦ BCR-2 ¦ AGV-2 ¦ CI chondrite reference ¦ N/A ¦ None | Source: Composition QC module | 2026-08-31 | defines: standard |  |  | BCR-2, AGV-2, JB-2, BR, JB-3 (stated Tables 1-2) | FER-2 — iron-formation reference material (CCRMP), 'used to validate of the proposed method' (§2.2) | JB-1, JB-2, JB-3, JA-1, JA-2, JA-3, JP-1, BHVO-1, AGV-1, PCC-1, DTS-1 — GSJ and USGS silicate reference materials (§2.3); the chondrites are samples | NASS-5; SAFe S1; SAFe D2 — NASS-5 (NRCC certified open-ocean seawater) and the SAFe inter-comparison samples (§2.6) | CAM-wuellerstorfi; CAM-Uvig-1; CAM-Uvig-2; CAM-Mix — the four Cambridge consistency standards (§2.2) | AGV-1, AGV-2, BCR-1, BCR-2, BHVO-1, BHVO-2, G-2, JR-1, KL2-G, ML3B-G, NIST SRM 612, BIR-1, OU-6, BCR-2G, BHVO-2G, BIR-1G, PCC-1 (stated Table 5)
Detection Limit | Detection limit, one per reported concentration variable (one per target species, these being the same set), with its units. For the procedure: the typical detection limits. For the session: the limits achieved in this session, where they are determined for the session as a whole (for example, from the blanks of the run). The calculation method is recorded separately in Detection Limit Method. Record 'N/A' where the procedure reports no concentrations. In the dataset: report the detection limit of each analysis, where the procedure computes one. | Advanced | Basic | Numeric + unit / Text |  | Source: Composition QC module | 2026-10-29 | reported property |  |  | N | Li: 0.024; Be: 0.006; Sc: 0.028; Cr: 0.043; Co: 0.009; Ni: 0.014; Cu: 0.052; Zn: 0.176; Ge: 0.015; Rb: 0.009; Sr: 0.008; Cs: 0.006; Ba: 0.018; other: N — method detection limits (MDL, 3 s) in ng/g (Table 2); instrumental detection limits are also given, in ng/mL | Ti: 4 µg/g; other: N — 3σ in rock; 21 ng/g in solution (Table 2b) | Mn: 0.007; Fe: 0.021; Co: 0.002; Ni: 0.026; Cu: 0.007; Zn: 0.005; Cd: 0.0006; Pb: 0.0002 — nM, 3 SD, for the extraction of a 12 mL sample (Table 5) | B/Ca: 2 µmol/mol; other: N — 'We report a B/Ca detection limit of 2 µmol/mol' (abstract) | all: about 0.1 to 10 ng/g sample equivalent for most elements — Figure 2
Detection Limit Method | Formula or approach used to calculate the detection limits, with a citation for the method where one exists. Must be consistent with the values reported in Detection Limit. | Basic | Read-Only | Controlled list / Text | 3σ blank ¦ 3σ background ¦ 3σ counting statistics ¦ 3× blank mean ¦ Poisson statistics ¦ N/A ¦ None | Source: Composition QC module | 2026-08-27 | reported property |  |  | N | all: 3 × SD of six procedural blanks (MDL) — Table 2 note; the IDL is 3 × SD of six replicate measurements of 2% HNO3 | all: 3σ — Table 2b | all: 3 SD of the reagent blank — Table 5, for a 12 mL sample | N | all: 3 s of total procedural blanks including spiking, 50 measurements in LR and 20 in HR — Limits of detection
Limit of Quantification (LOQ) Method | Reference or description of the method used to calculate the limit of quantification (LOQ): the lowest concentration reliably measurable with acceptable precision and accuracy. Required when concentrations near the LOD are reported. | Advanced | Basic | Controlled list / Text | 10σ blank ¦ 10σ background ¦ 10× LOD ¦ 3.3× LOD ¦ S/N ≥ 10 ¦ N/A ¦ None | Source: ICP-MS module | 2026-08-27 | reported property |  |  | N | N | N | N | N | all: RSD better than 10% on the low-concentration RM PCC-1, ca. 10 to 900 ng/g — 'about 10 to 20 times the LOD for most elements'
Counting Statistics Error | How the uncertainty predicted from counting statistics — the theoretical limit set by the Poisson distribution of the counts accumulated — is computed for each reported quantity, and the sigma level it is quoted at. It is derived from the counts on the target species together with those on any background or blank subtracted from it. Distinct from the scatter actually observed within a measurement or between repeated measurements, which is recorded separately. In the dataset: report, for each analysis and reported quantity, the counting-statistics uncertainty at the stated sigma level. | Advanced | Editable | Text (free) | e.g., '2SE predicted from counting statistics over 45 cycles' ¦ '1 sigma, counting statistics plus propagated uncertainties' ¦ 'N/A' | Source: Composition QC module | 2026-10-27 | reported property | Where a procedure reports both, agreement indicates the measurement is shot-noise limited, and a larger observed scatter indicates a further source of variance. |  | N | N | N | N | N | N
Internal (Within-Measurement) Analytical Precision and Assessment Method | How the precision of a single measurement is assessed from the scatter of the cycles, sweeps or integrations that make it up: the statistic (2SE, 2SD, 1s RSD) and the number of cycles it is computed over, for each reported quantity. Distinct from Counting Statistics Error, which records how the uncertainty predicted from the counts is computed rather than the scatter observed. In the dataset: report, for each measurement and reported quantity, the internal precision as the stated statistic. | Advanced | Editable | Text (free) | e.g., '2SE = 2SD/sqrt(n), n = 45 cycles' ¦ '2SE of individual measurements' ¦ 'N/A' | Source: ICP-MS module | 2026-10-27 | reported property | This is the finest of the three precision levels the library records, below within-session and between-session precision: it describes the repeatability of one analysis rather than agreement between analyses, and is normally the smallest of the three. |  | 206Pb/204Pb, 207Pb/204Pb: ±0.03–0.06 (±0.17–0.35%); 208Pb/204Pb: ±0.10–0.19 (±0.26–0.50%); 207Pb/206Pb: ±0.0006–0.0012 (±0.07–0.14%); 208Pb/206Pb: ±0.0020–0.0040 (0.09–0.18%) — 'Typical within-run precision (2 standards errors)' (§2.4); the ±0.001–0.002 in §2.3 is the MC-ICP-MS's | N | N | N | B/Ca: 1.0%; other: N — abstract | N
Within-Session Analytical Precision and Assessment Method | Precision of repeated measurements of a reference material within a single analytical session, and the method used to assess it. The assessment method specifies the reference material or standard measured, the number of replicates n, and the statistic reported (1s RSD, 2s RSD, 2SD, 2SE, 95% CI). For the procedure: the assessment method and the expected within-session precision. For the session: the precision achieved in this session, assessed by that method. Distinct from the internal precision of a single measurement, which derives from counting statistics over the cycles of that measurement rather than from repeated analyses. In the dataset: report the analyses of each secondary reference material, where they are included; the session value summarises them and must agree with them. | Basic | Basic | Text (free) | e.g., RSD of replicate standard measurements ¦ RSD of secondary RM replicates ¦ N/A ¦ None | Source: ICP-MS module | 2026-10-29 | standard x reported property |  |  | BCR-2, AGV-2, JB-2, BR, JB-3 [all: 2sd of repeat analyses] — Table 1; 'regular analyses of Tl-doped ~1 ppb solutions of several rock standards (unseparated)' (§2.4) | FER-2 [all: RSD <5%] — three analyses, Table 3 (§3) | all [Ti: RSD 3.6% (2.3–5.4%)] — TTi/93Nb RSD% (Table 2b) | NASS-5, SAFe S1, SAFe D2 [all: 95% confidence limit, n = 3] — Table 6 | all [B/Ca: 4.0% (2σ), average within-run external precision] — abstract | BHVO-1 [all: RSD of triplicate determinations on one digestion, generally better than 1%] — Table 4
Between-Session (Long-Term) Analytical Precision and Assessment Method | Precision of measurements across multiple analytical sessions over weeks to months — long-term or intermediate precision — and the method used to assess it. Specify the reference material, the number of measurements and sessions, the time span covered, and the statistic reported. For the procedure: the assessment method and the long-term precision on record when the procedure is registered. For the session: the long-term precision as it stands at this session, including this session's results where they contribute. | Advanced | Basic | Text (free) | e.g., RSD of secondary RM across sessions ¦ Control chart ¦ N/A ¦ None | Source: ICP-MS module | 2026-10-29 | standard x reported property | Long-term precision is normally poorer than within-session precision and is the figure a data user should carry when comparing results from different sessions. |  | N | N | N | N | CAM-wuellerstorfi, CAM-Uvig-1, CAM-Uvig-2, CAM-Mix [all: ±2σ of repeat analyses over 8 months] — Table 4; n = 180, 130, 100 and 150 | BHVO-1 [all: RSD of five independent digestions over 4 months, 1–3%] — abstract and Table 4; the sixteen other RMs by RSD of three to four independent analyses (Table 5)
Analytical Accuracy and Assessment Method | Offset between measured and accepted values for secondary reference materials, and the method used to assess it. Specify the reference material and the source of its accepted values, the number of analyses, and the quantities assessed. Express the offset in the form appropriate to what the procedure reports: percent relative bias for concentrations, or deviation in delta or ratio units for isotopic quantities. For the procedure: the assessment method, the accepted-value source and the expected offset. For the session: the offset found in this session, with any systematic bias and its likely cause. In the dataset: report the analyses of each secondary reference material, where they are included; the session value summarises them and must agree with them. | Basic | Basic | Text (free) | e.g., Percent recovery of secondary RM ¦ Z-score ¦ Bias from certified value ¦ N/A ¦ None | Source: ICP-MS module | 2026-10-29 | standard x reported property |  |  | BCR-2, AGV-2, JB-2, BR, JB-3 [all: % deviation from published Pb isotope values] — Table 1 and §3 | FER-2 [all: most ratios to literature values within 0.9–1.1] — Fig. 7 (§3) | JB-1, JB-2, JB-3, JA-1, JA-2, JA-3, JP-1, BHVO-1, AGV-1, PCC-1, DTS-1 [all: consistent with previous studies] — abstract; Tables 5 and 6 | NASS-5, SAFe S1, SAFe D2 [all: agreement with the NASS-5 certified and SAFe consensus values] — Table 6 | CAM-wuellerstorfi, CAM-Uvig-1, CAM-Uvig-2, CAM-Mix [B/Ca: across degrees of sample dilution] — Fig. 2 | BHVO-1 [all: within 1–2% of published ID data and 2–3% of all published data]; AGV-1, AGV-2, BCR-1, BCR-2, BCR-2G, BIR-1, BIR-1G, BHVO-2, BHVO-2G, G-2, JR-1, KL2-G, ML3B-G, NIST SRM 612, OU-6, PCC-1 [all: most within 3–4% of published data] — abstract; Table 5
Goodness-of-Fit or Dispersion Statistic | The statistic used to describe the scatter among the individual results contributing to a combined value, or to test whether that scatter exceeds what analytical uncertainty alone predicts: for example one standard deviation of the mean, or the MSWD with its probability of fit. A threshold for accepting it belongs with the inclusion criteria. In the dataset: report, for each combined value, the value of the statistic, and any further statistic (for example the standard error of the mean, or a 95% confidence interval) beside it, named. | Basic | Editable | Text (free) | e.g., One standard deviation of the mean ¦ MSWD with probability of fit | Source: Aggregation module | 2026-10-26 | reported property | Without the statistic named, the uncertainty given with a combined value cannot be interpreted, nor can whether the combined results form a single population be judged. |  | all: 2sd — '(2sd, n = 22)'; averages given with '±2sd%' | all: s, one standard deviation, and RSD — 'Mean ± 1 s (n = 3)'; 'RSD = standard deviation/mean × 100%' | all: RSD% — 'Averages of the reproducibility (RSD %)' | all: %RSD, with 1 S.D. for blanks — 'The precision is calculated as the percent relative standard deviation (% RSD)'; 'Mean blank ± 1S.D.' | N | all: RSD — Tables 4 and 5
Additional Notes | Any procedure- or analysis-specific information not captured by a structured field anywhere in this TAPP — including anomalies, deviations from the registered procedure, instrument modifications, and supplementary context. Scope is the whole document, not Group 6: this is the last field of the TAPP and covers all six groups. Use sparingly; a structured field is preferred for anything that can be formally categorised. | Advanced | Advanced | Text (free) |  | Source: Core module | 2026-08-26 | (none) |  |  | Nu Attom SC-SF-ICP-MS in single-collector mode; 30 sets x 2000 sweeps = 4.5 min total analysis; Tl-spiked matrix (1 ppb Tl) for mass bias correction; blank ~900 cps on 208Pb (stated section 2.3) | Chromatographic separation (AG1-X8 + TRUspec) performed before SF-ICP-MS; reflected power <2 W (stated Table 1); pulse counting detection only | Continuous sample introduction with an uptake time of 60 s (0.04 ml per measurement); quartz glass torch with sapphire injector (Table 1b) | Off-line pre-concentration on Toyopearl AF-Chelate-650M resin; enriched isotope spikes added before extraction; standard additions for Mn and Co (§2.2–2.3) | Teflon Scott-type single-pass spray chamber and platinum injector (1.8 mm I.D.) to reduce instrumental boron blanks; HF in the final matrix for rapid boron washout (§2.3) | Magnetic jump followed by electric scan (Table 3); acquisition 10 min per sample; ca. 20 min per sample in all, a mass spectrometer efficiency of almost 90%

## Legends
Procedure-Level Tier | Description | Analysis-Level Tier | Description
Basic | Mandatory for procedure registration. Must be provided to register a valid procedure. | Read-Only | Directly imported from the registered procedure; cannot be changed by the analyst. Changing this value means running a different procedure. Also used for fields relevant only at procedure level — the value is inherited from the procedure record and shown read-only in the analysis form.
Advanced | Optional for procedure registration. Strongly recommended but not required. | Editable | Imported from the registered procedure but may be adjusted within procedure-defined bounds (e.g., daily tuning, minor software updates). The procedure registers the target or typical value; the analyst confirms or adjusts it. Cannot be left void if the procedure-level tier is Basic.
N/A | Not applicable at procedure level. This field captures analysis-level information only. | Basic | Mandatory user input at analysis time. Value comes from the analysis itself and cannot be pre-specified in the procedure.
|  | Advanced | Optional user input at analysis time. Recommended for complete and reproducible documentation.
Mode Column | Definition
Purpose
Y | This field applies to this analytical mode and should be reported.
N | This field does not apply to this analytical mode.
Data Type | What to enter
Controlled list | Choose one of the values listed in "Example / Allowed Content". The list is CLOSED — it is meant to cover every case. If your value genuinely is not there, that is a gap in the list worth reporting, not a reason to write your own.
Controlled list / Text | Choose a listed value AND qualify it in the same cell — the detail is expected, not optional. Typically what was corrected, which target species or phases it applied to, the method or equation used, or the source it came from. An answer the list cannot express is also valid here.
Numeric (unit) | A number in the unit named in brackets. Enter the number only — the unit is fixed by the field.
Numeric + unit | A number AND its unit, because the unit varies between procedures. Write both, e.g. '50 us' or '0.5 s'.
Text (free) | Free text. No controlled vocabulary applies.
N/A ¦ None | Offered by every controlled list. "N/A" means the field does not apply to this procedure; "None" means it applies but nothing was used. Prefer either over leaving the cell blank.
Keyed By (Rule 7) | What the field's value repeats over
(none) | Scalar — one value per procedure or analysis. The default.
acquisition pass | One value per pass over the sample with its own instrument settings.
defines: acquisition pass | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per pass over the sample with its own instrument settings.
defines: monitored property per target species | ENUMERATES the monitored property domain — the header of the child table, not a column in it — and does so once per target species. One value per chemical species determined, at whatever granularity the procedure determines it.
defines: preparation step | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per stage in sample preparation.
defines: reported property | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
defines: sample | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it.
defines: sample > sampling unit | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per subdivision of the physical sample that carries its own row — grain, spot, aliquot, phase, sub-volume.
defines: standard | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per reference material or reference database entry.
defines: target material | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per material type the procedure is designed to analyse, as listed in Target Material.
defines: target species | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per chemical species determined, at whatever granularity the procedure determines it.
monitored property
preparation step | One value per stage in sample preparation.
reported property | One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
sample
standard x reported property | Cross-product, read as "for each standard, one value per reported property". One value per reference material or reference database entry. + One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
target material x target species | Cross-product, read as "for each target material, one value per target species". One value per material type the procedure is designed to analyse, as listed in Target Material. + One value per chemical species determined, at whatever granularity the procedure determines it.
target species | One value per chemical species determined, at whatever granularity the procedure determines it.