## TAPP
Metadata Item | Description | Procedure-Level Tier | Analysis-Level Tier | Data Type | Example / Allowed Content | Comments | Last Update | Keyed By | Purpose | Single-volume | Multi-volume stitching | Literature Assessment | Eckley 2024 (JSC Scan Record) Bennu particle Single-volume Nikon XTH 320 NASA JSC X-FaCT | Genge et al. 2025 (Nat. Commun.) Ryugu particle (A0180) Single-volume Zeiss Versa Not stated | Neuman et al. 2025 / Shearer et al. 2024 (JGR / Space Sci. Rev.) Apollo 17 core 73002 Multi-volume stitching NSI custom, UTCT | Neuman et al. 2025 (JGR Planets) Apollo 17 core 73001 Multi-volume stitching NSI custom, UTCT | Shearer et al. 2024 (Space Sci. Rev.) Apollo 17 73001 CSVC Single-volume Nikon XTH 320 NASA JSC | Shearer et al. 2024 (Space Sci. Rev.) Apollo 17 particles Single-volume Nikon XTH 320 NASA JSC X-FaCT | Tomkinson et al. 2015 (MAPS) NWA 5790 nakhlite Single-volume Nikon Metris XTH 225 U. Manchester | Glavin et al. 2023 (MAPS) Murchison CM2 Single-volume Nikon XTH 320 NASA JSC X-FaCT | Nascimento-Dias et al. 2019 (Appl. Radiat. Isot.) NWA 8277 + NWA 6963 Single-volume Bruker Skyscan 1173 | Richard et al. 2019 (Chem. Geol.) Olivine (melt incl.) Single-volume Zeiss Xradia 510 Versa UNAM Mexico | Richard et al. 2019 (Chem. Geol.) Synthetic quartz (fluid incl.) Whole sample (low-res) Nikon XTH 320/225 U. Strathclyde | Richard et al. 2019 (Chem. Geol.) Synthetic quartz (fluid incl.) ROI scan (high-res) Nikon XTH 320/225 U. Strathclyde | Richard et al. 2019 (Chem. Geol.) Fluid incl. minerals (C-I) Single-volume Phoenix Nanotom S U. Lorraine | Tait 2014 (Thesis) Watson 012 H7 chondrite Single-volume XRADIA XRM500 CSIRO Kensington
1. Procedure Identification | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Procedure Name | A short descriptive name for this analytical procedure, including a version number. Should identify the instrument, the technique, and the scope of what is measured — target species, material, or feature. | Basic | Read-Only | Text (free) | e.g., 'Lab-XCT Meteorite 3D Characterization v1.0' ¦ 'µXCT Lunar Core Multi-volume Stitching v1' | Source: Core module | 2026-09-01 | (none) |  | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Technique | Top-level analytical technique identifier. | Basic | Read-Only | Controlled list | Lab XCT ¦ Lab XCT (nano-CT) | Source: Core module | 2026-08-30 | (none) |  | Y | Y |  | Lab XCT | Lab XCT (nano-CT) | Lab XCT | Lab XCT | Lab XCT | Lab XCT | Lab XCT | Lab XCT | Lab XCT | Lab XCT (nano-CT) | Lab XCT | Lab XCT | Lab XCT (nano-CT) | Lab XCT (nano-CT)
Procedure Author | Person(s) or laboratory responsible for developing and registering this procedure. ORCID recommended for individuals. | Basic | Read-Only | Text (free) | e.g., 'J. Smith (ORCID: 0000-0001-2345-6789)' ¦ 'X-Ray Imaging Lab, University of X' | Source: Core module | 2026-07-24 | (none) |  | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Laboratory | Name of the laboratory or institution hosting the instrument. | Basic | Editable | Text (free) | e.g., 'UTCT Laboratory, University of Texas at Austin' ¦ 'NASA JSC Astromaterials X-ray CT Laboratory' | Source: Core module | 2026-05-15 | (none) |  | Y | Y |  | NASA JSC Astromaterials X-FaCT Lab | Not stated | UTCT Facility, U. Texas | UTCT Facility, U. Texas | NASA JSC Astromaterials X-FaCT Lab | NASA JSC Astromaterials X-FaCT Lab | Manchester X-ray Imaging Facility, U. Manchester | NASA JSC Astromaterials X-FaCT Lab | Nuclear Instrumentation Lab, COPPE, UFRJ, Brazil | Lab. de Microtomografía de Rayos X, UNAM, Mexico | U. Strathclyde, Glasgow, UK | U. Strathclyde, Glasgow, UK | Université de Lorraine, France | CSIRO, Kensington, Western Australia
Laboratory ID | Persistent identifier for the laboratory (e.g., ROR ID). | Advanced | Editable | URI / DOI | e.g., https://ror.org/00hj54h04 | Source: Core module | 2026-05-15 | (none) |  | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Procedure Start Date | First date this procedure configuration was used in production. | Basic | Read-Only | Date | YYYY-MM-DD, e.g., 2022-04-22 | Source: Core module | 2026-07-24 | (none) |  | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Funding Source for Procedure Development | Grants and other funding sources that supported instrument acquisition, major upgrades, procedure development, and associated personnel time. Include grant numbers and funding agencies where applicable. | Advanced | Read-Only | Text (free) | e.g., 'NSF EAR-2142500' ¦ 'NASA ROSES 80NSSC21K0153' ¦ 'None' | Source: Core module | 2026-07-24 | (none) |  | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Procedure Reference(s) | DOI or URL for peer-reviewed publications or technical reports describing, validating, or benchmarking this procedure. | Advanced | Read-Only | URI / DOI | e.g., doi:10.1111/maps.12888 ¦ doi:10.1016/j.chemer.2017.01.006 | Source: Core module | 2026-07-24 | (none) |  | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Procedure DOI | DOI or unique persistent identifier for the registered procedure used in this analysis, generated upon procedure submission to a data repository. If a DOI has been applied for but not yet minted, enter "pending". | N/A | Basic | URI / DOI | e.g., http://doi.org/10.60520/IEDA/114187 ¦ pending | Source: Core module | 2026-08-25 | (none) | Encourages formal procedure registration and ensures traceability. | Y | Y |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Session Identifier | Identifier for the analytical session this record describes — the laboratory's own run, sequence or batch identifier as generated by the instrument or acquisition software. Distinct from any persistent identifier a repository mints on submission. | N/A | Basic | Text (free) |  | Source: Core module | 2026-09-01 | (none) | The analysis record corresponds to one session, which may cover several samples, and this is the link back to the raw instrument files. | Y | Y |  | N (scan record dated 29 October 2024; operator named, no session identifier assigned) | N | N | N | N | N | N | N | N | N | N | N | N | N
Analyst | Name(s) of the analyst(s) who performed the analysis session. ORCID is recommended for persistent identification. | N/A | Basic | Text (free) | e.g., 'J. Smith (ORCID: 0000-0001-2345-6789)' | Source: Core module | 2026-08-25 | (none) |  | Y | Y |  | Scott Eckley | N | N | N | N | N | N | Scott A. Eckley | N | N | N | N | N | N
Analysis Start Date | Date on which the analytical session began. For sessions spanning multiple days, use the date of the first session. Format: YYYY-MM-DD. | N/A | Basic | Date | YYYY-MM-DD, e.g., 2022-04-22 | Source: Core module | 2026-08-25 | (none) |  | Y | Y |  | 2024-10-29 | N | N | N | N | N | N | N | N | N | N | N | N | N
Analysis End Date | Date on which the analytical session ended. May equal Analysis Start Date for single-day sessions. Format: YYYY-MM-DD. | N/A | Basic | Date | YYYY-MM-DD, e.g., 2022-04-22 | Source: Core module | 2026-08-25 | (none) |  | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Funding Source for Analysis | Grants and other funding sources that supported the specific sample analyses and data collection reported here. Include grant numbers and funding agencies. Distinct from procedure development funding, which may have been awarded years earlier. | N/A | Basic | Text (free) | e.g., 'NASA 80NSSC21K0153' ¦ 'None' | Source: Core module | 2026-08-25 | (none) |  | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Coupled Technique(s) | Other analytical techniques applied to the same sample(s) whose results are intended to be interpreted together with data from this procedure. Document coupling with any technique whose results are functionally linked to this dataset — providing calibration inputs, complementary spatial context, or required companion measurements. Name the instrument, laboratory or purpose alongside each technique where known, and join multiple techniques with '; '. Enter "None" if no coupling is intended. | Advanced | Editable | Text (free) | e.g., 'EPMA (Brown University CAMECA SX-100)' ¦ 'FIB-SEM (Crossbeam 540) for 3D tomography; EDS for mineral analysis' ¦ 'NanoSIMS; STXM; Raman' ¦ 'None' ¦ 'N/A' | Source: Core module | 2026-08-31 | (none) |  | Y | Y |  | N | SEM (post-XCT polished sections) | N | N | N | N | SEM-BSE (qualitative comparison) | N | Micro-XRF (same samples) | N | N | N | Confocal laser scanning microscopy (Sample I) | Optical microscopy; EBSD
Coupling Description | Description of how this procedure is coupled with the technique(s) listed above. Include: (1) the functional relationship — what data or context flows between techniques, or how results are combined (e.g. which output from the coupled technique serves as input to data reduction for this technique); and (2) the analytical sequence — which technique is performed first and why (e.g. non-destructive before destructive). Required when Coupled Technique(s) is not "None". | Advanced | Editable | Text (free) | e.g., 'NCT provides complementary attenuation contrast sensitive to H, Cl, and Ir, and discriminates Fe from Ni — information inaccessible to X-ray CT; XCT and NCT performed sequentially on the same sample. If thermoluminescence (TL) analyses are also planned, TL must be performed before any tomography. Note: NCT produces short-lived radioactivation (e.g., ²⁴Na, t₁/₂ = 15 h); hold samples at facility until below shipping limits.' ¦ 'EPMA provides phase identification used to guide XCT VOI selection; EPMA performed after XCT (non-destructive) and before polishing' | Source: Core module | 2026-08-25 | (none) |  | Y | Y |  | N | Serial polished sections cut post-XCT; SEM imaging | N | N | N | N | BSE images compared with XCT attenuation contrast | N | Micro-XRF performed on same meteorite fragments after XCT | N | N | N | CLSM performed on same inclusion for volumetric cross-validation | XCT qualitatively compared with thin sections and EBSD
Coupled Procedure DOI | Registered procedure DOI for the coupled technique named above. If the coupled procedure has not yet been registered, enter the DOI of a publication describing the coupled method, or "pending". Enter "None" if no coupling is planned. | N/A | Advanced | URI / DOI | e.g., 'http://doi.org/10.60520/IEDA/114187' ¦ pending ¦ None | Source: Core module | 2026-08-25 | (none) | Provides a stable, citable link to the companion method independent of whether a dataset has been deposited. | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Coupled Dataset or Publication Reference | DOI or other persistent identifier for the co-registered dataset or publication where both datasets are reported together. Accepts: a dedicated dataset DOI (if separately deposited), a shared dataset DOI (if co-submitted in the same package), or a publication DOI. Use "same submission" if the coupled dataset is included in this data package, or "pending" if not yet assigned. If coupling is documented through a shared sample identifier only, that information is already captured in Sample Persistent Identifier (Group 2). | N/A | Advanced | URI / DOI / Text | e.g., 'doi:10.26022/IEDA/112969' ¦ same submission ¦ pending ¦ None | Source: Core module | 2026-08-25 | (none) |  | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
2. Samples | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Target Material | General description of the material type(s) this procedure is designed to analyse. | Basic | Read-Only | Controlled list / Text | Chondrite meteorite ¦ Achondrite meteorite ¦ Lunar sample ¦ Mission-returned sample ¦ Drill core (rock) ¦ Sediment core ¦ Terrestrial rock ¦ Cosmic spherule ¦ N/A ¦ None | Source: Core module | 2026-09-28 | defines: target material | Used for discoverability and procedure matching, and because the material type constrains sample preparation, calibration and matrix-matching requirements. | Y | Y |  | C-type asteroid particle (Bennu, OSIRIS-REx) | C-type asteroid particle (Ryugu, Hayabusa2) | Lunar regolith (Apollo 17 drive-tube core) | Lunar regolith (Apollo 17 drive-tube core, steel sleeve) | Lunar regolith (Apollo 17 drive-tube CSVC) | Lunar regolith particles (Apollo 17 core) | Martian meteorite (nakhlite, NWA 5790) | CM2 carbonaceous chondrite (Murchison) | Lunar meteorite (NWA 8277); Martian meteorite (NWA 6963) | Olivine phenocryst with silicate melt inclusion | Synthetic quartz monocrystal with aqueous fluid inclusions | Synthetic quartz monocrystal with aqueous fluid inclusions | Minerals with fluid inclusions (quartz, garnet, emerald, wolframite, feldspar) | H7 ordinary chondrite (Watson 012)
Target Feature(s) | The 3D internal features or properties the procedure is designed to characterize. | Basic | Read-Only | Text (free) | e.g., 'Vesicle network and size distribution' ¦ 'Opaque phase (FeNi metal, sulfide) 3D distribution' ¦ 'Porosity and pore connectivity' ¦ 'Mineral phase boundaries and modal mineralogy' ¦ 'General 3D structural characterization' | Note for Phase 3 literature assessment: evaluate whether procedures in the literature are feature-specific (acquisition parameters tuned to a specific feature type, e.g., opaque-phase contrast) or general-purpose (designed to capture all structural features in a given sample type). This will inform whether the field warrants a controlled list or remains free text. | 2026-08-27 | (none) | Defines the scientific purpose of the scan and informs resolution and contrast requirements. Used for procedure discoverability: a researcher can identify whether this procedure is suitable for their scientific question. | Y | Y |  | Internal structure; mineralogy (reconnaissance scan) | Microchondrules (SSOs); 3D volume fraction | Voids, lithic clasts, stratigraphic layers in core | Voids, lithic clasts, stratigraphic layers in core (steel-sleeved) | Internal structure of core vacuum seal container | Individual particle internal structure | Modal mineralogy (vol%); secondary minerals | Radiation dose assessment (not mineralogy) | Internal structure; porosity; density contrast | Silicate melt inclusion morphology and phase volumes | Fluid inclusion morphology and phase volumes | Fluid inclusion morphology and phase volumes | Fluid inclusion morphology and phase volumes | Plagioclase network connectivity; partial melt evidence
Applicable Sample Dimension Range | The range of sample sizes compatible with this procedure at the target voxel size and geometric configuration. The upper bound is a hard constraint set by the instrument field of view (FOV) at the procedure's target magnification: FOV ≈ detector array width × pixel size / geometric magnification (divide additionally by optical objective multiplier for Versa-class systems using optical magnification). For Mode A (single-volume): all three sample dimensions must fall within the FOV. For Mode B (multi-volume stitching): sample diameter must fall within the FOV; sample length along the rotation axis is effectively unlimited through stitching. The lower bound is a practical guideline: ~10 voxels across the smallest dimension. | Basic | Read-Only | Text (free) | e.g., 'Single-volume: max 20 mm in all dimensions at 7.6 µm voxel' ¦ 'Multi-volume stitching: max 25 mm diameter, unlimited length at 12.9 µm voxel' ¦ 'Max 5 mm (all axes) at 0.85 µm voxel with 4× objective' |  | 2026-08-27 | (none) | Samples spanning fewer than that are difficult to characterize reliably. | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Sample Dimensions (L × W × H) | Physical dimensions of the sample in mm, reported as length × width × height (or equivalent three orthogonal measurements). | Advanced | Basic | Text (free) | e.g., '12 × 8 × 5 mm' ¦ '2.844 × 1.994 × 1.38 mm' ¦ '350 × 25 × 25 mm (drive tube)' |  | 2026-08-27 | (none) | The maximum dimension constrains the minimum achievable field of view and therefore the coarsest achievable voxel size. The minimum dimension determines how many voxels span the smallest feature of interest. | Y | Y |  | N | 1.592 × 0.756 × 0.985 mm | ~35 cm length core | ~35 cm length core | N | N (multiple particles) | ~1.1 × 1.2 × 0.8 cm | N (powder in 12 mL glass vial) | ~4 mm fragments | ~1 mm olivine | 3 × 5 × 2 cm | 3 × 5 × 2 cm | Varies: 0.3×0.4 to 11×11×7 mm | 8 mm diameter core
Sample Mass | Mass of the sample in grams, recorded for curatorial tracking of precious or limited materials. | N/A | Advanced | Numeric (g) | e.g., 5.59 g ¦ 0.003 g ¦ 19.64 g |  | 2026-08-27 | (none) | Particularly important for meteorite, mission-returned, and other restricted samples. The actual sample mass depends on the specific sample being scanned, not on the procedure design. | Y | Y |  | N | N | N | N | N | N | 2.7 g | 4.6430 g (Murchison B in vial) | N | N | N | N | N | N
Sample Mounting Method | Method used to mount or hold the sample on the instrument rotation stage. Mounting material should transmit X-rays at the selected voltage without dominating beam attenuation. Report the holder CLASS from the list and name the specific vessel or material alongside it — 'Tube or vial — 1 cm plastic straw', not 'Tube or vial'. Where the sample is sealed or bagged inside a further holder for contamination control, record both layers. Report any adhesive, support material and alignment aids used. | Advanced | Editable | Controlled list / Text | Free-standing on stage pin ¦ Free-standing without holder ¦ Tube or vial ¦ Glass capillary ¦ Sealed bag or wrap ¦ Embedded or adhered ¦ On a flat support or substrate ¦ Fixed to the stage or gantry ¦ N/A ¦ None |  | 2026-08-30 | (none) |  | Y | Y |  | N | Mounted in pipette tips | Plexiglass tube (vertical); triple-sealed Teflon bag | Custom PVC tube | N | Triple-bagged in Teflon, wrapped in cylinder, placed in 1-cm plastic straw | N | Glass vial placed in scanner | Polystyrene support | N | N | N | N | N
Sample Name | Name or identifier of each sample analysed in this session, as used in the laboratory — a sample mount, section or aliquot counts as one entry where that is what the laboratory tracks. Should match the identifier used in associated publications or data tables. | N/A | Basic | Text (free) | e.g., 'NWA 10597' ¦ 'Apollo 17 sample 71035,36' ¦ 'Ryugu A0159' | Source: Core module | 2026-08-25 | defines: sample | The analysis record corresponds to one session and may cover several samples. | Y | Y |  | OREX-800099-0 | A0180 (A0180-A and A0180-B) | Apollo 17 core 73002 | Apollo 17 core 73001 | 73001 CSVC | Multiple particles from 73001 and 73002 | NWA 5790 | Murchison B (USNM 5453,1) | NWA 8277; NWA 6963 | Sample A (olivine) | Sample B (synthetic quartz) | Sample B (synthetic quartz) | Samples C–I (various) | Watson 012
Sampling Unit Type | The kind of physical subdivision of the sample to which one row of reported values corresponds — the unit that is analysed and reported, as distinct from the sample as a whole. Where units nest (e.g. confined tracks within grains), state both levels. The units themselves are listed in Sampling Unit Name. | Basic | Read-Only | Controlled list / Text | Whole sample ¦ Aliquot ¦ Grain ¦ Spot ¦ Analysis point ¦ Phase ¦ Sub-volume ¦ Region of interest ¦ N/A ¦ None | Source: Core module | 2026-09-15 | (none) |  | Y | Y |  | Whole sample (single allocated Bennu particle; 928 slices) | Sub-volume > Grain (sub-samples A0180-A and A0180-B; size and shape factor reported per microchondrule / sulphide-silicate object) | Whole sample (six overlapping cone-beam volumes stitched into one continuous dataset per core; the sub-volumes are an acquisition unit, not a reporting unit) | Whole sample (six overlapping cone-beam volumes stitched into one continuous dataset per core; the sub-volumes are an acquisition unit, not a reporting unit) | Whole sample (the 73001 CSVC container assembly) | Grain (individual extracted particles >4 mm, each individually bagged and scanned) | Sub-volume (modal mineralogy reported for six 2-D XCT slices at ~1 mm spacing and for the entire chip volume) | Aliquot (~1 g crushed Murchison B in a glass vial) | Whole sample (one scan per meteorite specimen: NWA 8277, NWA 6963) | Region of interest (individual melt inclusion) > Phase (glass, clinopyroxene, spinel, vapour) | Region of interest (individual fluid inclusion) > Phase (vapour, liquid) | Region of interest (individual fluid inclusion) > Phase (vapour, liquid) | Region of interest (individual fluid inclusion) > Phase (vapour, liquid, oil, solid bitumen) | Whole sample (8 mm core) > Phase (plagioclase network)
Sampling Unit Name | The name or label of each sampling unit analysed in this session, as the laboratory records it, together with the sample it belongs to — e.g. a spot number, a grain label, a map or region-of-interest name, or an aliquot identifier. Where units are too numerous to name individually, such as map pixels or reconstructed voxels, name the acquisition area they belong to instead. | N/A | Basic | Text (free) | e.g., NWA 8657-1: spots 1–24 ¦ Zircon mount ZM-3: grains Z1–Z30 ¦ Core 73001: map areas A1–A5 ¦ Bennu aggregate OREX-803015-100: aliquot for Ti | Source: Core module | 2026-09-15 | defines: sample > sampling unit | Gives every field keyed by sampling unit a row to attach its values to, and ties each unit to the sample it belongs to; without it those fields collapse to one value per session. | Y | Y |  | Sample name only — the scan record's "Sample name: OREX-800099-0", "Sample type: Asteroid Bennu particle" (p.1); one whole-particle scan | Labelled: scans "A0180-A and A0180-B" of Ryugu sample A0180 (p.7) | Sample name only — core 73002, "A series of six volumes was acquired" and stitched "to create a continuous data set for each core" (Neuman et al. 2025, p.4); the volumes are counted, not labelled | Sample name only — core 73001, "A series of nine scan volumes was acquired along the length of the core" (p.5); the volumes are counted, not labelled | Sample name only — the 73001 CSVC; scans are identified by position and stage ("the bottom portion of 73001 still within the CSVC before (b) and after (c) piercing", "the top portion", Fig. 7, p.27), not labelled | Sample name only — "132 particles were scanned from 73002 ... and 220 particles from 73001" (p.28), each "individually bagged"; the particles are counted and shown by lithology (Fig. 8, p.28), not labelled | Labelled: 2-D slices of the one scanned volume — "XCT slice 500" to "XCT slice 1000" (Table 1, p.2) — of "a single 2.7 g chip" of NWA 5790 (p.3) | Labelled: the split "Murchison B (mass = 4.6430 g)", scanned in its sealed vial; its counterpart "Murchison A" was the unscanned control (p.3) | Sample name only — "fragments of about 4 mm of both meteorites", NWA 8277 and NWA 6963 (p.5); not otherwise labelled | Labelled: "a single olivine phenocryst (1.0 mm large) separated from Sample A" (p.2), hosting one silicate melt inclusion (Fig. 1, p.4) | Labelled: "Sample B was scanned both entirely and on a 1.4 × 1.4 × 1.4 mm region of interest" (p.2) — this column is the whole-sample scan; its inclusions are numbered ("#3", Fig. 2, p.5) | Labelled: the region-of-interest scan of Sample B, "shown in Fig. 2 (fluid inclusion #3)" (Table 1 note, p.3) | Labelled: "Sample #" C, D, E, F, G, H and I, each scanned on the Phoenix Nanotom S (Table 1, p.3) | Sample name only — "An 8 mm diameter core was drilled from the Watson 012 sample and then scanned" (p.9); not otherwise labelled
Sample Persistent Identifier | Globally unique, persistent identifier for each sample listed in Sample Name. IGSN (International Geo Sample Number) is the recommended standard for geological and cosmochemical samples, as used by Astromat, EarthChem and SESAR. Where a sample and its sub-samples are separately registered, record the identifier at the level actually analysed. | N/A | Advanced | URI / IGSN | e.g., IGSN: IEDA-100001 (https://www.igsn.org/) ¦ Smithsonian NMNH catalog number ¦ JAXA Hayabusa2 sample ID | Source: Core module | 2026-09-15 | sample |  | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Sample Preparation Notes | Any preparation steps applied to the sample before scanning, including cleaning, trimming, consolidation, or drying. Note any exceptions. | Advanced | Editable | Text (free) | e.g., 'None — scanned as received' ¦ 'Fragment trimmed to fit within FOV' ¦ 'Wrapped in PTFE tape to stabilize loose grains' ¦ 'Dried at 60°C for 24 h before scanning' |  | 2026-08-27 | (none) | XCT is typically non-destructive, with no surface preparation required. | Y | Y |  | N | Decanted for XCT; sample split along fractures during mounting into pipette tips. Post-XCT: embedded in Specifix resin, polished with 0.1 µm Al₂O₃. | Al inner sleeve triple-sealed in Teflon; stainless-steel outer sleeve removed prior to scan | Steel outer sleeve retained for scan | N | Triple-bagged (Teflon) per ANGSA curation protocol | None; chip used as received | Chips (~10.3 g) crushed with mortar and pestle; vortex mixed 3 min; split into two ~4.6 g portions | None stated (fragments purchased from IMCA member) | None stated for XCT; embedded in epoxy after XCT for microprobe | None stated | None stated | None stated | 8 mm diameter core drilled from meteorite prior to XCT
Sample Preparation Method | The form in which the sample is presented to the instrument, and the preparation that brought it to that form — for example mounting, sectioning, polishing, coating, crushing, fusion, or extraction of an electron-transparent section. Record 'None' where the material is analysed as received. | Basic | Editable | Controlled list / Text | Bulk specimen or fragment (as received) ¦ Core or trimmed billet ¦ Powder or crushed split ¦ Mounted in tube, straw or pipette tip ¦ Sealed or bagged for containment ¦ Polished block or epoxy mount ¦ N/A ¦ None | Source: Core module | 2026-09-28 | sample |  | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Sampling Unit Selection Criteria | The rules governing which sampling unit(s) within a sample are selected for analysis, and why. Covers the criteria applied when choosing grains, aliquots, spots, or a region of interest — size, morphology, clarity, freedom from inclusions or alteration, phase identity, or spatial position. Distinct from Target Material, which states the material type the procedure is designed for, and from Sampling Unit, which names the kind of subdivision one row of reported values corresponds to: this field states how, within such a sample, the unit actually analysed is picked out. | Basic | Editable | Text (free) |  | Source: Sampling Unit Selection module | 2026-09-01 | (none) |  | Y | Y |  | N (single allocated particle; no selection criteria stated) | N (whole particle scanned; the sample split into parts A and B along fractures during mounting, not by selection) | N | N | N (whole CSVC assembly scanned) | Particles >4 mm extracted during dissection are individually bagged and XCT scanned for classification and characterization, without destructive chipping, sectioning or dust removal | Single 2.7 g chip (~1.1 x 1.2 x 0.8 cm) taken from the outer part of one stone, selected because it provides a profile from the weathered exterior to the fresh interior; six 2-D slices at ~1 mm spacing extracted for modal analysis | N (whole crushed aliquot scanned in its vial) | N | N | No sectioning was carried out prior to HRXCT scanning; Sample B was scanned entirely | 1.4 x 1.4 x 1.4 mm region of interest containing fluid inclusion #3, scanned at higher resolution after the whole-sample scan | N | N
Pre-Analysis Imaging and Screening | Imaging or other characterisation performed before the measurement in order to select or locate the sampling unit to be analysed, including the technique, instrument and settings used, and how individual analyses are linked back to the images. Distinct from any imaging the procedure performs as its own measurement. Where the imaging is performed on a separate instrument, it should also be recorded in the Group 1 coupling fields. | Advanced | Editable | Text (free) |  | Source: Sampling Unit Selection module | 2026-09-01 | sample |  | Y | Y |  | N | Optical microscopy — morphological inspection of the sample exterior through the container window with the CLOXS digital optical microscope system on automated digital sample stages at JAXA/ISAS, before the sample was decanted for nano-XCT | N | N | N | N | N (XCT is itself the pre-screening step here: the chip was characterized in 3-D prior to destructive sampling) | N | N | N (epoxy mounting and secondary electron imaging were carried out after HRXCT, for subsequent microprobe work) | Optical photography of the sample showing the fluid inclusion array, with inclusions numbered for correspondence with the HRXCT reconstruction (Fig. 2) | Optical photography identifying the fluid inclusion array; inclusion #3 targeted for the high-resolution region-of-interest scan (Fig. 2) | Optical microscopy under UV illumination, used to identify hydrocarbon-bearing phases within the inclusions (Fig. 9) | N
3. Instrument & Software | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Instrument Manufacturer | Manufacturer of the instrument that performs the measurement, recorded as a controlled value. Where a procedure couples a sample-introduction system to an analysing instrument, this records the analysing instrument. Instrument Model gives the specific designation. | Basic | Read-Only | Controlled list | Nikon ¦ Zeiss ¦ Bruker ¦ GE / Waygate ¦ North Star Imaging ¦ RX Solutions ¦ Custom-built ¦ Unknown ¦ N/A ¦ None | Source: Core module | 2026-09-01 | (none) | Recording it as a controlled value lets procedures be found by vendor. | Y | Y |  | Nikon | Zeiss | North Star Imaging | North Star Imaging | Nikon | Nikon | Nikon | Nikon | Bruker | Zeiss | Nikon | Nikon | GE / Waygate (stated as "Phoenix Nanotom S") | Zeiss (stated as "XRADIA XRM500")
Instrument Model | Model designation of the instrument that performs the measurement, including any generation or configuration suffix. Conventionally written with the manufacturer name included; Instrument Manufacturer records the vendor separately, as a controlled value, so that procedures remain findable by vendor. | Basic | Read-Only | Text (free) | e.g., 'Nikon XTH 320/225' ¦ 'Zeiss Xradia Versa 520' ¦ 'Bruker SkyScan 1173' ¦ 'Custom NSI system (FeinFocus FXE 225.48 source + Perkin Elmer 2048×2048 flat-panel), UTCT Lab' | Source: Core module | 2026-08-25 | (none) |  | Y | Y |  | Nikon XT H 320 | Zeiss Versa (model not specified) | Custom NSI instrument (North Star Imaging) at UTCT | Custom NSI instrument (North Star Imaging) at UTCT | Nikon XTH 320 | Nikon XTH 320 | Nikon Metris XTH 225 | Nikon XTH 320 | Bruker/Skyscan 1173 | Zeiss Xradia 510 Versa | Nikon XTH 320/225 | Nikon XTH 320/225 | Phoenix Nanotom S | XRADIA XRM500
X-ray Source Configuration | Focal spot geometry and scale of the X-ray tube. | Basic | Read-Only | Controlled list / Text | Reflection target, microfocal ¦ Reflection target, rotating (high-power) ¦ Transmission target, nanofocal ¦ N/A ¦ None |  | 2026-08-30 | (none) | Reflection target configurations are most common in lab XCT (X-rays emitted from the face of the anode toward the sample); transmission target configurations allow smaller focal spots by thinning the anode into a membrane and collecting X-rays from the back, enabling higher spatial resolution at the cost of lower maximum power. | Y | Y |  | 180 kV transmission source | N | Feinfocus FXE 225.48 microfocal source | Feinfocus FXE 225.48 microfocal source | 225 kV multi-metal reflection target source | 180 kV nano-focus transmission target source | N | 180 kV nano-focus tungsten transmission source (~1 µm spot size) | N | N | Microfocus multi-metal target (225 kV) | Microfocus multi-metal target (225 kV) | N | N
X-ray Tube Anode Material | Material of the X-ray tube anode (target). | Advanced | Read-Only | Controlled list | Tungsten (W) ¦ Multi-metal (W/Mo/Cu) ¦ Silver (Ag) ¦ Molybdenum (Mo) ¦ Copper (Cu) ¦ N/A ¦ None |  | 2026-08-30 | (none) | Determines the characteristic line energies superimposed on the bremsstrahlung continuum, which can affect phase contrast at specific energies. Multi-metal anodes (e.g., W/Mo/Cu) extend the usable energy range. | Y | Y |  | Tungsten | Tungsten | N | N | N (multi-metal target) | Tungsten | Copper | Tungsten | N | N | Silver (for Sample B) | Silver (for Sample B) | N | N
X-ray Detector Type | Type of X-ray detector used to record projection images. | Basic | Read-Only | Controlled list / Text | Flat-panel (amorphous silicon, a-Si) ¦ CCD + scintillator (with optical objective) ¦ CMOS + scintillator ¦ N/A ¦ None |  | 2026-08-30 | (none) | Flat-panel detectors (amorphous silicon) are standard in Nikon, NSI, and similar systems. CCD or CMOS detectors coupled to a scintillator via an optical objective are used in Zeiss Xradia Versa systems, enabling optical magnification for sub-micrometre voxel sizes. | Y | Y |  | N | 2000 × 2000 CCD plane (16-bit) | 2048 × 2048 Perkin Elmer flat panel (4096 × 4096 effective via Subpix) | 2048 × 2048 Perkin Elmer flat panel (4096 × 4096 effective via Subpix) | N | N | N | 2000 × 2000 pixel Perkin Elmer flat panel CCD | Flat panel detector (2240 × 2240 pixels) | CCD ANDOR camera (1080 × 1080 px, 32-bit) | Flat panel photodetector (2000 × 2000 px, cell size 0.2 × 0.2 mm) | Flat panel photodetector (2000 × 2000 px, cell size 0.2 × 0.2 mm) | CCD Hamamatsu (2300 × 2300 px) | N
Detector Array Size | Number of pixels in the detector array (width × height). | Advanced | Read-Only | Text (free) | e.g., '2048 × 2048 pixels' ¦ '2000 × 2000 pixels' ¦ '2240 × 2240 pixels' |  | 2026-08-27 | (none) | Combined with detector pixel size and geometric magnification, determines the field of view. | Y | Y |  | N | 2000 × 2000 pixels | 2048 × 2048 (physical); 4096 × 4096 (effective, NSI Subpix) | 2048 × 2048 (physical); 4096 × 4096 (effective, NSI Subpix) | N | N | N | 2000 × 2000 pixels | 2240 × 2240 pixels | 1080 × 1080 pixels | 2000 × 2000 pixels | 2000 × 2000 pixels | 2300 × 2300 pixels | N
Detector Pixel Size | Physical size of a single pixel on the detector in micrometres. This is a fixed hardware property of the detector array, distinct from the reconstructed voxel size (which depends on geometric magnification and binning). | Advanced | Read-Only | Numeric (µm) | e.g., 127 µm ¦ 6.5 µm ¦ 12.7 µm ¦ 200 µm |  | 2026-05-15 | (none) |  | Y | Y |  | N | N | N | N | N | N | N | N | N | N | 0.2 × 0.2 mm cell size | 0.2 × 0.2 mm cell size | N | N
Optical Objective | Optical magnification objective installed between the scintillator and CCD/CMOS detector. Applies only to instruments using optical magnification (e.g., Zeiss Xradia Versa series). For systems using only geometric magnification, report "Not applicable". | Advanced | Read-Only | Text (free) | e.g., '4× objective' ¦ '20× objective' ¦ 'Not applicable — geometric magnification only' |  | 2026-05-15 | (none) |  | Y | Y |  | N/A | 4× objective lens | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 4.0× magnification | N/A | N/A | N/A | N/A
Acquisition Software | Instrument control, automation and data acquisition software used to collect the raw data, including version number. Distinct from Data Processing Software(s), which covers everything applied to the data after acquisition. | Basic | Editable | Text (free) | e.g., 'CT Agent Pro 3D v3.8 (Nikon)' ¦ 'Zeiss Scout-and-Scan Control System v14' ¦ 'NRecon SkyScan v1.6.9.4 (Bruker)' | Source: Core module | 2026-08-25 | (none) |  | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Reconstruction Software | Software used to reconstruct 2D projection images into a 3D CT volume, including version number. | Basic | Editable | Text (free) | e.g., 'CT Pro 3D v3.8 (Nikon)' ¦ 'Zeiss Reconstructor v14' ¦ 'NRecon v1.6.9.4 (Bruker)' ¦ 'Octopus Reconstruction v8.9' |  | 2026-08-27 | (none) | Often bundled with the instrument and proprietary to the manufacturer. | Y | Y |  | Nikon CTPro3D v5.4 | Zeiss Reconstructor software | N (UTCT custom; software not named) | N (UTCT custom; software not named) | N | N | Nikon proprietary filtered back projection (software not named) | Nikon CTAgentPro v5.4 (FBP algorithm) | NRecon SkyScan v1.6.9.4 | ZEISS XRM software | CT Pro 3D software | CT Pro 3D software | N | N
Data Processing Software(s) | All software applied to the data after acquisition in order to produce the reported quantities, including version numbers. List every package used. Distinct from Acquisition Software, which controls the instrument and collects the raw data. | Basic | Editable | Text (free) | e.g., 'Avizo 9.2 (Thermo Fisher)' ¦ 'Dragonfly 2021.1 (ORS Inc.)' ¦ 'Blob3D v2019 (UT Austin)' ¦ 'Fiji/ImageJ + 3D Suite plugin' ¦ 'VGStudio Max 3.4 (Volume Graphics)' ¦ 'CTAn v1.20 (Bruker)' | Source: Core module | 2026-08-25 | (none) |  | Y | Y |  | N | ImageJ + TANGO plugin (Ollion et al. 2013) | N | N | N | N | Avizo™ | N (3D rendering and TIFF export only) | DataViewer; CTVox; CTAn (Bruker proprietary suite) | ImageJ; Avizo 9.2 | Avizo 9.2.0 | Avizo 9.2.0 | Avizo 9.2.0 | AvizoFire 8.0; Matlab; Drishti 2.0
4. Measurement Information | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Analytical Mode | Primary analytical mode(s) executed under this procedure. Records one value for a single-mode procedure and all applicable modes for a multi-mode one. Applies even where a procedure defines no mode flag columns. | Basic | Read-Only | Controlled list | Single-volume ¦ Multi-volume stitching | Source: Core module | 2026-09-09 | (none) | Serves as the procedure-level declaration of measurement type, distinct from the mode flag columns, which state per-field applicability. Ensures every procedure record is self-describing and comparable across the library. | Y | Y |  | Single-volume | Single-volume | Multi-volume stitching | Multi-volume stitching | Single-volume | Single-volume | Single-volume | Single-volume | Single-volume | Single-volume | Single-volume | Single-volume | Single-volume | Single-volume
Reported Variables and Units | The final variable(s) this procedure reports and their units — distinct from the fields recording what was *acquired* rather than what is reported. Record every reported variable, including intermediate quantities reported alongside final ones (e.g. both the 206Pb/238U ratio and the 206Pb/238U date). Where a reported variable is a nominal property with no magnitude (e.g. a mineral identification), record the variable and give the unit as 'N/A — nominal property'. | Basic | Read-Only | Text (free) | e.g., 'Modal mineralogy (vol%: augite, mesostasis, olivine, titanomagnetite)' ¦ 'Porosity (vol%); bulk density (g cm-3)' ¦ 'Inclusion volume (mm3); vapour volumetric fraction (%)' ¦ 'Equivalent diameter (um); shape factor (dimensionless)' ¦ 'Phase interconnectivity (N/A - nominal property)' ¦ 'Absorbed x-ray dose (Gy)' | Source: Core module | 2026-09-01 | defines: reported property | A procedure may acquire many channels and report a small number of derived quantities; without this field a data consumer cannot tell which. | Y | Y |  | N (instrument and scan-parameter record; no derived variables reported) | microchondrule and sulphide-silicate object diameter (µm); shape factor; object abundance; particle volume — Microchondrule / sulphide-silicate object diameter (um) and shape factor; object abundance; particle volume | N | N | Space between the bottom tip of the CSVC and the Teflon cap; Teflon cap location and integrity (nominal) | N (particles scanned for classification and characterization; no quantitative variables stated) | Modal mineralogy (vol%) for augite, mesostasis, olivine and titanomagnetite; slice area (mm2) | Absorbed x-ray dose (~180 Gy, the maximum a Bennu sample would receive during an XCT imaging experiment) | Density and porosity; proportions, volume, size, shape and spatial distribution of internal structure | Phase volumes within the melt inclusion (glass, clinopyroxene, spinel, vapour) | Total fluid inclusion volume (mm3); vapour volume (mm3); liquid volume (mm3); vapour volumetric fraction phi_vap (%) | Total fluid inclusion volume (mm3); vapour volume (mm3); liquid volume (mm3); vapour volumetric fraction phi_vap (%) | Total fluid inclusion volume (mm3); phase volumes (mm3); vapour volumetric fraction phi_vap (%) | Plagioclase network interconnectivity
Accelerating Voltage | X-ray tube accelerating voltage in kilovolts (kV). If the voltage was varied across samples within the same procedure or session, report the full range applied (e.g., 90–115 kV). Note: some sources report this parameter as 'X keV' (maximum Bremsstrahlung photon energy) rather than 'X kV' (tube voltage). For polychromatic lab XCT, these are numerically equivalent: E_max [keV] = V [kV]. Record the value as originally reported, and add a parenthetical note if the unit used is keV. | Basic | Editable | Numeric (kV) | e.g., 80 kV ¦ 130 kV ¦ 160 kV ¦ 190 kV; for variable settings: 90–115 kV |  | 2026-08-27 | (none) | Determines the maximum X-ray photon energy and controls penetration depth and phase contrast. Higher voltages provide greater penetration for dense or large samples; lower voltages improve contrast between low-density phases. | Y | Y |  | 110 kV | 90 kV | 180 kV | 190 kV | 215 kV | 90 kV (typical; optimized per sample) | 120 keV (reported in paper as 'accelerating voltage of 120 keV'; likely typo for 120 kV) | 160 keV (paper reports as X-ray photon energy; equivalent to 160 kV tube voltage) | 50 kV | 30 kV | 160 kV | 160 kV | 90–115 kV (varies by sample) | 70 kV
Tube Current | X-ray tube current in microamperes (µA). If the current was varied across samples within the same procedure or session, report the full range applied (e.g., 65–115 µA). | Basic | Editable | Numeric (µA) | e.g., 89 µA ¦ 108 µA ¦ 180 µA; for variable settings: 65–115 µA |  | 2026-08-27 | (none) | Controls photon flux and therefore projection image SNR. Higher current improves SNR but increases thermal loading on the source and may increase focal spot size. | Y | Y |  | 27 µA | 89 µA | 0.18 mA (180 µA) | N | 179 mA (possibly typo for µA per source) | N | N | 38 µA | 160 µA | N | 71 µA (Ag target) | 46 µA (Ag target) | 65–115 µA (varies by sample) | 86 µA
X-ray Power | X-ray tube power in watts (W). Derivable as voltage (kV) × current (mA) = kV × µA / 1000. If power was varied across samples within the session, report the full range applied (e.g., 7–13 W). | Advanced | Editable | Numeric (W) | e.g., 9 W ¦ 14 W ¦ 25 W; for variable settings: 7–13 W |  | 2026-08-27 | (none) | Power constrains the achievable focal spot size for microfocal tubes: higher power requires a larger focal spot, trading spatial resolution for SNR. Retained as a standalone field because it is routinely reported in the literature and directly constrains instrument operating limits. | Y | Y |  | 3.0 W | N | N | N | N | 3.0 W | N | N | N | N | 11.4 W | 7.4 W | N | 7 W
X-ray Pre-filter | Material and thickness of the beam-hardening filter placed between the X-ray source and the sample. Both filter material and thickness must be reported. Instrument-proprietary filter codes should be decoded where possible. | Basic | Editable | Text (free) | e.g., '1 mm Cu' ¦ '0.72 mm Al' ¦ '2 mm Al' ¦ 'LE2 (calcium fluoride, CaF₂, Zeiss)' ¦ 'HE1 (copper composite, Zeiss)' ¦ 'None' |  | 2026-08-27 | (none) | Hardens the beam by attenuating low-energy photons, reducing beam hardening artifacts and improving CT number stability across the sample. | Y | Y |  | 0.25 mm aluminum | Inbuilt LE4 filter (beam hardening reduction) | 0.72 mm Al | None (steel sleeve acts as effective filter) | N | 0.1–0.25 mm Al (Bennu PE protocol range stated) | N | None (intentionally unfiltered; worst-case dose experiment) | 1.0 mm aluminum | N | None (conditions did not require filtering) | None (conditions did not require filtering) | N | N
Source-to-Object Distance (SOD) | Distance from the X-ray source focal spot to the centre of the sample rotation axis, in mm. | Advanced | Editable | Numeric (mm) | e.g., 12 mm ¦ 55 mm ¦ 83 mm ¦ 150 mm |  | 2026-08-27 | (none) | Together with SDD, determines geometric magnification (M = SDD/SOD). Shorter SOD increases magnification and reduces voxel size. | Y | Y |  | 20.39 mm | N | N | N | N | N | N | 39.22 mm | N | N (optimized for 2.06 µm/px; exact value not stated) | N | N | 6–55 mm (varies by sample) | N
Source-to-Detector Distance (SDD) | Distance from the X-ray source focal spot to the detector surface, in mm. Voxel size ≈ detector pixel size / M (before binning; divide additionally by optical objective for Versa-class systems). | Advanced | Editable | Numeric (mm) | e.g., 250 mm ¦ 400 mm ¦ 900 mm ¦ 1000 mm |  | 2026-08-27 | (none) | Together with SOD, determines geometric magnification M = SDD/SOD. | Y | Y |  | 679.51 mm | N | N | N | N | N | N | 679.51 mm | N | N | N | N | N | N
Voxel Size | Isotropic voxel edge length of the reconstructed 3D volume in micrometres. The target voxel size is set based on the smallest feature to be resolved (target voxel size ≤ ~1/3 of that feature size; see the criterion recorded under Partial Volume Effect Criteria). Record the achieved voxel size as reported by the reconstruction software, which may differ slightly from the target due to final geometric calibration. | Basic | Editable | Numeric (µm) | e.g., 0.85 µm ¦ 7.63 µm ¦ 22.44 µm (isotropic); "µm/pixel" is an acceptable equivalent for isotropic scans; for anisotropic voxels: "5 µm × 5 µm × 8 µm"; µm³ and µm³/voxel denote voxel volume — do not use for voxel size |  | 2026-08-27 | (none) |  | Y | Y |  | 6.00 µm | 0.625 µm (A0180-A); 0.672 µm (A0180-B) | N (voxel size not stated; see Ketcham et al. 2022) | N (voxel size not stated; same as 73002 per paper; see Ketcham et al. 2022) | 38.49 µm voxel edge | 2.8–20.6 µm (optimized per sample) | 10.3 × 10.3 × 10.3 µm³ | 11.54 µm (cubic voxel edge) | 5.34 µm | 2.06 µm/px | 25 µm | 7.7 µm | 0.77–3.5 µm (varies by sample) | 1.923 µm
Rotation Range | Total angular range over which projection images are acquired, in degrees. | Basic | Editable | Numeric (°) | e.g., 360 ¦ 180 |  | 2026-08-27 | (none) | Full 360° rotation is standard for cone-beam lab XCT. 180° rotation (half-scan) is sometimes used for faster acquisition but may introduce additional artifacts. | Y | Y |  | N | 360° | N | N | N | N | N | 360° (continuous) | 360° | N | N | N | N | 360°
Number of Projections | Total number of equally-spaced projection images acquired over the full rotation range. For Mode B (multi-volume stitching), this is the number per sub-volume. | Basic | Editable | Integer | e.g., 1811 ¦ 2046 ¦ 3141 ¦ 3700 |  | 2026-08-27 | (none) | More projections improve reconstruction quality by reducing under-sampling streak artifacts but increase scan time. | Y | Y |  | 4000 | 2401 | N | N | N | N | 3143 | 3141 | N (0.4° step, 360° implied = 900 views; not stated directly) | 1601 | 3141 | 3141 | 1200–2000 (varies by sample) | 2000
Rotation Step Size | Angular increment between successive projection images, in degrees. Equal to Rotation Range divided by Number of Projections when both are reported; however, some sources report step size as the primary rotation parameter without stating the total number of projections explicitly. | Advanced | Editable | Numeric (°) | e.g., 0.115° ¦ 0.4° ¦ 1° |  | 2026-08-27 | (none) | Finer step sizes improve angular sampling and reduce streak artifacts at the cost of longer scan time. | Y | Y |  | N | N | N | N | N | N | N | 0.115° | 0.4°/step | N | N | N | N | N
Exposure Time per Projection | Duration of X-ray exposure for each individual projection image frame, in seconds. Adjustment is warranted for samples that are unusually dense or unusually transparent. If the exposure time was varied across samples within the same procedure or session, report the full range applied (e.g., 0.5–2 s). | Basic | Editable | Numeric (s) | e.g., 0.5 s ¦ 1.42 s ¦ 6 s ¦ 28 s; for variable settings: 0.5–2 s |  | 2026-08-27 | (none) | Set as a deliberate choice for the target sample type and voxel size: longer exposures increase SNR at the cost of total scan time. | Y | Y |  | 4.00 s | 33 s (A0180-A); 28 s (A0180-B) | N | N | N | N | 20 s | 1.00 s per frame | N | 8 s | 0.708 s | 1.415 s | 750–1250 ms (varies by sample) | N
Frames Averaged per Projection | Number of individual detector frames acquired and averaged to produce each saved projection image. The effective exposure per projection = exposure time per frame × frames averaged. | Advanced | Editable | Integer | e.g., 1 ¦ 2 ¦ 3 ¦ 4 |  | 2026-08-27 | (none) | Frame averaging reduces random electronic noise. | Y | Y |  | 1 frame | N | N | N | N | N | N | 8 frames averaged per projection | N | ~200 reference (blank) images taken; number of frames averaged per projection N | 1 frame per projection | 1 frame per projection | 3–6 (varies by sample) | N
Detector Binning | Detector pixel binning factor applied during acquisition. Binning combines adjacent pixels (e.g., 2×2 combines 4 pixels into one). | Advanced | Editable | Controlled list | 1×1 (no binning) ¦ 2×2 ¦ 4×4 ¦ N/A ¦ None |  | 2026-08-30 | (none) | Reduces effective resolution while increasing per-pixel SNR and reducing file size. | Y | Y |  | 1×1 (no binning) | No binning | N | N | N | N | N | N | N | N | N | N | N | N
Flat Field Correction | Whether dark-field (detector read with X-ray source off; electronic noise baseline) and bright-field (source on, no sample; gain calibration) reference images are acquired and applied to normalize detector response before reconstruction. | Basic | Editable | Controlled list / Text | e.g., 'Yes — flat-field applied; residual Poisson noise from flat-fielding noted' ¦ 'No — flux normalization not applied' ¦ N/A ¦ None |  | 2026-08-30 | (none) | Standard practice in all quantitative lab XCT. | Y | Y |  | Flux normalization: no | N | N | N | N | N | N | N | N | N | N | N | N | N
Rotation Mode | Whether projection images are acquired during continuous sample rotation (faster; may introduce slight motion blur at high rotation speeds) or at discrete step positions with the stage stationary during each exposure (stop-and-shoot; eliminates motion blur). | Basic | Editable | Controlled list | Step rotation (stop-and-shoot) ¦ Continuous rotation ¦ N/A ¦ None |  | 2026-08-30 | (none) | Continuous rotation has been associated with sub-volume rotational mismatch artifacts in multi-volume stitching workflows (Eckley et al. 2025). | Y | Y |  | Continuous rotation | N | Continuous rotation (each Subpix sub-acquisition) | Non-continuous rotation (to avoid rotational mismatch of continuous mode) | N | N | N | Continuous 360° rotation | N | N | N | N | N | N
Number of Sub-volumes | Total number of overlapping sub-volume scans acquired to cover the full sample length. Applies only to Mode B (multi-volume stitching). | N/A | Basic | Integer | e.g., 6 ¦ 9 |  | 2026-08-27 | (none) | The number of sub-volumes depends on the length of the specific sample being scanned and cannot be fixed in the procedure in advance. | N | Y |  | N/A | N/A | 6 volumes | 9 volumes | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Minimum Sub-volume Overlap | Minimum number of reconstructed slices that must overlap between adjacent sub-volumes in a stitched dataset. | Basic | Read-Only | Integer | e.g., 500 ¦ 300 |  | 2026-08-27 | (none) | Eckley et al. (2025) report that 500 slices of overlap were used in the Apollo 73001 stitching workflow, while 180 slices proved insufficient for their cone-beam geometry. | N | Y |  | N/A | N/A | ~380 slices overlap between sub-volumes | N | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Total Scan Duration | Total elapsed time for the complete XCT acquisition, from the start of the first projection to the end of the last. For multi-volume acquisitions, this is the combined scan time across all sub-volumes (not including setup or reconstruction time). | N/A | Advanced | Text (free) | e.g., 37 min ¦ 1.5 h ¦ 240 min |  | 2026-08-27 | (none) | Useful for assessing beam stability and sample integrity concerns over long scans. | Y | Y |  | N | N | N | N | N | N | N | N | N | N | 37 min | 75 min | N | N
5. Data Processing | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Reconstruction Algorithm | Mathematical algorithm used to reconstruct the 3D CT volume from 2D projection images. | Basic | Read-Only | Controlled list / Text | FDK (Feldkamp-Davis-Kress, cone-beam) ¦ FBP (filtered back-projection, parallel-beam) ¦ Iterative reconstruction ¦ SART ¦ N/A ¦ None |  | 2026-08-30 | (none) | The Feldkamp-Davis-Kress (FDK) algorithm is standard for cone-beam lab XCT geometry; iterative methods offer better artifact suppression at the cost of computation time. | Y | Y |  | N (Nikon CTPro3D; algorithm not named) | N (Zeiss Reconstructor; algorithm not named) | N (UTCT custom; not named) | N (UTCT custom; not named) | N | N | Filtered back projection (Nikon proprietary) | Filtered back projection (FBP) | N (NRecon; algorithm not named) | N (ZEISS XRM; algorithm not named) | N (CT Pro 3D; algorithm not named) | N (CT Pro 3D; algorithm not named) | N | N
Reconstruction Convolution Filter | Convolution (apodization) filter kernel applied during back-projection reconstruction. | Advanced | Editable | Text (free) | e.g., 'Hann filter' ¦ 'Ram-Lak filter' ¦ 'Shepp-Logan filter' ¦ 'Hamming filter' ¦ 'Software default (not reported)' |  | 2026-08-27 | (none) | Sharper filters (Ram-Lak) enhance edge definition but amplify high-frequency noise; smoother filters (Hann, Hamming) reduce noise at the cost of edge sharpness. | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Beam Hardening Correction Method | Strategy used to reduce beam hardening artifacts caused by preferential attenuation of low-energy photons in polychromatic cone-beam XCT. The specific correction parameter value is recorded separately. | Basic | Read-Only | Controlled list / Text | e.g., 'Hardware pre-filtering only (no software correction)' ¦ 'Software polynomial correction (instrument-specific algorithm)' ¦ 'Ketcham & Hanna (2014) iterative BH correction' ¦ 'None applied' ¦ N/A |  | 2026-08-27 | (none) | Beam hardening causes cupping artifacts (CT numbers elevated at sample edges relative to the interior) and can bias segmentation thresholds. Hardware pre-filtering (Group 4: X-ray Pre-filter) and software correction are complementary approaches. | Y | Y |  | Hardware filter (0.25 mm Al) + software BHC preset | Hardware filter (LE4 inbuilt) | Software BHC applied during reconstruction | Software BHC (same factor as 73002 re-used) | N | N | N (not mentioned) | None (intentionally unfiltered; no software BHC mentioned) | Hardware filter (1.0 mm Al); software BHC in NRecon not mentioned | N | None (conditions did not saturate detector; not required) | None (conditions did not saturate detector; not required) | Software BHC applied per respective software | N
Beam Hardening Correction Parameter | Numerical value or setting applied in the software beam hardening correction algorithm for this specific analysis. Companion to Beam Hardening Correction Method. | Advanced | Editable | Text (free) | e.g., '0.05 (constant correction in Zeiss Reconstructor; Zhang et al. 2021)' ¦ 'BH reduction level = 3 (NRecon, Bruker)' ¦ 'Not applicable — hardware filtering only' |  | 2026-08-27 | (none) | May be tuned empirically per material type or per scan session. | Y | Y |  | Preset 4 | N | N | N | N | N | N | N | N | N | N | N | N | N
Ring Artifact Correction Method | Procedure specification for how ring artifacts are handled. Whether correction was applied and its outcome are recorded separately in Group 6. | Advanced | Editable | Text (free) | e.g., 'Software ring correction applied during reconstruction (instrument algorithm)' ¦ 'Post-reconstruction ring artifact filter (ImageJ ring artifact plugin)' ¦ 'Defective pixel masking applied before reconstruction' ¦ 'None specified' |  | 2026-08-27 | (none) | Ring artifacts appear as concentric circular bands centred on the rotation axis and arise from defective or miscalibrated detector pixels. | Y | Y |  | N | N | Post-reconstruction ring corrections applied | Post-reconstruction ring corrections applied | N | N | N | N | N | N | N | N | Software ring artifact correction filter applied during analysis | N
Output Data Format | File format of the reconstructed CT volume as output by the reconstruction software. Bit depth is recorded separately in Output Bit Depth. | Basic | Editable | Controlled list / Text | TIFF image stack ¦ VGI/VOL (VGStudio) ¦ RAW volume ¦ DICOM ¦ N/A ¦ None |  | 2026-08-30 | (none) |  | Y | Y |  | TIFF (16-bit, 928 slices) | N | N | N | N | N | N | TIFF (continuous series of 2D TIFF images) | N | N | N | N | N | N
CT Number Calibration | Whether the raw CT grayscale values have been calibrated to physically meaningful units using reference materials. | Advanced | Editable | Controlled list / Text | e.g., 'None — raw instrument grayscale' ¦ 'Compared to MuCalc calculated LAC values (Hanna & Ketcham 2017)' ¦ 'Calibrated to LAC (cm⁻¹) using air and known mineral standards' ¦ 'Calibrated to Hounsfield Units (HU) relative to water' ¦ N/A |  | 2026-08-27 | (none) | Calibrated CT numbers enable cross-instrument and cross-session comparison and support phase identification by comparison to calculated linear attenuation coefficient (LAC) values (e.g., using MuCalc; Hanna & Ketcham 2017). Uncalibrated grayscale values are instrument- and session-specific. | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Segmentation Method | Method and software used to separate distinct phases or features in the reconstructed 3D volume, turning the grayscale volume into labelled regions. | Basic | Editable | Controlled list / Text | Global threshold (single grayscale boundary) ¦ Multi-threshold (separate range per phase) ¦ Manual threshold per phase ¦ Semi-automated: threshold + morphological filtering ¦ Region growing ¦ Manual tracing ¦ Machine learning (specify model) ¦ N/A ¦ None |  | 2026-08-26 | (none) | The method must be reported for any quantitative result to be reproducible. | Y | Y |  | N | ImageJ threshold-based; TANGO plugin for 3D object detection | N | N | N | N | Manual segmentation of 2D slices; grayscale threshold applied to 3D volume | N (no segmentation performed; dose assessment study only) | Adaptive (mean) thresholding | Grayscale threshold range selection (ImageJ) | Manual grayscale threshold | Manual grayscale threshold | Manual grayscale threshold per phase | Dedicated workflow (Godel 2013) modified for plagioclase network segmentation
Segmentation Threshold Values or Criteria | Specific CT number range(s) or quantitative criteria used to define each segmented phase or feature. For LAC-calibrated datasets, report values in cm⁻¹. | Advanced | Editable | Text (free) | e.g., 'Voids: <500; Silicates: 500–1800; Opaques: >1800 (16-bit grayscale)' ¦ 'LAC: Voids <3 cm⁻¹; Matrix 3.5–5 cm⁻¹; Carbonates 5.5–6.5 cm⁻¹; Opaques >9 cm⁻¹' |  | 2026-08-27 | sample > sampling unit | Documenting these values enables reproducibility assessment and cross-study comparison. | Y | Y |  | N | N (visual sulphide-rim contrast; numerical thresholds N) | N | N | N | N | Thresholds not given numerically; SEM images used to define phases; errors from threshold + subvoxel variability generally <5% for grains >125 voxels | N | Adaptive mean threshold; total analyzed VOI for NWA 8277 = 5.39 mm³ | Threshold range selected to include all voxels of one phase while excluding adjacent phases; specific values N | Manual threshold: vapor = darker, liquid = brighter pixels; specific values N | Manual threshold: vapor = darker, liquid = brighter pixels; specific values N | Manual threshold per phase; specific values N | Dedicated workflow; specific threshold values N
Phase Identification Method | Method used to assign reconstructed CT number ranges to specific mineral phases or material types. Approaches include comparison to calculated linear attenuation coefficients (LAC), cross-validation with independent analytical techniques, or empirical calibration. | Advanced | Editable | Text (free) | e.g., 'Comparison to calculated LAC values using MuCalc (Hanna & Ketcham 2017)' ¦ 'Cross-validated against SEM-EDS phase maps' ¦ 'Empirical calibration using known mineral standards (olivine, chromite)' ¦ 'Not performed — structural analysis only' |  | 2026-08-12 | sample > sampling unit |  | Y | Y |  | N | Sulphide rims of SSOs identified by bright contrast; silicate glass by contrast relative to matrix | N | N | N | N | Phases: augite, mesostasis, olivine, titanomagnetite, pore space — identified by grayscale correlated to linear attenuation coefficients; verified with SEM | N | Three density regions: high, medium, low density; porosity separate | Phase identified by grayscale contrast (vapor vs. silicate/mineral) | Phase identified by grayscale contrast (vapor darker, liquid brighter) | Phase identified by grayscale contrast (vapor darker, liquid brighter) | Phases by grayscale contrast; minerals by differential attenuation | Plagioclase targeted; olivine + pyroxene not separately resolved; troilite/oxyhydroxide indistinguishable
VOI Selection Criteria | Rules specifying how the Volume of Interest (VOI) is to be defined for quantitative analysis. Common criteria exclude cone-beam artifact zones at sample edges, beam hardening halos near dense inclusions, and sample holder signal. The actual VOI applied in a specific analysis is recorded separately at analysis level. | Advanced | Editable | Text (free) | e.g., 'Exclude outer 15% of sample radius to avoid cone-beam artifacts' ¦ 'Exclude sample holder signal; VOI diameter = 95% of sample diameter' ¦ 'No restriction — full reconstructed volume used' |  | 2026-08-27 | (none) |  | Y | Y |  | N | N (full scan volume) | N (full core length per sub-volume) | N (full core length per sub-volume) | N | N | N (entire chip volume; six 2D slices at ~1 mm spacing selected for modal analysis) | N (full vial volume) | N (total analyzed volume stated as 5.39 mm³ for NWA 8277) | N (full scan volume) | N (full scan volume) | N (full scan volume) | N (full scan volume) | N (full 8 mm core)
VOI Applied | Actual Volume of Interest used, including dimensions or defining criteria. Companion to VOI Selection Criteria. | N/A | Basic | Text (free) | e.g., 'Cylindrical VOI, radius = 85% of sample radius, full height' ¦ 'Full reconstructed volume' ¦ 'Spherical VOI, 2 mm radius, centred on primary inclusion' |  | 2026-08-27 | (none) |  | Y | Y |  | N | Full scan volume | Full core length per sub-volume | Full core length per sub-volume | N | N | Entire chip volume (~250 thin-section equivalents); six interspaced 2D slices at ~1 mm spacing for modal mineralogy | Full vial volume (2000 × 2000 × 2000 voxels) | 5.39 mm³ total analyzed volume (NWA 8277) | Full scan volume | Full scan volume | Full scan volume | Full scan volume per sample | Full 8 mm core volume
Sub-volume Stitching and Registration Method | Method used to register adjacent sub-volume datasets to each other and stitch them into a single continuous 3D volume. Report the alignment strategy (manual, automated, fiducial-based), the software used, and any correction steps applied. Where rotational mismatch has been corrected via raw projection re-alignment, document it here. | Basic | Editable | Text (free) | e.g., 'Overlap cross-correlation; rotational mismatch corrected via raw projection re-alignment (Eckley et al. 2025); VGStudio Max 3.4' ¦ 'Automated stitching, VGStudio Max' ¦ 'FIJI BigStitcher plugin' |  | 2026-08-27 | (none) | Rotational mismatch artifacts have been identified in continuous-rotation acquisitions (~0.35° misalignment in Eckley et al. 2025). | N | Y |  | N/A | N/A | Digital stitching (method not named); post-reconstruction ring and distortion corrections applied | Digital stitching; manual realignment of raw projection images for rotational misalignment (up to 0.35°); additional cone-beam, BH, and scattering corrections applied | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Sub-volume Overlap | Actual number of reconstructed slices overlapping between adjacent sub-volumes. May differ from the minimum sub-volume overlap required specified in the procedure (Group 4) if sample geometry or operator decisions resulted in greater or lesser overlap. | N/A | Basic | Integer | e.g., 500 ¦ 300 ¦ 180 |  | 2026-08-27 | (none) |  | N | Y |  | N/A | N/A | ~380 slices per sub-volume overlap | N | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Output Bit Depth | Bit depth of the reconstructed 3D volume (number of bits used to encode each voxel's grayscale value). Common values are 8-bit (256 gray levels), 16-bit (65,536 gray levels), or 32-bit floating point. A required output bit depth may be specified if downstream analysis workflows depend on a consistent grayscale range. | Advanced | Editable | Controlled list | e.g., 16-bit ¦ 32-bit ¦ 8-bit ¦ N/A ¦ None |  | 2026-08-30 | (none) | Bit depth affects the dynamic range available for phase segmentation and quantitative attenuation analysis. | Y | Y |  | 16-bit (full grayscale range) | 16-bit (CCD pixel depth) | N | N | N | N | N | N | N | 32-bit (ZEISS XRM output) | N | N | N | N
Calibration Factor and Determination Method | An externally-calibrated factor that converts the measured quantity into the reported quantity, how it was determined, and its uncertainty. Applies where the conversion depends on a factor calibrated against a reference of independently known value, rather than on the instrument response alone. Distinct from the fields that name the calibration material and that state which approach applies to which target species, where the technique has them: this field records the resulting factor itself. | Basic | Editable | Text (free) |  | Source: Calibration Factor module | 2026-09-01 | reported property |  | Y | Y |  | N | N | N | N | N | N | N (segmentation performed on uncalibrated grayscale values; no attenuation calibration stated) | N | N | N (segmentation performed on relative grey values; no attenuation calibration stated) | N (segmentation performed on relative grey values; no attenuation calibration stated) | N (segmentation performed on relative grey values; no attenuation calibration stated) | N (segmentation performed on relative grey values; no attenuation calibration stated) | N
Constants and Reference Values Used | Physical constants and reference values used in data reduction to calculate the final reported quantity (e.g., decay constants for age calculation, standard isotope ratios, or other citable reference values used in a correction or calculation), together with their source. Distinct from the Group 6 reference-material fields, which document accepted values for specific calibration/validation materials rather than universal physical constants. Record "None" if no citable, revisable physical constants feed into this procedure's data reduction. | Basic | Editable | Text (free) | e.g., "λ238U = 1.55125×10⁻¹⁰ yr⁻¹, λ235U = 9.8485×10⁻¹⁰ yr⁻¹ (Jaffey et al. 1971); 238U/235U = 137.818 (Hiess et al. 2012)" ¦ "None" | Source: Core module | 2026-09-01 | (none) |  | Y | Y |  | N | N | N | N | N | N | N (minerals discriminated by contrasting linear attenuation coefficients; no coefficient values stated) | N | N (mass attenuation coefficients are discussed for the micro-XRF Rayleigh/Compton work, not for the micro-CT) | N | N | N | N | N (olivine and pyroxene were not segmented owing to similar attenuation coefficients under the imaging conditions; no values stated)
6. Quality Control & Uncertainty | [9 empty] | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Effective Spatial Resolution (PSF/MTF) | Effective spatial resolution of the reconstructed CT volume, which typically differs from (and is coarser than) the voxel size due to the detector point spread function (PSF), geometric unsharpness, and the reconstruction filter. Formal measurement uses the PSF method described in Ketcham & Hildebrandt (2014) or the modulation transfer function (MTF); resolution can also be reported per ASTM E1441-11. When formal measurement is unavailable, an estimate based on acquisition geometry (e.g., the Brenner-Weiss formula) or a stated multiple of the voxel size may be recorded. Report the value in µm along with the method used to determine or estimate it. Note that most XCT papers use the bare phrase "spatial resolution" for the voxel size; record that under Voxel Size, and leave this field as not reported unless an effective resolution was actually determined. | N/A | Advanced | Text (free) | e.g., '~5 µm effective (PSF measured via wire phantom at 1.295 µm voxel; Ketcham & Hildebrandt 2014)' ¦ '~30 µm estimated (3x voxel size; not formally measured)' ¦ 'PSF FWHM = 8.2 µm at 2.5 µm voxel (Brenner-Weiss formula)'; report in µm; µm/pixel is equivalent for cubic voxels; µm³/voxel denotes voxel volume and is not a resolution unit |  | 2026-08-27 | (none) | The Nyquist limit sets a theoretical floor at 2× the voxel size. | Y | Y |  | N (voxel 6 µm; no formal spatial resolution measurement stated) | N — the paper's "spatial resolutions (in voxels)" are the voxel sizes (0.625 / 0.672 µm); no effective (PSF/MTF) resolution reported | N (voxel size not stated in this paper; see Ketcham et al. 2022) | N (voxel size not stated; same as 73002 per paper; see Ketcham et al. 2022) | N — voxel size only (38.49 µm); no effective resolution reported. The paper's "spatial resolution of 60 µm" is the multispectral core imager, not this XCT scan | N — voxel size only (2.8–20.6 µm); no effective resolution reported. The paper's "spatial resolution of 60 µm" is the multispectral core imager, not this XCT scan | N — the paper states "a resolution of 10.3 x 10.3 x 10.3 µm³ per voxel", i.e. the voxel size; no effective resolution reported | ~30 µm — stated as "resolution limits of ~30 µm (around 3x the voxel size)" against an 11.54 µm voxel. Rule-of-thumb estimate; no PSF or MTF measurement | N — voxel size only (5.34 µm); the paper describes resolution only generically ("of the order of microns") | N — the paper's "spatial resolution of 2.06 µm/px (8.7 µm³/vx)" is the voxel size; no effective resolution reported | N — the paper's "spatial resolution" figures are given in µm/px with voxel volumes, i.e. the voxel size (25 µm); no effective resolution reported | N — the paper's "spatial resolution" figures are given in µm/px with voxel volumes, i.e. the voxel size (7.7 µm); no effective resolution reported | N — the paper's "spatial resolution" figures are given in µm/px with voxel volumes, i.e. the voxel size (0.77–3.5 µm); no effective resolution reported | N — voxel size only (1.923 µm); no spatial resolution reported
Beam Hardening Artifact Assessment | Assessment of whether residual beam hardening (cupping artifact) is present in the reconstructed volume after any corrections applied during reconstruction. Record whether this artifact is detectable in the final dataset, what correction or mitigation was applied (e.g., linearisation, pre-filtering, iterative correction), and whether any residual effect influences the analysis region. | N/A | Advanced | Text (free) | e.g., 'Cupping visible at 100 kV; minimized at 160 kV (Wu et al. 2024)' ¦ 'Negligible — sample diameter < 3× attenuation length at selected voltage' ¦ 'Corrected; residual <2% CT number variation across sample diameter' |  | 2026-08-27 | (none) | Beam hardening in polychromatic lab XCT produces a characteristic darker interior and brighter edges in the reconstructed CT number profile, even in compositionally uniform material. | Y | Y |  | N | N | N (corrections applied; extent not documented) | BH artifact from steel sleeve: uneven brightness and scattering; asymmetric + symmetric corrections applied | N | N | High noise in central portion of sample attributed to BH; augite/mesostasis difficult to distinguish in central region | N | N | N | N | N | N | N
Ring Artifact Severity and Correction Outcome | Assessment of whether ring artifacts — concentric rings centred on the rotation axis arising from differential sensitivity or gain drift among detector elements — were present and whether correction was effective. Note that ring correction algorithms modify image intensity in narrow annular bands. Record whether rings were present, whether correction was applied, and whether any residual rings or correction artifacts fall within the analysis volume of interest. | N/A | Advanced | Text (free) | e.g., 'Minor ring artifacts in outer 5% of sample radius; ring correction applied in Avizo; no residual rings in analysis VOI' ¦ 'No ring artifacts detected' ¦ 'Rings present near sample centre; correction applied but introduced banding in carbonate vein region; vein excluded from modal analysis' |  | 2026-08-27 | (none) | In samples containing linear geological features oriented tangentially to the rotation axis, ring correction can alter or introduce spurious linear features in those orientations. | Y | Y |  | N | N | Post-reconstruction ring corrections applied; severity not documented | Post-reconstruction ring corrections applied; severity not documented | N | N | N | N | N | N | N | N | Software ring artifact correction applied; outcome not documented | N
Metal Streak Artifact Assessment | Assessment of whether high-density inclusions (FeNi metal, platinum-group minerals, dense oxides, or other high-Z phases) produce streak or starburst artifacts in the reconstructed volume. Common forms include radial starburst streaks from high-Z inclusions and shadowing (anomalously low CT number) in beam-shadow regions behind dense phases. Record which artifact type(s) are present, their spatial extent, and any exclusion zones or correction steps applied. | N/A | Advanced | Text (free) | e.g., 'FeNi grains produce starburst artifacts; 10-voxel exclusion zone applied around each metal grain; silicate analysis unaffected' ¦ 'Shadowing artifact in olivine region behind 1.5 mm FeNi inclusion; region excluded from modal analysis' ¦ 'No metal streak artifacts — meteorite lacks metallic phases' |  | 2026-05-15 | (none) |  | Y | Y |  | N | N | N | Steel outer sleeve caused secondary BH and scattering artifact resulting in uneven brightness; corrected with asymmetric + symmetric corrections | N | N | N | N | N | N | N | N | N | N
Partial Volume Effect Criteria | Specification of how partial volume effects (PVE) are managed in quantitative analysis. PVE correction can be implemented via PSF-based deconvolution tools such as Blob3D. Record the minimum feature size criterion adopted for the procedure (in voxels or µm), the basis for it, the treatment of boundary voxels in modal abundance or size distribution calculations, and whether PVE correction is required or optional. State whether the criterion follows the Withers et al. (2021) convention — a feature must span at least 3 voxels to be positively identified and at least 10 for reliable shape and volume characterisation — or is SNR-limited, PVE-limited or analyst-defined. | Advanced | Editable | Text (free) | e.g., 'Features < 5 voxels equivalent diameter excluded from size distribution analysis' ¦ 'Minimum 3-voxel border exclusion around phase boundaries for modal abundance calculation' ¦ 'Not specified — analyst judgement at analysis level' |  | 2026-08-27 | sample > sampling unit | PVE produces intermediate CT numbers at phase boundaries and at the surfaces of small features, because each voxel integrates the attenuation of all material within its volume. The criterion materially changes reported modal abundances and size distributions, so two datasets are not comparable without it. | Y | Y |  | N | SSOs ≥ 5.4 µm identified (implicit: features ≥ voxel size = 0.625 µm); sulphide rims ≥ 1 voxel required for detection | N | N | N | N | Errors from threshold + PVE generally <5% for grains >125 voxels; features smaller than ~3 voxels unreliable | N | N | N | N | N | N | N
Partial Volume Effect Assessment | Record of PVE severity and how it was handled. Documents the minimum feature size relative to the voxel size, the fraction of the feature population affected, and any PVE correction method applied (e.g., PSF-based deconvolution via Blob3D). | N/A | Advanced | Text (free) | e.g., 'Features <5 voxels excluded; affects ~3% of vesicle population by count' ¦ 'PVE correction applied (Blob3D PSF deconvolution); vein width recoverable to ±0.5 voxel' ¦ 'PVE negligible — smallest features of interest are >20× voxel size' |  | 2026-08-27 | (none) |  | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Signal-to-Noise Ratio | Quantitative measure of image quality in the reconstructed CT volume. Typically calculated as the mean CT number divided by the standard deviation of CT number in a homogeneous reference region (e.g., an inclusion-free zone of the matrix). Record the measured value and the region used for measurement, or note if SNR was not formally measured. | N/A | Advanced | Text (free) | e.g., 'SNR = 35 (measured in glass matrix region)' ¦ 'Not formally measured' |  | 2026-08-27 | (none) | SNR is controlled by photon flux (source power, exposure time, number of projections) and sample attenuation. | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Cross-Validation Procedure Requirement | Specification of what independent analytical validation is required to confirm CT segmentation results, phase identification, or quantitative measurements. Common approaches include BSE imaging, SEM-EDS or EPMA modal analysis, He pycnometry for bulk porosity, and Raman or SIMS phase mapping. Record the required validation method(s) and the sampling fraction (e.g., every sample, one per session, or a representative subset). | Advanced | Editable | Text (free) | e.g., 'CT modal abundances to be validated against SEM-EDS phase maps for ≥1 sample per session' ¦ 'CT porosity to be compared against Archimedes density measurement' ¦ 'None specified' |  | 2026-08-27 | (none) | Cross-validation is relevant where CT grey-values alone cannot uniquely distinguish phases with similar attenuation coefficients. | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N
Cross-Validation Outcome | Record of what independent validation was performed and its result. Report the validation method(s) used, the quantitative agreement achieved (e.g., relative difference in modal abundance, porosity), and any discrepancies and their likely causes. Where BSE imaging is the validation method, note that BSE provides a 2D section while CT provides a 3D volume. | N/A | Advanced | Text (free) | e.g., 'CT modal abundances compared to EPMA WDS phase maps: mean difference 4.2% for olivine, 6.8% for low-Ca pyroxene' ¦ 'CT bulk porosity 8.3% vs. He pycnometry 7.9% (difference 5.1%)' ¦ 'Cross-validation not performed — sample insufficient for destructive sectioning' |  | 2026-08-27 | (none) |  | Y | Y |  | N | 61 SSOs by polished SEM sections vs. 108 by XCT; discrepancy attributed to disruption by polishing; XCT detects more SSOs | N (XCT pre-dates dissection; qualitative comparison with QEMSCAN post-dissection) | N | N | N | BSE contrast differences between phases comparable to XCT attenuation contrast; XCT lower spatial resolution | N (dose experiment; organic chemistry of scanned vs. unscanned split compared) | Micro-XRF structural observations consistent with meteorite classification | N | N | N | Sample I: CLSM gives φ_vap = 10±2%; XCT gives 9.8% — compatible | XCT qualitatively compared with thin sections and EBSD; cm-scale plagioclase confirmed
Additional Notes | Any procedure- or analysis-specific information not captured by a structured field anywhere in this TAPP — including anomalies, deviations from the registered procedure, instrument modifications, and supplementary context. Scope is the whole document, not Group 6: this is the last field of the TAPP and covers all six groups. Use sparingly; a structured field is preferred for anything that can be formally categorised. | Advanced | Advanced | Text (free) | e.g., 'Detector replaced mid-session; scans 14-22 acquired on the new detector' ¦ 'None' | Source: Core module | 2026-09-01 | (none) |  | Y | Y |  | N | N | N | N | N | N | N | N | N | N | N | N | N | N

## Legends
Procedure-Level Tier | Description | Analysis-Level Tier | Description
Basic | Mandatory for procedure registration. Must be provided to register a valid procedure. | Read-Only | Directly imported from the registered procedure; cannot be changed by the analyst. Changing this value means running a different procedure. Also used for fields relevant only at procedure level — the value is inherited from the procedure record and shown read-only in the analysis form.
Advanced | Optional for procedure registration. Strongly recommended but not required. | Editable | Imported from the registered procedure but may be adjusted within procedure-defined bounds (e.g., daily tuning, minor software updates). The procedure registers the target or typical value; the analyst confirms or adjusts it. Cannot be left void if the procedure-level tier is Basic.
N/A | Not applicable at procedure level. This field captures analysis-level information only. | Basic | Mandatory user input at analysis time. Value comes from the analysis itself and cannot be pre-specified in the procedure.
|  | Advanced | Optional user input at analysis time. Recommended for complete and reproducible documentation.
Mode Column | Definition
Purpose
Single-volume
Multi-volume stitching
Y | This field applies to this analytical mode and should be reported.
N | This field does not apply to this analytical mode.
Data Type | What to enter
Controlled list | Choose one of the values listed in "Example / Allowed Content". The list is CLOSED — it is meant to cover every case. If your value genuinely is not there, that is a gap in the list worth reporting, not a reason to write your own.
Controlled list / Text | Choose a listed value AND qualify it in the same cell — the detail is expected, not optional. Typically what was corrected, which target species or phases it applied to, the method or equation used, or the source it came from. An answer the list cannot express is also valid here.
Numeric (unit) | A number in the unit named in brackets. Enter the number only — the unit is fixed by the field.
Numeric + unit | A number AND its unit, because the unit varies between procedures. Write both, e.g. '50 us' or '0.5 s'.
Text (free) | Free text. No controlled vocabulary applies.
N/A ¦ None | Offered by every controlled list. "N/A" means the field does not apply to this procedure; "None" means it applies but nothing was used. Prefer either over leaving the cell blank.
Keyed By (Rule 7) | What the field's value repeats over
(none) | Scalar — one value per procedure or analysis. The default.
defines: reported property | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
defines: sample | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it.
defines: sample > sampling unit | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per subdivision of the physical sample that carries its own row — grain, spot, aliquot, phase, sub-volume.
defines: target material | ENUMERATES the key domain rather than being keyed by it — the header of the child table, not a column in it. One value per material type the procedure is designed to analyse, as listed in Target Material.
reported property | One value per reported quantity or nominal property, at any point in the chain — ratios and dates alike, plus their uncertainties.
sample
sample > sampling unit | Nested — sampling unit exists only within sample. One value per subdivision of the physical sample that carries its own row — grain, spot, aliquot, phase, sub-volume.